## Data Set Properties
Data Set Properties
Data Set Name: | Conway2024v61is
Hospital Name: | Conway Medical Center
Status: | Released
Data Set ID: | 02d298b43-R

## ContBolus Non-Anesth Drugs
Continuous/Bolus Non-Anesthesia Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount (- - = custom concentration) | Drug Amount Units | Diluent Volume (- - = custom concentration) | Diluent Volume Units | Final Concentration | Final Concentration Units | Module Availability | Dosing Units | Dose Limits Soft Min | Dose Limits Soft Max | Dose Limits Hard Max | Dose Initial Value | Bolus Dose Units | Bolus Dose Limits Soft Min | Bolus Dose Limits Soft Max | Bolus Dose Limits Hard Max | Bolus Dose Limits Initial Value | Bolus Dose Admin. Rate Units | Bolus Dose Admin. Rate Limits Soft Min | Bolus Dose Admin. Rate Limits Soft Max | Bolus Dose Admin. Rate Limits Hard Max | Bolus Dose Admin. Rate Initial Value | Conc. Limit Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max | Clinical Advisory
1 | Adult | adenosine (CATH LAB) | .Standard wt < 115kg | 90 | mg | 90 | mL | 1 | mg/mL | Pump | mcg/kg/min | 135 | 145 |  | 140 | [10 empty] | n/a
2 | Adult | adenosine (CATH LAB) | Weight >= 115 kg | 180 | mg | 90 | mL | 2 | mg/mL | Pump | mcg/kg/min | 135 | 145 |  | 140 | [10 empty] | n/a |  |  |  | Double Concentrated
3 | Adult | alteplase -stroke | Stroke < 101 kg | - - | mg | - - | mL |  | mg/mL | Pump | mg/kg/h | 0.7 | 0.9 |  | 0.81 | [10 empty] | mg/mL | 0.85 | 0.99 | 1.01 | Alteplase wt based
4 | Adult | alteplase -stroke | Stroke 101kg or more | 81 | mg | 81 | mL | 1 | mg/mL | Pump | mg/h | 71 | 90 |  | 81 | [10 empty] | n/a
5 | Adult | amiodarone Drip |  | 360 | mg | 200 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.51 | [11 empty] | n/a |  |  |  | In-line filter
6 | Adult | amiodarone Drip |  | 450 | mg | 250 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.51 | [11 empty] | n/a |  |  |  | In-line filter
7 | Adult | amiodarone Drip |  | 900 | mg | 500 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.51 | [11 empty] | n/a |  |  |  | In-line filter
8 | Adult | argatroban | .Standard Dose | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 10 | 11.5 | 2 | [10 empty] | n/a
9 | Adult | argatroban | .Standard Dose | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 10 | 11.5 | 2 | [10 empty] | n/a
10 | Adult | argatroban | Hepatic Dose | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.1 | 5 | 10.1 | 0.5 | [10 empty] | n/a
11 | Adult | argatroban | Hepatic Dose | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.1 | 5 | 10.1 | 0.5 | [10 empty] | n/a
12 | Adult | argatroban | Undergoing PCI | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 15 | 30 | 34 | 25 | mcg/kg | 149 | 351 |  |  | mcg/kg/min | 29 | 117 |  |  | n/a
13 | Adult | argatroban | Undergoing PCI | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 15 | 30 | 34 | 25 | mcg/kg | 149 | 351 |  |  | mcg/kg/min | 29 | 117 |  |  | n/a
14 | Adult | atracurium |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 0.5 | 29.5 | 39 |  | mg/kg | 0.07 | 0.5 | 0.58 |  | mg/kg/min | 0.07 | 0.25 | 0.29 |  | mcg/mL | 200 | 5000 | 5100 | Neuromuscular BA
15 | Adult | bivalirudin | .Standard Dose | 250 | mg | 50 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.1 | 1.8 | 2 | 1.75 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.76 |  | 0.75 | n/a
16 | Adult | bivalirudin | .Standard Dose | 500 | mg | 100 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.1 | 1.8 | 2 | 1.75 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.76 |  | 0.75 | n/a
17 | Adult | bivalirudin | .Standard Dose | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.1 | 1.8 | 2 | 1.75 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.76 |  | 0.75 | n/a
18 | Adult | bivalirudin | ESRD | 250 | mg | 50 | mL | 5 | mg/mL | Pump | mcg/kg/h | 0.1 | 1.2 | 1.4 | 0.25 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
19 | Adult | bivalirudin | ESRD | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/h | 0.1 | 1.2 | 1.4 | 0.25 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
20 | Adult | bivalirudin | Renal Dose CrCl < 30 | 250 | mg | 50 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.24 | 1.2 | 1.4 | 1 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
21 | Adult | bivalirudin | Renal Dose CrCl < 30 | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.24 | 1.2 | 1.4 | 1 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
22 | Adult | bivalirudin | Renal Dose CrCl < 30 | 500 | mg | 100 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.24 | 1.2 | 1.4 | 1 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
23 | Adult | bumetanide continuos |  | 24 | mg | 96 | mL | 0.25 | mg/mL | Pump | mg/h | 0.1 | 2 |  | 0.5 | mg | 0.25 | 4 |  |  | mg/min | 0.25 |  | 0.5 |  | n/a |  |  |  | Protect From Light
24 | Adult | cangrelor |  | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 2 | 4.2 | 4.6 | 4 | mcg/kg | 26 |  | 35 | 30 | mcg/kg/min | 26 |  | 35 |  | n/a
25 | Adult | cisatracurium |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | mcg/kg/min | 0.5 | 5 | 7 | 3 | [10 empty] | n/a |  |  |  | Neuromuscular BA
26 | Adult | conivaptan |  | 20 | mg | 100 | mL | 0.2 | mg/mL | Pump | mg/h | 0.6 | 1.7 | 2 | 0.8 | mg | 10 | 20 |  | 20 | mg/min | 0.333 | 0.7 |  | 0.667 | n/a
27 | Adult | conivaptan |  | 20 | mg | 250 | mL | 0.08 | mg/mL | Pump | mg/h | 0.6 | 1.7 | 2 | 0.8 | mg | 10 | 20 |  | 20 | mg/min | 0.333 | 0.7 |  | 0.667 | n/a
28 | Adult | dexmedetomidine |  | 200 | mcg | 50 | mL | 4 | mcg/mL | Pump | mcg/kg/h | 0.05 | 0.7 | 1.5 | 0.2 | mcg/kg | 0.9 | 1.1 | 2 | 1 | mcg/kg/min | 0.01 | 0.11 | 0.2 | 0.1 | n/a
29 | Adult | dexmedetomidine |  | 400 | mcg | 100 | mL | 4 | mcg/mL | Pump | mcg/kg/h | 0.05 | 0.7 | 1.5 | 0.2 | mcg/kg | 0.9 | 1.1 | 2 | 1 | mcg/kg/min | 0.01 | 0.11 | 0.2 | 0.1 | n/a
30 | Adult | dilTIAZem |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | mg/h | 0.5 | 15 | 25 | 5 | [10 empty] | n/a
31 | Adult | dilTIAZem |  | 125 | mg | 125 | mL | 1 | mg/mL | Pump | mg/h | 0.5 | 15 | 25 | 5 | [10 empty] | n/a
32 | Adult | DOBUTamine | ..Standard | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a
33 | Adult | DOBUTamine | ..Standard | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a
34 | Adult | DOBUTamine | .Double Concentrated | 500 | mg | 250 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a |  |  |  | Double Concentrated
35 | Adult | DOBUTamine | .Double Concentrated | 1000 | mg | 500 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a |  |  |  | Double Concentrated
36 | Adult | DOBUTamine | Echo Stress Test | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 4.9 | 50 |  | 5 | [10 empty] | mcg/mL | 900 |  | 5000
37 | Adult | DOBUTamine | Stress Test | 500 | mg | 250 | mL | 2 | mg/mL | Pump | mcg/kg/min | 5 | 50 |  | 10 | [10 empty] | n/a
38 | Adult | DOPamine | ..Standard | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 0.9 | 20 | 30 | [11 empty] | n/a
39 | Adult | DOPamine | .Double Concentrated | 800 | mg | 250 | mL | 3.2 | mg/mL | Pump | mcg/kg/min | 0.9 | 20 | 30 | [11 empty] | n/a |  |  |  | Double Concentrated
40 | Adult | DOPamine | Renal Dose | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 0.9 | 5 | 6 | 1 | [10 empty] | n/a
41 | Adult | doxapram |  | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mg/min | 0.1 | 6 | [12 empty] | n/a
42 | Adult | doxapram |  | - - | mg | - - | mL |  | mg/mL | Pump | mg/min | 0.1 | 6 | [12 empty] | mg/mL | 0.89 | 0.9 | 2.1
43 | Adult | epidural w/ fentanyl |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 0.01 | 12 | 15 |  | unit | 2 | 7 | 8 | 5 | unit/min | 4 | 14 | 16 | 10 | n/a |  |  |  | 1 unit = 1 mL
44 | Adult | epidural w/o fentany |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 0.01 | 12 | 15 |  | unit | 2 | 7 | 8 | 5 | unit/min | 4 | 14 | 16 | 10 | n/a |  |  |  | 1 unit = 1 mL
45 | Adult | EPINEPHrine | ..Standard | 5 | mg | 250 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 40 | [12 empty] | n/a |  |  |  | Dedicated IV Line
46 | Adult | EPINEPHrine | ..Standard | 10 | mg | 500 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 40 | [12 empty] | n/a |  |  |  | Dedicated IV Line
47 | Adult | EPINEPHrine | Cardiac Arrest | 30 | mg | 250 | mL | 0.12 | mg/mL | Pump | mcg/min | 0.33 | 40 |  | 33.33 | [10 empty] | n/a |  |  |  | Dedicated IV Line
48 | Adult | epoprostenol |  | 1.5 | mg | 50 | mL | 0.03 | mg/mL | Syringe | nanogram/kg/min | 5 | 51 |  | 50 | [10 empty] | n/a
49 | Adult | eptifibatide | .Standard & wt>121kg | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mg/h | 13 | 17 |  | 15 | [10 empty] | n/a |  |  |  | SCr<2.0
50 | Adult | eptifibatide | .Standard & wt>121kg | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mg/h | 13 | 17 |  | 15 | [10 empty] | n/a |  |  |  | SCr<2.0
51 | Adult | eptifibatide | .Standard Dose | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mcg/kg/min | 0.5 | 2.2 | 2.5 | 2 | [10 empty] | n/a |  |  |  | SCr<2.0
52 | Adult | eptifibatide | .Standard Dose | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 2.2 | 2.5 | 2 | [10 empty] | n/a |  |  |  | SCr<2.0
53 | Adult | eptifibatide | Renal Dose | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mcg/kg/min | 0.25 | 1.2 | 1.4 | 1 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
54 | Adult | eptifibatide | Renal Dose | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.25 | 1.2 | 1.4 | 1 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
55 | Adult | eptifibatide | Renal Dose wt >121kg | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mg/h | 6 | 8 | 9 | 7.5 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
56 | Adult | eptifibatide | Renal Dose wt >121kg | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mg/h | 6 | 8 | 9 | 7.5 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
57 | Adult | esmolol |  | 2500 | mg | 250 | mL | 10 | mg/mL | Pump | mcg/kg/min | 25 | 300 |  | 50 | mcg/kg | 100 | 500 |  |  | mcg/kg/min | 100 | 500 |  |  | n/a
58 | Adult | fenoldopam |  | 10 | mg | 250 | mL | 0.04 | mg/mL | Pump | mcg/kg/min | 0.03 | 1.6 | 1.8 | 0.1 | [10 empty] | n/a
59 | Adult | fentaNYL |  | 2500 | mcg | 100 | mL | 25 | mcg/mL | Pump | mcg/h | 0.0001 | 250 | 750 |  | mcg | 0.0001 | 250 | 500 |  | mcg/min | 0.0001 | 250 | 500 |  | n/a
60 | Adult | furosemide |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mg/h | 1 | 40 | [12 empty] | n/a |  |  |  | Protect From Light
61 | Adult | glucagon |  | 10 | mg | 50 | mL | 0.2 | mg/mL | Pump | mg/h | 0.1 | 10 | [12 empty] | n/a
62 | Adult | glucagon |  | 10 | mg | 100 | mL | 0.1 | mg/mL | Pump | mg/h | 0.1 | 10 | [12 empty] | n/a
63 | Adult | heparin | ..Std. Cardiac/PVD | 25000 | unit | 500 | mL | 50 | unit/mL | Pump | unit/kg/h | 1 | 100 |  | 12 | unit/kg | 24 |  | 62 |  | unit/kg/min | 12 |  | 31 |  | n/a |  |  |  | Heparin
64 | Adult | heparin | .VTE (DVT, PE) | 25000 | unit | 500 | mL | 50 | unit/mL | Pump | unit/kg/h | 1 | 100 |  | 18 | unit/kg | 24 |  | 82 |  | unit/kg/min | 12 |  | 41 |  | n/a |  |  |  | Heparin
65 | Adult | heparin | Cerebrovascular | 25000 | unit | 500 | mL | 50 | unit/mL | Pump | unit/kg/h | 1 | 100 |  | 10 | [10 empty] | n/a |  |  |  | Heparin
66 | Adult | insulin |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 1 | 20 | 30 |  | unit | 1 | 11 |  |  | unit/min | 0.5 | 15 |  | 10 | n/a |  |  |  | 2nd RN to verify
67 | Adult | insulin |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/kg/h | 0.001 | 0.3 | 0.8 |  | unit/kg | 0.001 | 0.3 |  | 0.1 | unit/min | 0.5 | 15 |  | 10 | n/a |  |  |  | 2nd RN to verify
68 | Adult | isoproterenol |  | 1 | mg | 250 | mL | 0.004 | mg/mL | Pump | mcg/min | 0.15 | 4 | 30 | 1 | [10 empty] | n/a
69 | Adult | isoproterenol |  | 2 | mg | 500 | mL | 0.004 | mg/mL | Pump | mcg/min | 0.15 | 4 | 30 | 1 | [10 empty] | n/a
70 | Adult | ketamine ANES |  | - - | mg | - - | mL |  | mg/mL | Pump | mg/kg/h | 0.05 | 0.5 | 1 | 0.1 | [10 empty] | mg/mL | 0.1 | 0.5 | 1.5
71 | Adult | labetalol |  | 200 | mg | 200 | mL | 1 | mg/mL | Pump | mg/min | 0.05 | 2 |  | 2 | [10 empty] | n/a
72 | Adult | lidocaine |  | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | mg/min | 0.5 | 4 | 4.6 | 2 | [10 empty] | n/a
73 | Adult | mag OB maintenance |  | 20 | gram | 500 | mL | 0.04 | gram/mL | Pump | gram/h | 0.5 | 3 | 5 | [11 empty] | n/a
74 | Adult | midazolam | Anti-Convulsant | 100 | mg | 100 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.0001 | 2 |  | 0.1 | mg/kg | 0.0001 | 0.21 |  | 0.2 | mg/min | 1.9 | 2.1 |  | 2 | n/a
75 | Adult | midazolam | Anti-Convulsant | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.0001 | 2 |  | 0.1 | mg/kg | 0.0001 | 0.21 |  | 0.2 | mg/min | 1.9 | 2.1 |  | 2 | n/a
76 | Adult | midazolam | Sedation | 100 | mg | 100 | mL | 1 | mg/mL | Pump | mg/h | 0.0001 | 5 | 10 |  | mg | 0.0001 | 1 | 5 |  | mg/min | 0.0001 | 1 | 5 |  | n/a
77 | Adult | milrinone |  | 40 | mg | 200 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 0.19 | 0.75 |  | 0.5 | mcg/kg | 25 | 50.1 |  | 50 | mcg/kg/min | 1 | 5.01 |  | 5 | n/a
78 | Adult | naloxone |  | 4 | mg | 100 | mL | 0.04 | mg/mL | Pump | mg/h | 0.25 | 2 | [12 empty] | n/a
79 | Adult | nesiritide |  | 1.5 | mg | 250 | mL | 0.006 | mg/mL | Pump | mcg/kg/min | 0.005 | 0.03 | 0.035 | 0.01 | [10 empty] | n/a
80 | Adult | niCARdipine |  | 20 | mg | 200 | mL | 0.1 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
81 | Adult | niCARdipine |  | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
82 | Adult | niCARdipine |  | 40 | mg | 200 | mL | 0.2 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
83 | Adult | niCARdipine |  | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
84 | Adult | niCARdipine |  | 50 | mg | 500 | mL | 0.1 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
85 | Adult | nitroglycerin | ..Standard | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mcg/min | 5 | 200 | 400 | 5 | [10 empty] | n/a
86 | Adult | nitroglycerin | .Double Concentrated | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mcg/min | 5 | 200 | 400 | 5 | [10 empty] | n/a
87 | Adult | nitroprusside |  | 50 | mg | 500 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.1 |  | 10 | 0.3 | [10 empty] | n/a |  |  |  | Protect From Light
88 | Adult | norepinephrine | ..Standard | 4 | mg | 250 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Levophed
89 | Adult | norepinephrine | ..Standard | 8 | mg | 500 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Levophed
90 | Adult | norepinephrine | .Double Concentrated | 16 | mg | 500 | mL | 0.032 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Double Concentrated
91 | Adult | octreotide |  | 600 | mcg | 240 | mL | 2.5 | mcg/mL | Pump | mcg/h | 25 | 50 |  |  | mcg | 25 | 150 |  | 50 | mcg/min | 1 | 50 |  | 10 | n/a
92 | Adult | oxytocin |  | 30 | unit | 500 | mL | 0.06 | unit/mL | Pump | milliunit/min | 1 | 41 |  |  | unit | 0.9 | 30.1 |  | 10 | unit/min | 0.1 | 1.01 |  | 0.999 | n/a
93 | Adult | pancuronium |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 0.01 | 2 | 3 |  | mcg/kg | 0.06 | 0.1 |  |  | mcg/kg/min | 0.06 | 1 |  |  | mcg/mL | 99 | 100 | 2001 | Neuromuscular BA
94 | Adult | pantoprazole drip |  | 40 | mg | 100 | mL | 0.4 | mg/mL | Pump | mg/h | 7 | 9 | 11 | 8 | [10 empty] | n/a
95 | Adult | pantoprazole drip |  | 96 | mg | 240 | mL | 0.4 | mg/mL | Pump | mg/h | 7 | 9 | 11 | 8 | [10 empty] | n/a
96 | Adult | phenylephrine |  | 10 | mg | 250 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Phenylephrine
97 | Adult | phenylephrine |  | 20 | mg | 500 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Phenylephrine
98 | Adult | phenylephrine |  | 40 | mg | 1000 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Phenylephrine
99 | Adult | phenylephrine ANES |  | 10 | mg | 101 | mL | 0.099 | mg/mL | Pump | mcg/min | 5 | 120 | 180 | [11 empty] | n/a
100 | Adult | phenylephrine ANES |  | 10 | mg | 251 | mL | 0.0398 | mg/mL | Pump | mcg/min | 5 | 120 | 180 | [11 empty] | n/a
101 | Adult | procainamide | Loading Dose | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | mg/min | 15 | 50 |  | 33.33 | [10 empty] | n/a
102 | Adult | procainamide | Maintenance Dose | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | mg/kg/min | 0.01 | 0.06 |  | 0.05 | [10 empty] | n/a
103 | Adult | propofol |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | mcg/kg/min | 1 | 55 | 100 | 5 | [10 empty] | n/a
104 | Adult | rocuronium |  | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 4 | 12 | 16 | 8 | [10 empty] | n/a |  |  |  | Neuromuscular BA
105 | Adult | tirofiban | CrCl <= 60 | 12.5 | mg | 250 | mL | 0.05 | mg/mL | Pump | mcg/kg/min | 0.065 |  | 0.085 | 0.075 | mcg/kg | 22 |  | 28 | 25 | mcg/kg/min | 4 |  | 6 | 5 | n/a |  |  |  | Bleeding Precautions
106 | Adult | tirofiban | CrCl > 60 | 12.5 | mg | 250 | mL | 0.05 | mg/mL | Pump | mcg/kg/min | 0.13 |  | 0.17 | 0.15 | mcg/kg | 22 |  | 28 | 25 | mcg/kg/min | 4 |  | 6 | 5 | n/a |  |  |  | Bleeding Precautions
107 | Adult | vasopressin | ..Standard | 40 | unit | 100 | mL | 0.4 | unit/mL | Pump | unit/min | 0.01 | 0.05 | 0.06 | 0.04 | [10 empty] | n/a
108 | Adult | vasopressin | .Double Concentrated | 80 | unit | 100 | mL | 0.8 | unit/mL | Pump | unit/min | 0.01 | 0.05 | 0.06 | 0.04 | [10 empty] | n/a |  |  |  | Double Concentrated
109 | Adult | vecuronium |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | mcg/kg/min | 0.01 | 1.3 | 1.67 | 1 | mg | 5 | 10 |  |  | mg/min | 5 | 20 |  |  | n/a |  |  |  | Neuromuscular BA
110 | Adult | zz amiodarone DRIP |  | 360 | mg | 200 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.41 | [11 empty] | n/a |  |  |  | Code only drugs
111 | Adult | zz amiodarone DRIP |  | 900 | mg | 500 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.41 | [11 empty] | n/a |  |  |  | Code only drugs
112 | Adult | zz DOBUTamine | ..Standard | 500 | mg | 250 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a |  |  |  | Code only drugs
113 | Adult | zz DOBUTamine | .Double Concentrated | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a |  |  |  | Code only drugs
114 | Adult | zz DOPamine | ..Standard | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 0.9 | 20 | 30 | [11 empty] | n/a |  |  |  | Code only drugs
115 | Adult | zz DOPamine | .Double Concentrated | 800 | mg | 250 | mL | 3.2 | mg/mL | Pump | mcg/kg/min | 0.9 | 20 | 30 | [11 empty] | n/a |  |  |  | Code only drugs
116 | Adult | zz EPINEPHrine |  | 5 | mg | 250 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 40 | [12 empty] | n/a |  |  |  | Code only drugs
117 | Adult | zz EPINEPHrine |  | 10 | mg | 500 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 40 | [12 empty] | n/a |  |  |  | Code only drugs
118 | Adult | zz lidocaine |  | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | mg/min | 0.5 | 4 | 4.6 | 2 | [10 empty] | n/a |  |  |  | Code only drugs
119 | Adult | zz nitroglycerin | ..Standard | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mcg/min | 5 | 200 | 400 | 5 | [10 empty] | n/a |  |  |  | Code only drugs
120 | Adult | zz nitroglycerin | .Double Concentrated | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mcg/min | 5 | 200 | 400 | 5 | [10 empty] | n/a
121 | Adult | zz nitroprusside |  | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 0.1 |  | 10 | 0.3 | [10 empty] | n/a |  |  |  | Code only drugs
122 | Adult | zz norepinephrine | ..Standard | 4 | mg | 250 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Code only drugs
123 | Adult | zz norepinephrine | ..Standard | 8 | mg | 500 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Code only drugs
124 | Adult | zz norepinephrine | .Double Concentrated | 16 | mg | 500 | mL | 0.032 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Code only drugs
125 | Adult | zz phenylephrine |  | 10 | mg | 250 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Code only drugs
126 | Adult | zz phenylephrine |  | 20 | mg | 500 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Code only drugs
127 | Adult | zz phenylephrine |  | 40 | mg | 1000 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Code only drugs
128 | Nursery | No drugs exist in this section
129 | ONC Infusion Center | No drugs exist in this section
130 | OutPt InfusionCenter | No drugs exist in this section
131 | Pediatrics | alprostadil |  | 500 | mcg | 250 | mL | 2 | mcg/mL | Pump | mcg/kg/min | 0.05 | 0.1 | 0.4 | 0.05 | [10 empty] | n/a
132 | Pediatrics | alprostadil |  | - - | mcg | - - | mL |  | mcg/mL | Pump | mcg/kg/min | 0.05 | 0.1 | 0.4 | 0.05 | [10 empty] | mcg/mL | 0.5 | 1 | 20
133 | Pediatrics | aminophylline |  | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.2 | 1.1 |  |  | mg/kg | 0.001 |  | 6 |  | mg/kg/min | 0.19 |  | 0.36 | 0.2 | n/a
134 | Pediatrics | DOBUTamine |  | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | 2 | [10 empty] | n/a
135 | Pediatrics | DOPamine |  | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 2 |  | 20 | 2 | [10 empty] | n/a
136 | Pediatrics | EPINEPHrine |  | 5 | mg | 250 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 20 | 30 | [11 empty] | n/a
137 | Pediatrics | insulin |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 1 | 20 | 30 |  | unit | 1 | 11 |  |  | unit/min | 0.5 | 15 |  | 10 | n/a |  |  |  | 2nd RN to verify
138 | Pediatrics | insulin |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/kg/h | 0.001 | 0.3 | 0.8 |  | unit/kg | 0.001 | 0.3 |  | 0.1 | unit/min | 0.5 | 15 |  | 10 | n/a |  |  |  | 2nd RN to verify
139 | Pediatrics | lidocaine |  | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | mg/min | 0.5 | 4 | 4.6 | 2 | [10 empty] | n/a
140 | Pediatrics | milrinone |  | 20 | mg | 100 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 0.125 | 0.75 |  | 0.5 | mcg/kg | 25 | 50 |  |  | mcg/kg/min | 1 | 5 |  |  | n/a
141 | Pediatrics | naloxone |  | 2 | mg | 500 | mL | 0.004 | mg/mL | Pump | mg/h | 0.25 | 2 | [12 empty] | n/a
142 | Pediatrics | niCARdipine |  | 20 | mg | 200 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.01 | 4 | 5 | 0.5 | mcg/kg | 29 |  | 50 | 30 | mcg/kg/min | 29 |  | 50 | 30 | n/a
143 | Pediatrics | niCARdipine |  | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.01 | 4 | 5 | 0.5 | mcg/kg | 29 |  | 50 | 30 | mcg/kg/min | 29 |  | 50 | 30 | n/a
144 | Pediatrics | niCARdipine |  | 40 | mg | 200 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 0.01 | 4 | 5 | 0.5 | mcg/kg | 29 |  | 50 | 30 | mcg/kg/min | 29 |  | 50 | 30 | n/a
145 | Pediatrics | niCARdipine |  | 50 | mg | 500 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.01 | 4 | 5 | 0.5 | mcg/kg | 29 |  | 50 | 30 | mcg/kg/min | 29 |  | 50 | 30 | n/a
146 | Pediatrics | nitroglycerin |  | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mcg/min | 5 | 150 | 400 | 5 | [10 empty] | n/a
147 | Pediatrics | nitroprusside |  | 50 | mg | 500 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.1 | 8 | 8.1 | 1 | [10 empty] | n/a |  |  |  | Protect From Light
148 | Pediatrics | norepinephrine |  | 4 | mg | 250 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 45 | [12 empty] | n/a
149 | Pediatrics | norepinephrine |  | 8 | mg | 500 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 45 | [12 empty] | n/a
150 | Pediatrics | propofol |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | mcg/kg/min | 1 | 155 | 310 | [11 empty] | n/a
Note on AE4: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on AE5: Dose and volume to be infused need to be set according to weight based protocol to avoid potential over infusion of entire contents of bag.
Note on AE7: Infusion requires the use of an in-line filter.
Note on AE8: Infusion requires the use of an in-line filter.
Note on AE9: Infusion requires the use of an in-line filter.
Note on J16: Units Only concentration and Dosing Units have different units of measure
Note on AA16: Units Only concentration and Dosing Units have different units of measure
Note on AE16: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE25: Protect this drug from light.
Note on AE27: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE36: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on AE37: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on J38: Units Only concentration and Dosing Units have different units of measure
Note on AA38: Units Only concentration and Dosing Units have different units of measure
Note on AE41: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on AE45: 1 unit = 1 mL
Note on AE46: 1 unit = 1 mL
Note on AE47: Do not infuse any other drugs through this line.
Note on AE48: Do not infuse any other drugs through this line.
Note on AE49: Do not infuse any other drugs through this line.
Note on AE51: SCr < 2.0
Note on AE52: SCr < 2.0
Note on AE53: SCr < 2.0
Note on AE54: SCr < 2.0
Note on AE55: SCr 2.0 - 4.0
Note on AE56: SCr 2.0 - 4.0
Note on AE57: SCr 2.0 - 4.0
Note on AE58: SCr 2.0 - 4.0
Note on AE62: Protect this drug from light.
Note on AE65: Monitor PTT. Do not administer other drugs through same IV line.
Note on AE66: Monitor PTT. Do not administer other drugs through same IV line.
Note on AE67: Monitor PTT. Do not administer other drugs through same IV line.
Note on AE68: 2 RNs to verify and document concentration and programming prior to infusion.
Note on AE69: 2 RNs to verify and document concentration and programming prior to infusion.
Note on AE89: Protect this drug from light.
Note on AE90: *****Must have physician order for rates greater than 60 mcg/min*****
Note on AE91: *****Must have physician order for rates greater than 60 mcg/min*****
Note on AE92: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on J95: Units Only concentration and Dosing Units have different units of measure
Note on AA95: Units Only concentration and Dosing Units have different units of measure
Note on AE95: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE98: *****Must have physician order for rates greater than 300 mcg/min*****
Note on AE99: *****Must have physician order for rates greater than 300 mcg/min*****
Note on AE100: *****Must have physician order for rates greater than 300 mcg/min*****
Note on AE106: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE107: Monitor for signs of bleeding.
Note on AE108: Monitor for signs of bleeding.
Note on AE110: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on AE111: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE112: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE113: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE114: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE115: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE116: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE117: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE118: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE119: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE120: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE121: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE123: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE124: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE125: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE126: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE127: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE128: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE129: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE139: 2 RNs to verify and document concentration and programming prior to infusion.
Note on AE140: 2 RNs to verify and document concentration and programming prior to infusion.
Note on AE149: Protect this drug from light.

## ContBolus Anesth Drugs
Continuous/Bolus Anesthesia Only Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount (- - = custom concentration) | Drug Amount Units | Diluent Volume (- - = custom concentration) | Diluent Volume Units | Final Concentration | Final Concentration Units | Module Availability | Dosing Units | Dose Limits Soft Min | Dose Limits Soft Max | Dose Limits Hard Max | Dose Initial Value | Bolus Dose Units | Bolus Dose Limits Soft Min | Bolus Dose Limits Soft Max | Bolus Dose Limits Hard Max | Bolus Dose Limits Initial Value | Bolus Dose Admin. Rate Units | Bolus Dose Admin. Rate Limits Soft Min | Bolus Dose Admin. Rate Limits Soft Max | Bolus Dose Admin. Rate Limits Hard Max | Bolus Dose Admin. Rate Initial Value | Conc. Limit Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max
1 | Adult | ketamine ANES |  | 50 | mg | 50 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.05 | 0.5 | 1 | 0.1 | [10 empty] | n/a
2 | Adult | ketamine ANES |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.05 | 0.5 | 1 | 0.1 | [10 empty] | n/a
3 | Adult | phenylephrine ANES |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/min | 5 | 120 | 180 | [11 empty] | mcg/mL | 10 | 38 | 110
4 | Adult | propofol ANES |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | mcg/kg/min | 1 |  | 300 | 5 | mg | 10 | 70 |  | 30 | mg/min | 50 | 140 |  | 120 | n/a
5 | Adult | propofol ANES |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | mcg/kg/min | 1 |  | 300 | 5 | mg | 10 | 70 |  | 30 | mg/min | 50 | 140 |  | 120 | n/a
6 | Adult | remifentanil ANES |  | 1000 | mcg | 40 | mL | 25 | mcg/mL | Pump | mcg/kg/min | 0.025 | 2 | 4 | 0.05 | mcg/kg | 0.1 |  | 2 | 1 | mcg/min | 1 | 2000 | 4000 | 1000 | n/a
7 | Adult | remifentanil ANES |  | 2000 | mcg | 80 | mL | 25 | mcg/mL | Pump | mcg/kg/min | 0.025 | 2 | 4 | 0.05 | mcg/kg | 0.1 |  | 2 | 1 | mcg/min | 1 | 2000 | 4000 | 1000 | n/a
8 | Adult | remifentanil ANES |  | - - | mcg | - - | mL |  | mcg/mL | Pump | mcg/kg/min | 0.025 | 2 | 4 | 0.05 | mcg/kg | 0.1 |  | 2 | 1 | mcg/min | 1 | 2000 | 4000 | 1000 | mcg/mL | 8.9 | 9 | 250
9 | Adult | rocuronium ANES |  | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 4 | 12 | 16 | 8 | [10 empty] | n/a
10 | Adult | rocuronium ANES |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 4 | 12 | 16 | 8 | [10 empty] | mcg/mL | 890 | 900 | 1100
11 | Adult | SUFentanil ANES |  | 100 | mcg | 100 | mL | 1 | mcg/mL | Pump | mcg/kg/h | 0.1 | 1 | [12 empty] | n/a
12 | Adult | SUFentanil ANES |  | - - | mcg | - - | mL |  | mcg/mL | Pump | mcg/kg/h | 0.1 | 1 | [12 empty] | mcg/mL | 0.1 | 0.2 | 4
13 | Adult | vecuronium |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 0.01 | 1.3 | 1.67 | 1 | mg | 5 | 10 |  |  | mg/min | 5 | 20 |  |  | mcg/mL | 440 | 450 | 550
14 | Nursery | No drugs exist in this section
15 | ONC Infusion Center | No drugs exist in this section
16 | OutPt InfusionCenter | No drugs exist in this section
17 | Pediatrics | No drugs exist in this section
Note on J5: Units Only concentration and Dosing Units have different units of measure
Note on AA5: Units Only concentration and Dosing Units have different units of measure
Note on J12: Units Only concentration and Dosing Units have different units of measure
Note on AA12: Units Only concentration and Dosing Units have different units of measure
Note on J15: Units Only concentration and Dosing Units have different units of measure
Note on AA15: Units Only concentration and Dosing Units have different units of measure

## Intermittent Drugs
Intermittent Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount (- - = custom concentration) | Drug Amount Units | Diluent Volume (- - = custom concentration) | Diluent Volume Units | Final Concentration | Final Concentration Units | Module Availability | Available As | Duration Limits (hh:mm) Hard Min | Duration Limits (hh:mm) Soft Min | Duration Limits (hh:mm) Soft Max | Duration (hh:mm) Initial Value | Dosing Units | Dose Limits Soft Min | Dose Limits Soft Max | Dose Limits Hard Max | Conc. Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max | Clinical Advisory
1 | Adult | acetaminophen | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 500 | 1000 |  | n/a
2 | Adult | acetaminophen | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | gram | 0.5 | 1 |  | n/a
3 | Adult | acetaZOLAMIDE |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 00:40 | 00:30 | mg | 100 | 1000 |  | n/a
4 | Adult | acetylcysteine | 1st Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg/kg | 142.5 |  | 157.5 | mg/mL | 1 |  | 75
5 | Adult | acetylcysteine | 1st Dose >100kg | 15000 | mg | 275 | mL | 54.5455 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 13500 | 16500 | 18000 | n/a
6 | Adult | acetylcysteine | 2nd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg/kg | 47.5 |  | 52.5 | mg/mL | 1 |  | 10
7 | Adult | acetylcysteine | 2nd Dose >100kg | 5000 | mg | 525 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg | 4000 | 5500 | 6500 | n/a
8 | Adult | acetylcysteine | 3rd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg/kg | 95 |  | 105 | mg/mL | 1 |  | 10
9 | Adult | acetylcysteine | 3rd Dose >100kg | 10000 | mg | 1050 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg | 9000 | 11000 | 12500 | n/a
10 | Adult | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 100 | 2000 |  | mg/mL | 0.1 | 1 | 7
11 | Adult | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg/kg | 1 | 10.1 |  | mg/mL | 0.1 | 1 | 7
12 | Adult | albumin 25% |  | 5 | gram | 20 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:08 | 01:00 | 00:20 | gram | 4.9 | 25.1 |  | n/a
13 | Adult | albumin 25% |  | 12.5 | gram | 50 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:16 | 02:00 | 00:50 | gram | 4.9 | 25.1 |  | n/a
14 | Adult | albumin 25% |  | 25 | gram | 100 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:32 | 04:00 | 01:40 | gram | 4.9 | 25.1 |  | n/a
15 | Adult | alteplase-non-stroke | Central IV Occlusion | 2 | mg | 50 | mL | 0.04 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:00 | 00:15 | mg | 2 | 4 |  | n/a
16 | Adult | alteplase-non-stroke | Central IV Occlusion | 2 | mg | 100 | mL | 0.02 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:00 | 00:15 | mg | 2 | 4 |  | n/a
17 | Adult | alteplase-non-stroke | DVT / Ischemic Limb | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:00 |  | 24:00 |  | mg | 0.9 |  | 24 | mg/mL | 0.004 | 0.008 | 0.51
18 | Adult | alteplase-non-stroke | Pulmonary Embolism | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary Only |  | 01:45 | 02:15 | 02:00 | mg | 49 | 115 | 130 | n/a
19 | Adult | amikacin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 |  | mg/kg | 1 | 21 | 39 | mg/mL | 0.2 |  | 5.1
20 | Adult | aminocaproic acid | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 500 | 5000 |  | mg/mL | 1 | 5 | 50
21 | Adult | aminocaproic acid | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | gram | 0.5 | 5.7 |  | gram/mL | 0.001 | 0.005 | 0.05
22 | Adult | amiodarone *Loading* |  | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:08 | 00:15 | 00:10 | mg | 149 | 151 |  | n/a
23 | Adult | amiodarone *Loading* |  | 150 | mg | 103 | mL | 1.4563 | mg/mL | Pump | Primary and Secondary |  | 00:08 | 00:15 | 00:10 | mg | 149 | 151 |  | n/a
24 | Adult | ampho B (liposomal) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only |  | 00:55 | 04:00 | 02:00 | mg/kg | 1.9 | 10.5 |  | mg/mL | 0.8 | 0.9 | 2.1 | AmBisome
25 | Adult | amphotericin B |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 02:00 | 03:00 | 10:00 | 06:00 | mg/kg | 0.05 | 1 | 1.5 | mg/mL | 0.01 |  | 0.1
26 | Adult | ampicillin | .MG Dosing | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
27 | Adult | ampicillin | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
28 | Adult | ampicillin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
29 | Adult | ampicillin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
30 | Adult | ampicillin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
31 | Adult | ampicillin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
32 | Adult | ampicillin | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.1 | 2.4 | n/a
33 | Adult | ampicillin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.1 | 2.4 | n/a
34 | Adult | ampicillin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.1 | 2.4 | n/a
35 | Adult | ampicillin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.1 | 2.4 | n/a
36 | Adult | ampicillin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.1 | 2.4 | n/a
37 | Adult | ampicillin/sulbact | .MG Dosing | 1500 | mg | 50 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | mg | 1200 | 3400 |  | n/a
38 | Adult | ampicillin/sulbact | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | mg | 1200 | 3400 |  | n/a
39 | Adult | ampicillin/sulbact | .MG Dosing | 3000 | mg | 100 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 04:00 | 01:00 | mg | 1200 | 3400 |  | n/a
40 | Adult | ampicillin/sulbact | GM dosing | 1.5 | gram | 50 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | gram | 1.2 | 3.4 |  | n/a
41 | Adult | ampicillin/sulbact | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | gram | 1.2 | 3.4 |  | n/a
42 | Adult | ampicillin/sulbact | GM dosing | 3 | gram | 100 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 04:00 | 01:00 | gram | 1.2 | 3.4 |  | n/a
43 | Adult | andulafungin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 04:00 | 02:00 | mg | 50 | 200 |  | mg/mL | 0.1 | 0.35 | 0.5
44 | Adult | antivenin crotalidae | Anavip | 4 | unit | 250 | mL | 0.016 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 10.1 | 11 | n/a |  |  |  | Antivenin
45 | Adult | antivenin crotalidae | Anavip | 10 | unit | 250 | mL | 0.04 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 10.1 | 11 | n/a |  |  |  | Antivenin
46 | Adult | antivenin crotalidae | Anavip | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 10.1 | 11 | unit/mL | 0.003 |  | 0.045 | Antivenin
47 | Adult | antivenin crotalidae | Crofab | 4 | unit | 250 | mL | 0.016 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 6.1 | 13 | n/a |  |  |  | Antivenin
48 | Adult | antivenin crotalidae | Crofab | 6 | unit | 250 | mL | 0.024 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 6.1 | 13 | n/a |  |  |  | Antivenin
49 | Adult | antivenin crotalidae | Crofab | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 6.1 | 13 | unit/mL | 0.004 |  | 0.025 | Antivenin
50 | Adult | avibactam/ceftazidim |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 01:42 | 03:20 | 02:00 | gram | 0.93 |  | 2.8 | gram/mL | 0.009 |  | 0.051
51 | Adult | azithromycin |  | 250 | mg | 250 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:10 | 01:00 | mg | 210 | 575 |  | n/a
52 | Adult | azithromycin |  | 500 | mg | 250 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:10 | 01:00 | mg | 210 | 575 |  | n/a
53 | Adult | aztreonam | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
54 | Adult | aztreonam | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
55 | Adult | aztreonam | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg | 250 | 2000 |  | n/a
56 | Adult | aztreonam | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg | 250 | 2000 |  | n/a
57 | Adult | aztreonam | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.25 | 2 |  | n/a
58 | Adult | aztreonam | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.25 | 2 |  | n/a
59 | Adult | aztreonam | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | gram | 0.25 | 2 |  | n/a
60 | Adult | aztreonam | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 02:00 | gram | 0.25 | 2 |  | n/a
61 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 110 | mL | 6.3636 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:20 | 06:00 | 00:22 | mg | 699 | 701 |  | n/a
62 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 160 | mL | 4.375 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 06:00 | 00:31 | mg | 699 | 701 |  | n/a
63 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 210 | mL | 3.3333 | mg/mL | Pump | Primary and Secondary | 00:36 | 00:37 | 06:00 | 00:41 | mg | 699 | 701 |  | n/a
64 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 310 | mL | 2.2581 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 06:00 | 01:00 | mg | 699 | 701 |  | n/a
65 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 110 | mL | 6.3636 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:20 | 06:00 | 00:22 | mg | 699 | 701 |  | n/a
66 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 160 | mL | 4.375 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 06:00 | 00:31 | mg | 699 | 701 |  | n/a
67 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 210 | mL | 3.3333 | mg/mL | Pump | Primary and Secondary | 00:36 | 00:37 | 06:00 | 00:41 | mg | 699 | 701 |  | n/a
68 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 310 | mL | 2.2581 | mg/mL | Pump | Primary and Secondary | 00:59 | 01:00 | 06:00 | 01:10 | mg | 699 | 701 |  | n/a
69 | Adult | bamlanivimab |  | 700 | mg | 270 | mL | 2.5926 | mg/mL | Pump | Primary and Secondary | 01:00 |  | 10:00 | 01:00 | mg | 650 | 750 |  | n/a |  |  |  | Low Binding Filter
70 | Adult | caffeine citrate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 06:00 |  | mg | 1 | 1200 |  | mg/mL |  | 0.01 | 10
71 | Adult | caffeine/sod benzoa |  | 500 | mg | 1000 | mL | 0.5 | mg/mL | Pump | Primary and Secondary | 01:01 |  | 08:00 | 01:02 | mg | 50 | 501 | 1001 | n/a |  |  |  | 0.22 Filter Needed
72 | Adult | caffeine/sod benzoa |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 04:00 | 01:00 | mg | 50 | 501 | 1001 | mg/mL |  | 0.24 | 6 | 0.22 Filter Needed
73 | Adult | calcium chloride | .MG Dosing | 1000 | mg | 60 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | mg | 100 | 2000 |  | n/a
74 | Adult | calcium chloride | .MG Dosing | 1000 | mg | 110 | mL | 9.0909 | mg/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | mg | 100 | 2000 |  | n/a
75 | Adult | calcium chloride | .MG Dosing | 2000 | mg | 120 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 | 02:00 | mg | 100 | 2000 |  | n/a
76 | Adult | calcium chloride | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 |  | mg | 100 | 2000 |  | mg/mL |  | 0.1 | 20
77 | Adult | calcium chloride | GM dosing | 1 | gram | 60 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | gram | 0.1 | 10 |  | n/a
78 | Adult | calcium chloride | GM dosing | 1 | gram | 110 | mL | 0.0091 | gram/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | gram | 0.1 | 10 |  | n/a
79 | Adult | calcium chloride | GM dosing | 2 | gram | 120 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 | 02:00 | gram | 0.1 | 10 |  | n/a
80 | Adult | calcium chloride | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 |  | gram | 0.1 | 10 |  | gram/mL |  | 0.001 | 0.04
81 | Adult | calcium gluconate | .MG Dosing | 1000 | mg | 60 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | mg | 500 | 2400 |  | n/a
82 | Adult | calcium gluconate | .MG Dosing | 1000 | mg | 110 | mL | 9.0909 | mg/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | mg | 500 | 2400 |  | n/a
83 | Adult | calcium gluconate | .MG Dosing | 2000 | mg | 70 | mL | 28.5714 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | mg | 500 | 2400 |  | n/a
84 | Adult | calcium gluconate | .MG Dosing | 2000 | mg | 120 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | mg | 500 | 2400 |  | n/a
85 | Adult | calcium gluconate | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:25 | 06:00 |  | mg | 500 | 2400 |  | mg/mL | 1 |  | 40
86 | Adult | calcium gluconate | GM dosing | 1 | gram | 60 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
87 | Adult | calcium gluconate | GM dosing | 1 | gram | 110 | mL | 0.0091 | gram/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
88 | Adult | calcium gluconate | GM dosing | 2 | gram | 70 | mL | 0.0286 | gram/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
89 | Adult | calcium gluconate | GM dosing | 2 | gram | 120 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
90 | Adult | calcium gluconate | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 | 00:25 | 06:00 |  | gram | 0.5 | 2.4 |  | gram/mL | 0.001 |  | 0.04
91 | Adult | casirivimab/imdevima |  | 1200 | mg | 260 | mL | 4.6154 | mg/mL | Pump | Primary and Secondary | 00:48 | 00:49 | 04:00 | 01:00 | mg | 1190 | 1201 | 1210 | n/a |  |  |  | Low Binding Filter
92 | Adult | casirivimab/imdevima |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:18 | 00:19 | 04:00 |  | mg | 1190 | 1201 | 1210 | No |  |  |  | Low Binding Filter
93 | Adult | caspofungin |  | 35 | mg | 100 | mL | 0.35 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg | 35 | 90 |  | n/a |  |  |  | Dedicated IV Line
94 | Adult | caspofungin |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg | 35 | 90 |  | n/a |  |  |  | Dedicated IV Line
95 | Adult | caspofungin |  | 70 | mg | 250 | mL | 0.28 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg | 35 | 90 |  | n/a |  |  |  | Dedicated IV Line
96 | Adult | ceFAZolin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 |  | mg | 480 | 3100 |  | n/a
97 | Adult | ceFAZolin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 3100 |  | n/a
98 | Adult | ceFAZolin | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 3100 |  | n/a
99 | Adult | ceFAZolin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 3100 |  | n/a
100 | Adult | ceFAZolin | .MG Dosing | 3000 | mg | 100 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 3100 |  | n/a
101 | Adult | ceFAZolin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 3.1 |  | n/a
102 | Adult | ceFAZolin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 3.1 |  | n/a
103 | Adult | ceFAZolin | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 3.1 |  | n/a
104 | Adult | ceFAZolin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 3.1 |  | n/a
105 | Adult | ceFAZolin | GM dosing | 3 | gram | 100 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 3.1 |  | n/a
106 | Adult | cefepime | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
107 | Adult | cefepime | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
108 | Adult | cefepime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
109 | Adult | cefepime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
110 | Adult | cefepime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
111 | Adult | cefepime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
112 | Adult | cefepime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
113 | Adult | cefepime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
114 | Adult | cefepime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
115 | Adult | cefepime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
116 | Adult | cefiderocol | .MG Dosing | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
117 | Adult | cefiderocol | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
118 | Adult | cefiderocol | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
119 | Adult | cefiderocol | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
120 | Adult | cefiderocol | GM dosing | 0.75 | gram | 100 | mL | 0.0075 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
121 | Adult | cefiderocol | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
122 | Adult | cefiderocol | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
123 | Adult | cefiderocol | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
124 | Adult | cefotaxime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
125 | Adult | cefotaxime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
126 | Adult | cefotaxime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
127 | Adult | cefotaxime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
128 | Adult | cefotaxime | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 40
129 | Adult | cefotaxime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
130 | Adult | cefotaxime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
131 | Adult | cefotaxime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
132 | Adult | cefotaxime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
133 | Adult | cefotaxime | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.04
134 | Adult | cefoTEtan | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
135 | Adult | cefoTEtan | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
136 | Adult | cefoTEtan | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
137 | Adult | cefoTEtan | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
138 | Adult | cefoTEtan | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
139 | Adult | cefoTEtan | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
140 | Adult | cefoTEtan | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
141 | Adult | cefoTEtan | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
142 | Adult | cefOXitin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
143 | Adult | cefOXitin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
144 | Adult | cefOXitin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 20
145 | Adult | cefOXitin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
146 | Adult | cefOXitin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
147 | Adult | cefOXitin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.02
148 | Adult | ceftaroline |  | 200 | mg | 50 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
149 | Adult | ceftaroline |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
150 | Adult | ceftaroline |  | 300 | mg | 50 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
151 | Adult | ceftaroline |  | 300 | mg | 100 | mL | 3 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
152 | Adult | ceftaroline |  | 400 | mg | 50 | mL | 8 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
153 | Adult | ceftaroline |  | 400 | mg | 100 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
154 | Adult | ceftaroline |  | 600 | mg | 50 | mL | 12 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
155 | Adult | ceftaroline |  | 600 | mg | 100 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
156 | Adult | ceftaroline |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:07 | 04:00 | 01:00 | mg | 200 | 600 |  | mg/mL |  | 0.1 | 14
157 | Adult | cefTAZidime | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
158 | Adult | cefTAZidime | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
159 | Adult | cefTAZidime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
160 | Adult | cefTAZidime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
161 | Adult | cefTAZidime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
162 | Adult | cefTAZidime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
163 | Adult | cefTAZidime | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
164 | Adult | cefTAZidime | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
165 | Adult | cefTAZidime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
166 | Adult | cefTAZidime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
167 | Adult | cefTAZidime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
168 | Adult | cefTAZidime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
169 | Adult | ceftolozane/tazo | .MG Dosing | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 149 | 3010 |  | n/a
170 | Adult | ceftolozane/tazo | .MG Dosing | 375 | mg | 100 | mL | 3.75 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 149 | 3010 |  | n/a
171 | Adult | ceftolozane/tazo | .MG Dosing | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 149 | 3010 |  | n/a
172 | Adult | ceftolozane/tazo | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 149 | 3010 |  | n/a
173 | Adult | ceftolozane/tazo | .MG Dosing | 3000 | mg | 100 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 06:00 | 03:00 | mg | 149 | 3010 |  | n/a
174 | Adult | ceftolozane/tazo | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 149 | 3010 |  | mg/mL | 1 | 1.4 | 30.1
175 | Adult | ceftolozane/tazo | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | gram | 0.149 | 3.1 |  | n/a
176 | Adult | ceftolozane/tazo | GM dosing | 3 | gram | 100 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 06:00 | 03:00 | gram | 0.149 | 3.1 |  | n/a
177 | Adult | ceftolozane/tazo | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | gram | 0.149 | 3.1 |  | gram/mL | 0.001 |  | 0.031
178 | Adult | cefTRIAXone | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
179 | Adult | cefTRIAXone | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
180 | Adult | cefTRIAXone | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
181 | Adult | cefTRIAXone | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
182 | Adult | cefTRIAXone | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | mg/mL | 4.9 | 9 | 40
183 | Adult | cefTRIAXone | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
184 | Adult | cefTRIAXone | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
185 | Adult | cefTRIAXone | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
186 | Adult | cefTRIAXone | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
187 | Adult | cefTRIAXone | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | gram/mL | 0.004 | 0.009 | 0.04
188 | Adult | cefuroxime | .MG Dosing | 750 | mg | 50 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
189 | Adult | cefuroxime | .MG Dosing | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
190 | Adult | cefuroxime | .MG Dosing | 1500 | mg | 50 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
191 | Adult | cefuroxime | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
192 | Adult | cefuroxime | GM dosing | 0.75 | gram | 50 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
193 | Adult | cefuroxime | GM dosing | 0.75 | gram | 100 | mL | 0.0075 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
194 | Adult | cefuroxime | GM dosing | 1.5 | gram | 50 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
195 | Adult | cefuroxime | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
196 | Adult | chloramphenicol |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg/kg | 3 | 25 | 30 | mg/mL | 5 | 10 | 110
197 | Adult | chlorothiazide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 1000 | 2300 | mg/mL | 1 | 2 | 20
198 | Adult | chlorproMAZINE |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 02:00 | 00:15 | mg | 1 | 51 |  | mg/mL |  | 0.001 | 1.1 | Patient Lie Down
199 | Adult | ciprofloxacin |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
200 | Adult | ciprofloxacin |  | 400 | mg | 200 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
201 | Adult | clindamycin |  | 300 | mg | 50 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
202 | Adult | clindamycin |  | 300 | mg | 100 | mL | 3 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
203 | Adult | clindamycin |  | 600 | mg | 50 | mL | 12 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
204 | Adult | clindamycin |  | 600 | mg | 100 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
205 | Adult | clindamycin |  | 900 | mg | 50 | mL | 18 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
206 | Adult | clindamycin |  | 900 | mg | 100 | mL | 9 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
207 | Adult | colistimethate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 10 | 300 |  | mg/mL | 0.1 | 0.5 | 5
208 | Adult | colistimethate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 0.3 |  | 5 | mg/mL | 0.1 | 0.5 | 5
209 | Adult | conivaptan Loading |  | 20 | mg | 100 | mL | 0.2 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 15 | 25 |  | n/a
210 | Adult | dalbavancin | .MG Dosing | 375 | mg | 250 | mL | 1.5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
211 | Adult | dalbavancin | .MG Dosing | 500 | mg | 250 | mL | 2 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
212 | Adult | dalbavancin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
213 | Adult | dalbavancin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
214 | Adult | dalbavancin | .MG Dosing | 1125 | mg | 250 | mL | 4.5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
215 | Adult | dalbavancin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
216 | Adult | dalbavancin | GM dosing | 0.375 | gram | 250 | mL | 0.0015 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
217 | Adult | dalbavancin | GM dosing | 0.5 | gram | 250 | mL | 0.002 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
218 | Adult | dalbavancin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
219 | Adult | dalbavancin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
220 | Adult | dalbavancin | GM dosing | 1.125 | gram | 250 | mL | 0.0045 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
221 | Adult | dalbavancin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
222 | Adult | dantrolene |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg/kg | 1 | 2.5 |  | mg/mL | 0.05 | 0.1 | 100
223 | Adult | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg | 100 | 2000 |  | mg/mL | 1 | 2 | 20
224 | Adult | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg/kg | 4 | 12.1 |  | mg/mL | 1 | 2 | 20
225 | Adult | deferoxamine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 04:00 | 08:00 | 24:00 |  | mg | 500 | 6000 |  | mg/mL | 1 |  | 95
226 | Adult | deferoxamine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 04:00 | 08:00 | 24:00 |  | mg/kg | 30 | 60 |  | mg/mL | 1 |  | 95
227 | Adult | desmopressin |  | - - | mcg | - - | mL |  | mcg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mcg | 0.1 | 22 |  | mcg/mL | 0.005 | 0.01 | 4
228 | Adult | desmopressin |  | - - | mcg | - - | mL |  | mcg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mcg/kg | 0.25 | 0.41 |  | mcg/mL | 0.005 | 0.01 | 4
229 | Adult | dexamethasone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:07 | 02:00 | 00:30 | mg | 1 | 20 |  | mg/mL | 0.01 |  | 10
230 | Adult | digoxin immune fab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 10 | 800 |  | mg/mL | 0.05 | 0.1 | 1
231 | Adult | doripenem |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 100 | 500 |  | n/a
232 | Adult | doripenem |  | 250 | mg | 100 | mL | 2.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 100 | 500 |  | n/a
233 | Adult | doripenem |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 100 | 500 |  | n/a
234 | Adult | doxycycline |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
235 | Adult | doxycycline |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
236 | Adult | ertapenem | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
237 | Adult | ertapenem | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
238 | Adult | ertapenem | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
239 | Adult | ertapenem | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
240 | Adult | ertapenem | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
241 | Adult | ertapenem | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
242 | Adult | ertapenem | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
243 | Adult | ertapenem | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
244 | Adult | erythromycin |  | 250 | mg | 100 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
245 | Adult | erythromycin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
246 | Adult | erythromycin |  | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:30 | 01:00 | mg | 100 | 1000 |  | n/a
247 | Adult | erythromycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | mg/mL | 0.05 | 1 | 5
248 | Adult | famotidine |  | 20 | mg | 50 | mL | 0.4 | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg | 8 | 44 | 50 | n/a
249 | Adult | famotidine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 02:00 |  | mg | 8 | 44 | 50 | mg/mL | 0.09 | 0.1 | 3.5
250 | Adult | fat emulsion 20% |  | 25 | gram | 125 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 11:00 | 24:00 |  | gram | 24 | 51 |  | n/a
251 | Adult | fat emulsion 20% |  | 50 | gram | 250 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 11:00 | 24:00 |  | gram | 24 | 51 |  | n/a
252 | Adult | ferric carboxymaltos |  | 750 | mg | 265 | mL | 2.8302 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg | 250 | 750 | 1501 | n/a |  |  |  | Hypersensitivity
253 | Adult | ferric carboxymaltos |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg | 250 | 750 | 1501 | mg/mL | 1.9 | 2 | 4 | Hypersensitivity
254 | Adult | ferric carboxymaltos |  | 750 | mg | 265 | mL | 2.8302 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg/kg | 5 | 16 |  | n/a |  |  |  | Hypersensitivity
255 | Adult | ferric carboxymaltos |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg/kg | 5 | 16 |  | mg/mL | 1.9 | 2 | 4 | Hypersensitivity
256 | Adult | ferumoxytol | .MG Dosing | 510 | mg | 117 | mL | 4.359 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 509 | 1021 | 1022 | n/a |  |  |  | Hypersensitivity
257 | Adult | ferumoxytol | .MG Dosing | 1020 | mg | 284 | mL | 3.5915 | mg/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | mg | 509 | 1021 | 1022 | n/a |  |  |  | Hypersensitivity
258 | Adult | ferumoxytol | GM dosing | 0.51 | gram | 117 | mL | 0.0044 | gram/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | gram | 0.509 | 1.021 | 1.022 | n/a |  |  |  | Hypersensitivity
259 | Adult | ferumoxytol | GM dosing | 1.02 | gram | 284 | mL | 0.0036 | gram/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | gram | 0.509 | 1.021 | 1.022 | n/a |  |  |  | Hypersensitivity
260 | Adult | fluconazole |  | 100 | mg | 50 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 04:00 | 00:30 | mg | 50 | 801 |  | n/a
261 | Adult | fluconazole |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 04:00 | 01:00 | mg | 50 | 801 |  | n/a
262 | Adult | fluconazole |  | 400 | mg | 200 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 01:00 | 04:00 | 02:00 | mg | 50 | 801 |  | n/a
263 | Adult | fluconazole |  | 800 | mg | 400 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 02:00 | 08:00 | 04:00 | mg | 50 | 801 |  | n/a
264 | Adult | folic acid |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 0.1 | 5 |  | mg/mL | 0.001 | 0.009 | 0.06
265 | Adult | fomepizole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 9 | 16 |  | mg/mL | 0.5 | 1 | 20
266 | Adult | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:11 | 01:00 | 00:20 | mg | 50 | 1500 |  | mg/mL | 0.5 | 1 | 30
267 | Adult | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:11 | 01:00 | 00:20 | mg/kg | 1 | 20 |  | mg/mL | 0.5 | 1 | 30
268 | Adult | ganciclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg/kg | 0.625 | 8 |  | mg/mL | 0.5 | 1 | 10 | In-line filter
269 | Adult | gentamicin |  | 80 | mg | 100 | mL | 0.8 | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 | 00:30 | mg | 10 | 1000 |  | n/a
270 | Adult | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 | 01:00 | mg | 10 | 1000 |  | mg/mL | 0.05 | 0.1 | 10
271 | Adult | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 |  | 01:20 | 01:00 | mg/kg | 0.5 | 10 |  | mg/mL | 0.05 | 0.1 | 10
272 | Adult | hetastarch 6% |  | 30 | gram | 500 | mL | 0.06 | gram/mL | Pump | Primary and Secondary |  | 00:31 | 24:00 |  | gram | 29 | 31 |  | n/a
273 | Adult | hydrocortisone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:30 | mg | 100 | 1000 |  | mg/mL | 0.5 | 1 | 20
274 | Adult | ibutilide | pt weight < 60 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:07 |  | 00:30 | 00:10 | mg/kg | 0.009 | 0.011 |  | mg/mL | 0.001 |  | 0.02
275 | Adult | ibutilide | pt weight >= 60 kg | 1 | mg | 60 | mL | 0.0167 | mg/mL | Pump | Primary and Secondary | 00:07 |  | 00:30 | 00:10 | mg | 0.4 | 1.25 |  | n/a
276 | Adult | imipenem/cilastatin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 199 | 1000 |  | n/a
277 | Adult | imipenem/cilastatin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 199 | 1000 |  | n/a
278 | Adult | imipenem/cilastatin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 01:00 | mg | 199 | 1000 |  | mg/mL | 1 | 2 | 5
279 | Adult | imipenem/cilastatin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 0.5 | 1 |  | n/a
280 | Adult | imipenem/cilastatin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | gram | 0.5 | 1 |  | n/a
281 | Adult | imipenem/cilastatin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 01:00 | gram | 0.5 | 1 |  | gram/mL | 0.001 | 0.003 | 0.006
282 | Adult | immune glob (IVIG) |  | 5 | gram | 50 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:00 |  | gram | 1 | 10.1 |  | n/a |  |  |  | Hypersensitivity
283 | Adult | immune glob (IVIG) |  | 10 | gram | 100 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:40 | 08:00 |  | gram | 1 | 10.1 |  | n/a |  |  |  | Hypersensitivity
284 | Adult | immune glob (IVIG) |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 02:00 | 36:00 |  | gram/kg | 1 | 2.4 |  | gram/mL | 0.01 | 0.05 | 0.1 | Hypersensitivity
285 | Adult | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg | 50 | 2000 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
286 | Adult | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg/kg | 2 | 10 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
287 | Adult | iron sucrose |  | 100 | mg | 105 | mL | 0.9524 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 08:00 | 00:30 | mg | 99 |  | 510 | n/a |  |  |  | Hypersensitivity
288 | Adult | iron sucrose |  | 200 | mg | 110 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 08:00 | 00:30 | mg | 99 |  | 510 | n/a |  |  |  | Hypersensitivity
289 | Adult | iron sucrose |  | 300 | mg | 265 | mL | 1.1321 | mg/mL | Pump | Primary and Secondary | 01:25 |  | 08:00 | 01:30 | mg | 99 |  | 510 | n/a |  |  |  | Hypersensitivity
290 | Adult | iron sucrose |  | 400 | mg | 270 | mL | 1.4815 | mg/mL | Pump | Primary and Secondary | 02:25 |  | 08:00 | 02:30 | mg | 99 |  | 510 | n/a |  |  |  | Hypersensitivity
291 | Adult | iron sucrose |  | 500 | mg | 275 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 03:25 |  | 08:00 | 04:00 | mg | 99 |  | 510 | n/a |  |  |  | Hypersensitivity
292 | Adult | itraconazole |  | 200 | mg | 60 | mL | 3.3333 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 100 | 400 |  | n/a |  |  |  | In-line filter
293 | Adult | itraconazole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 100 | 400 |  | mg/mL | 3.29 | 3.3 | 3.4 | In-line filter
294 | Adult | Kcentra (Factor IX) | .INR 2 to < 4 | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | unit/kg | 24 | 26 |  | unit/mL | 19 | 20 | 30 | INR 2-4
295 | Adult | Kcentra (Factor IX) | .INR 4 - 6 | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | unit/kg | 34 | 36 |  | unit/mL | 19 | 20 | 30 | INR 4-6
296 | Adult | Kcentra (Factor IX) | INR > 6 | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | unit/kg | 49 | 53 |  | unit/mL | 19 | 20 | 30 | INR > 6
297 | Adult | Kinevac (sincalide) |  | - - | mcg | - - | mL |  | mcg/mL | Pump | Primary and Secondary |  | 00:12 | 02:00 | 00:16 | mcg/kg | 0.018 | 0.13 |  | mcg/mL | 0.01 |  | 0.5
298 | Adult | lacosamide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 25 | 200 |  | mg/mL | 0.25 | 0.5 | 4.1
299 | Adult | levETIRAcetam |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:06 | 02:00 | 00:15 | mg | 250 | 2001 |  | mg/mL | 2 | 2.5 | 20.1
300 | Adult | levETIRAcetam |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:06 | 02:00 | 00:15 | mg/kg | 1 | 60 |  | mg/mL | 2 | 2.5 | 20.1
301 | Adult | levoFLOXacin |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
302 | Adult | levoFLOXacin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
303 | Adult | levoFLOXacin |  | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 01:15 | 02:00 | 01:30 | mg | 250 | 750 |  | n/a
304 | Adult | linezolid |  | 600 | mg | 300 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 04:00 | 01:00 | mg | 400 | 800 |  | n/a
305 | Adult | mag sulfate |  | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 02:00 | gram | 0.5 | 7 | 8 | n/a
306 | Adult | mag sulfate |  | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:30 | 06:00 | 04:00 | gram | 0.5 | 7 | 8 | n/a
307 | Adult | mag sulfate |  | 4 | gram | 100 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 09:00 |  | gram | 0.5 | 7 | 8 | n/a
308 | Adult | mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:30 | 12:00 |  | gram | 0.5 | 7 | 8 | gram/mL | 0.005 | 0.009 | 0.12
309 | Adult | mannitol 20% |  | 100 | gram | 500 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 00:31 | 04:00 |  | gram | 10 | 100 |  | n/a |  |  |  | In-line filter
310 | Adult | mannitol 20% |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:31 | 08:00 |  | gram | 10 | 100 |  | gram/mL | 0.15 | 0.18 | 0.21 | In-line filter
311 | Adult | mannitol 20% |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:31 | 08:00 |  | gram/kg | 0.1 | 2.1 |  | gram/mL | 0.15 | 0.18 | 0.21 | In-line filter
312 | Adult | meropenem | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
313 | Adult | meropenem | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
314 | Adult | meropenem | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
315 | Adult | meropenem | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
316 | Adult | meropenem | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
317 | Adult | meropenem | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
318 | Adult | meropenem | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
319 | Adult | meropenem | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
320 | Adult | meropenem | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
321 | Adult | meropenem | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
322 | Adult | meropenem/vaborbacta |  | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
323 | Adult | meropenem/vaborbacta |  | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
324 | Adult | meropenem/vaborbacta |  | 2 | gram | 250 | mL | 0.008 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
325 | Adult | meropenem/vaborbacta |  | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
326 | Adult | meropenem/vaborbacta |  | 4 | gram | 250 | mL | 0.016 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
327 | Adult | meropenem/vaborbacta |  | 4 | gram | 500 | mL | 0.008 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
328 | Adult | meropenem/vaborbacta |  | 4 | gram | 1000 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
329 | Adult | methocarbamol |  | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:10 |  | 04:00 |  | mg | 249 |  | 1001 | n/a
330 | Adult | methocarbamol |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 04:00 |  | mg | 249 |  | 1001 | mg/mL | 0.5 | 1 | 15
331 | Adult | methylPREDNISolone |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | n/a
332 | Adult | methylPREDNISolone |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | n/a
333 | Adult | methylPREDNISolone |  | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 125 | 2000 |  | n/a
334 | Adult | methylPREDNISolone |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 125 | 2000 |  | n/a
335 | Adult | methylPREDNISolone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | mg/mL | 0.5 | 1 | 25
336 | Adult | metroNIDAZOLE |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 250 | 1000 |  | n/a
337 | Adult | metroNIDAZOLE |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 250 | 1000 |  | n/a
338 | Adult | micafungin |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
339 | Adult | micafungin |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
340 | Adult | micafungin |  | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
341 | Adult | MVI + Thiame + Folic |  | 10 | unit | 1011 | mL | 0.0099 | unit/mL | Pump | Primary and Secondary | 01:01 |  | 24:00 | 08:06 | unit | 9.9 | 10 |  | n/a
342 | Adult | MVI Only |  | 10 | unit | 510 | mL | 0.0196 | unit/mL | Pump | Primary and Secondary | 00:31 |  | 24:00 | 04:05 | unit | 9.9 | 10.1 |  | n/a
343 | Adult | MVI Only |  | 10 | unit | 1010 | mL | 0.0099 | unit/mL | Pump | Primary and Secondary | 01:01 |  | 24:00 | 08:05 | unit | 9.9 | 10.1 |  | n/a
344 | Adult | nafcillin | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
345 | Adult | nafcillin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
346 | Adult | nafcillin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
347 | Adult | nafcillin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
348 | Adult | nafcillin | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | mg | 200 | 2000 |  | n/a
349 | Adult | nafcillin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | mg | 200 | 2000 |  | n/a
350 | Adult | nafcillin | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
351 | Adult | nafcillin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
352 | Adult | nafcillin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
353 | Adult | nafcillin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
354 | Adult | nafcillin | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | gram | 0.2 | 2 |  | n/a
355 | Adult | nafcillin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | gram | 0.2 | 2 |  | n/a
356 | Adult | ondansetron |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:30 | mg | 2 | 32 |  | mg/mL | 0.02 | 0.04 | 0.64
357 | Adult | pamidronate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 02:00 | 20:00 | 04:00 | mg | 29 | 99 | 110 | mg/mL | 0.001 |  | 0.4 | Dedicated IV Line
358 | Adult | pantoprazole bolus |  | 80 | mg | 100 | mL | 0.8 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:15 | mg | 40 | 80 |  | n/a
359 | Adult | penG 2.5 Mill Units |  | 2.5 | unit | 50 | mL | 0.05 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | unit | 2 | 3 |  | n/a |  |  |  | 1unit=1million
360 | Adult | penG 2.5 Mill Units |  | 2.5 | unit | 100 | mL | 0.025 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | unit | 2 | 3 |  | n/a |  |  |  | 1unit=1million
361 | Adult | penG 3 Million Units |  | 3 | unit | 50 | mL | 0.06 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | unit | 2.8 | 3.2 | 3.3 | n/a |  |  |  | 1unit=1million
362 | Adult | penG 5 Million Units |  | 5 | unit | 50 | mL | 0.1 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | unit | 4.5 | 5.5 |  | n/a |  |  |  | 1unit=1million
363 | Adult | penG 5 Million Units |  | 5 | unit | 100 | mL | 0.05 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | unit | 4.5 | 5.5 |  | n/a |  |  |  | 1unit=1million
364 | Adult | penG x Million Units |  | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | unit | 1 | 6 |  | unit/mL | 0.01 | 0.02 | 0.12 | 1unit=1million
365 | Adult | PHENobarbital |  | 130 | mg | 52 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:04 |  | 04:00 |  | mg | 5 | 270 |  | n/a
366 | Adult | PHENobarbital |  | 130 | mg | 102 | mL | 1.2745 | mg/mL | Pump | Primary and Secondary | 00:07 |  | 04:00 |  | mg | 5 | 270 |  | n/a
367 | Adult | PHENobarbital |  | 260 | mg | 54 | mL | 4.8148 | mg/mL | Pump | Primary and Secondary | 00:04 |  | 04:00 |  | mg | 5 | 270 |  | n/a
368 | Adult | PHENobarbital |  | 260 | mg | 104 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:07 |  | 04:00 |  | mg | 5 | 270 |  | n/a
369 | Adult | PHENobarbital |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:04 |  | 04:00 |  | mg | 5 | 270 |  | No
370 | Adult | phenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:30 | mg/kg | 0.5 | 20 |  | mg/mL | 0.5 | 1 | 10 | In-line filter
371 | Adult | phytonadione |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 01:00 | 00:30 | mg | 1 | 10 |  | mg/mL | 0.005 | 0.01 | 0.2
372 | Adult | piperacillin/tazo |  | 2.25 | gram | 50 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:30 | 00:30 | gram | 3 |  | 4.6 | n/a
373 | Adult | piperacillin/tazo |  | 2.25 | gram | 100 | mL | 0.0225 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:30 | 00:30 | gram | 3 |  | 4.6 | n/a
374 | Adult | piperacillin/tazo |  | 3.375 | gram | 50 | mL | 0.0675 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 | 04:00 | gram | 3 |  | 4.6 | n/a
375 | Adult | piperacillin/tazo |  | 3.375 | gram | 100 | mL | 0.0338 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 | 04:00 | gram | 3 |  | 4.6 | n/a
376 | Adult | piperacillin/tazo |  | 4.5 | gram | 50 | mL | 0.09 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
377 | Adult | piperacillin/tazo |  | 4.5 | gram | 100 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
378 | Adult | potassium acetate |  | 40 | mEq | 520 | mL | 0.0769 | mEq/mL | Pump | Primary and Secondary | 02:00 | 03:30 | 12:00 | 04:00 | mEq | 1 | 46 | 48 | n/a
379 | Adult | potassium acetate |  | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:45 |  | 12:00 |  | mEq | 1 | 46 | 48 | mEq/mL |  | 0.01 | 0.46
380 | Adult | potassium chloride | Central Line | 10 | mEq | 100 | mL | 0.1 | mEq/mL | Pump | Primary and Secondary | 00:14 | 00:45 | 06:00 | 01:00 | mEq | 5 |  | 41 | n/a
381 | Adult | potassium chloride | Central Line | 20 | mEq | 100 | mL | 0.2 | mEq/mL | Pump | Primary and Secondary | 00:25 | 01:30 | 08:00 | 02:00 | mEq | 5 |  | 41 | n/a
382 | Adult | potassium chloride | Central Line | 40 | mEq | 520 | mL | 0.0769 | mEq/mL | Pump | Primary and Secondary | 00:55 | 03:15 | 12:00 | 04:00 | mEq | 5 |  | 41 | n/a
383 | Adult | potassium chloride | Central Line | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:55 | 02:15 | 12:00 |  | mEq | 5 |  | 41 | mEq/mL | 0.01 | 0.02 | 0.21
384 | Adult | potassium chloride | Peripheral Line | 10 | mEq | 100 | mL | 0.1 | mEq/mL | Pump | Primary and Secondary | 00:25 | 00:45 | 06:00 | 01:00 | mEq | 5 |  | 41 | n/a
385 | Adult | potassium chloride | Peripheral Line | 40 | mEq | 520 | mL | 0.0769 | mEq/mL | Pump | Primary and Secondary | 03:15 |  | 12:00 | 04:00 | mEq | 5 |  | 41 | n/a
386 | Adult | potassium chloride | Peripheral Line | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:50 | 01:45 | 12:00 |  | mEq | 5 |  | 41 | mEq/mL | 0.04 | 0.07 | 0.11
387 | Adult | potassium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 01:00 | 02:00 | 09:00 |  | mmol | 15 | 40 |  | mmol/mL | 0.005 | 0.01 | 0.24
388 | Adult | pralidoxime | Load <67kg WtBasedGM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:15 | gram/kg | 0.001 |  | 0.034 | gram/mL |  | 0.009 | 0.051
389 | Adult | pralidoxime | Load <67kg WtBasedMG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:15 | mg/kg | 1 |  | 34 | mg/mL |  | 9 | 51
390 | Adult | pralidoxime | Load >67kg WtBasedGM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:20 | gram | 1 | 21 |  | gram/mL |  | 0.009 | 0.051
391 | Adult | pralidoxime | Load >67kg WtBasedMG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:20 | mg | 500 |  | 2300 | mg/mL |  | 9 | 51
392 | Adult | pralidoxime | Load NonWtBased MG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:20 | mg | 900 |  | 2300 | mg/mL |  | 9 | 51
393 | Adult | pralidoxime | Load NonWtBasedGM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:15 | gram | 0.9 |  | 2.3 | gram/mL |  | 0.009 | 0.051
394 | Adult | pralidoxime | Maintenance Dose GM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 |  | 24:00 |  | gram | 0.9 |  | 2.3 | gram/mL |  | 0.009 | 0.051
395 | Adult | pralidoxime | Maintenance Dose MG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 24:00 |  | mg | 900 |  | 2300 | mg/mL | 5 | 9 | 51
396 | Adult | promethazine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | mg | 6.25 | 25 |  | mg/mL | 0.031 |  | 1
397 | Adult | raNITIdine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg | 25 | 51 |  | mg/mL | 0.1 | 0.5 | 2.5
398 | Adult | remdesivir |  | 100 | mg | 230 | mL | 0.4348 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
399 | Adult | remdesivir |  | 100 | mg | 250 | mL | 0.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
400 | Adult | remdesivir |  | 200 | mg | 210 | mL | 0.9524 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
401 | Adult | remdesivir |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
402 | Adult | rifAMPin |  | 300 | mg | 250 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
403 | Adult | rifAMPin |  | 600 | mg | 500 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
404 | Adult | rifAMPin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | mg/mL | 0.1 | 0.5 | 6
405 | Adult | riTUXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:15 | 02:45 | 25:00 |  | mg | 250 | 1100 |  | mg/mL | 0.1 | 0.2 | 4 | Premedicate
406 | Adult | SMX/TMP |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:00 | 02:00 | 01:30 | mg/kg | 1.5 | 7 |  | mg/mL | 0.05 | 0.1 | 1.2 | 48 Hour Tubing
407 | Adult | sodium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 00:45 |  | 12:00 |  | mmol | 15 | 60 |  | mmol/mL | 0.005 | 0.01 | 0.25
408 | Adult | sotrovimab |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:22 | 00:23 | 04:00 | 00:30 | mg | 490 | 510 | 511 | n/a |  |  |  | Low Binding Filter
409 | Adult | sotrovimab |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:22 | 00:23 | 04:00 | 00:30 | mg | 490 | 510 | 511 | n/a |  |  |  | Low Binding Filter
410 | Adult | tedizolid phosphate |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 01:15 | 01:00 | mg | 175 | 225 | 250 | n/a
411 | Adult | thiamine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:15 | 00:30 | mg | 90 | 550 |  | mg/mL | 0.001 | 0.7 | 10.1
412 | Adult | ticarcillin/clav |  | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
413 | Adult | ticarcillin/clav |  | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
414 | Adult | ticarcillin/clav |  | 3.1 | gram | 50 | mL | 0.062 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
415 | Adult | ticarcillin/clav |  | 3.1 | gram | 100 | mL | 0.031 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
416 | Adult | tigecycline |  | 50 | mg | 50 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:15 | 00:30 | mg | 25 | 100 |  | n/a
417 | Adult | tigecycline |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:15 | 00:30 | mg | 25 | 100 |  | n/a
418 | Adult | tigecycline |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:15 | 00:30 | mg | 25 | 100 |  | n/a
419 | Adult | tobramycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 |  | 01:00 | 00:30 | mg/kg | 0.5 | 9 |  | mg/mL | 0.1 | 0.5 | 10
420 | Adult | tocilizumab |  | 800 | mg | 100 | mL | 8 | mg/mL | Pump | Primary Only | 00:55 |  | 04:00 | 01:00 | mg | 161 |  | 801 | n/a
421 | Adult | tocilizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 00:55 |  | 04:00 | 01:00 | mg/kg | 3.9 |  | 8.1 | mg/mL | 1 | 1.1 | 8.1
422 | Adult | tranexamic BOLUS | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:09 |  | 00:40 | 00:15 | mg | 500 | 2000 |  | n/a
423 | Adult | tranexamic BOLUS | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:07 | 00:20 | 00:40 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
424 | Adult | tranexamic BOLUS | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary | 00:10 |  | 00:40 | 00:15 | gram | 0.5 | 2 |  | n/a
425 | Adult | tranexamic BOLUS | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:07 | 00:20 | 00:40 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
426 | Adult | valproate sodium |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg | 250 | 2000 |  | mg/mL | 0.1 | 1 | 11
427 | Adult | valproate sodium |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg/kg | 1 | 40 |  | mg/mL | 0.1 | 1 | 11
428 | Adult | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 500 | 2000 |  | n/a
429 | Adult | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
430 | Adult | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
431 | Adult | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
432 | Adult | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
433 | Adult | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | mg | 500 | 2000 |  | n/a
434 | Adult | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg | 500 | 2000 |  | n/a
435 | Adult | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg | 500 | 2000 |  | n/a
436 | Adult | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
437 | Adult | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
438 | Adult | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
439 | Adult | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
440 | Adult | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | mg | 500 | 2000 |  | mg/mL | 1 | 2.9 | 5.1
441 | Adult | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg/kg | 5 |  | 30 | n/a
442 | Adult | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
443 | Adult | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
444 | Adult | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
445 | Adult | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
446 | Adult | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
447 | Adult | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
448 | Adult | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
449 | Adult | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
450 | Adult | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
451 | Adult | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
452 | Adult | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
453 | Adult | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | mg/kg | 5 |  | 30 | mg/mL | 1 | 2.9 | 5.1
454 | Adult | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.5 | 2 |  | n/a
455 | Adult | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
456 | Adult | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
457 | Adult | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
458 | Adult | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
459 | Adult | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | gram | 0.5 | 2 |  | n/a
460 | Adult | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram | 0.5 | 2 |  | n/a
461 | Adult | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram | 0.5 | 2 |  | n/a
462 | Adult | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
463 | Adult | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
464 | Adult | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
465 | Adult | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
466 | Adult | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.002 | 0.006
467 | Adult | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram/kg | 0.005 |  | 0.03 | n/a
468 | Adult | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
469 | Adult | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
470 | Adult | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
471 | Adult | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
472 | Adult | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
473 | Adult | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
474 | Adult | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
475 | Adult | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
476 | Adult | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
477 | Adult | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
478 | Adult | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
479 | Adult | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | gram/kg | 0.005 |  | 0.03 | gram/mL | 0.001 | 0.002 | 0.006
480 | Adult | voriconazole |  | 200 | mg | 50 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 02:00 | 01:00 | mg | 200 | 400 |  | n/a
481 | Adult | voriconazole |  | 400 | mg | 100 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 01:40 | 04:00 | 02:00 | mg | 200 | 400 |  | n/a
482 | Adult | voriconazole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 | 01:00 | 04:00 | 02:00 | mg/kg | 2 | 7 |  | mg/mL |  | 0.5 | 5
483 | Adult | zidovudine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 24:00 |  | mg/kg | 1 | 2.2 | 3 | mg/mL | 0.001 | 0.01 | 4.2
484 | Adult | zoledronic acid |  | 4 | mg | 100 | mL | 0.04 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | n/a
485 | Adult | zoledronic acid |  | 5 | mg | 100 | mL | 0.05 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | n/a
486 | Adult | zoledronic acid |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | mg/mL | 0.028 | 0.029 | 0.051
487 | Adult | zz amiodarone BOLUS |  | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:08 | 00:15 | 00:10 | mg | 149 | 151 |  | n/a |  |  |  | Code only drugs
488 | Adult | zz mag sulfate |  | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary | 00:07 | 00:14 | 02:00 | 00:20 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
489 | Adult | zz mag sulfate |  | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary | 00:07 | 00:14 | 02:00 | 00:20 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
490 | Adult | zz mag sulfate |  | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary | 00:14 | 00:28 | 04:00 | 00:40 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
491 | Adult | zz mag sulfate |  | 4 | gram | 100 | mL | 0.04 | gram/mL | Pump | Primary and Secondary | 00:28 | 00:56 | 08:00 | 01:20 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
492 | Adult | zz mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:14 | 00:28 | 08:00 |  | gram | 1 | 6 |  | gram/mL | 0.005 | 0.01 | 0.08 | Code only drugs
493 | Nursery | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:50 | 00:55 | 02:00 | 01:00 | mg/kg | 4 | 21 | 22 | mg/mL | 1 | 4 | 7
494 | Nursery | ampicillin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:07 | 00:08 | 01:10 | 00:30 | mg/kg | 24 | 101 | 105 | mg/mL |  | 10 | 105
495 | Nursery | ceFAZolin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:07 | 00:08 | 01:05 | 00:30 | mg/kg | 8 | 55 | 60 | mg/mL |  | 45 | 55
496 | Nursery | cefotaxime |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:12 | 00:13 | 01:05 | 00:30 | mg/kg | 12.5 | 80 | 81 | mg/mL |  | 45 | 55
497 | Nursery | cefTRIAXone |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:45 | 01:10 | 01:00 | mg/kg | 20 | 135 | 140 | mg/mL |  | 35 | 45 | CeftriSyringe
498 | Nursery | clindamycin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:08 | 00:09 | 01:05 | 00:30 | mg/kg | 3 | 7.1 | 7.2 | mg/mL |  | 13 | 17
499 | Nursery | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:25 | 02:15 | 00:30 | mg/kg | 0.5 | 7 | 7.1 | mg/mL |  | 3 | 7
500 | Nursery | nafcillin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:25 | 01:10 | 00:30 | mg/kg | 10 | 53 | 55 | mg/mL |  | 35 | 45
501 | Nursery | PenG Potass Syringe |  | - - | unit | - - | mL |  | unit/mL | Syringe | Primary Only | 00:15 |  | 02:00 | 00:30 | unit/kg | 20 | 52.5 | 152 | unit/mL | 10 |  | 150 | PenG-Syringe
502 | Nursery | piperacillin/tazo |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:25 | 04:15 | 00:30 | mg/kg | 49 | 125 | 130 | mg/mL |  | 15 | 85
503 | Nursery | vancomycin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:52 | 00:55 | 03:10 | 01:00 | mg/kg | 4 | 22 | 23 | mg/mL |  | 4.5 | 6.5
504 | ONC Infusion Center | abatacept |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
505 | ONC Infusion Center | abatacept |  | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
506 | ONC Infusion Center | abatacept |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
507 | ONC Infusion Center | aprepitant |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:30 | mg | 99 | 131 |  | n/a |  |  |  | DEHP
508 | ONC Infusion Center | aprepitant |  | 130 | mg | 100 | mL | 1.3 | mg/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:30 | mg | 99 | 131 |  | n/a |  |  |  | DEHP
509 | ONC Infusion Center | atezolizumab |  | 840 | mg | 250 | mL | 3.36 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 799 | 1681 |  | n/a
510 | ONC Infusion Center | atezolizumab |  | 1200 | mg | 250 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 799 | 1681 |  | n/a
511 | ONC Infusion Center | atezolizumab |  | 1680 | mg | 250 | mL | 6.72 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 799 | 1681 |  | n/a
512 | ONC Infusion Center | Bendamusti (Bendeka) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:08 | 00:15 | 00:10 | mg/m2 | 25 | 120 |  | mg/mL |  | 1.85 | 5.6
513 | ONC Infusion Center | Bendamusti (Treanda) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:20 | 00:30 | mg/m2 | 25 | 120 |  | mg/mL |  | 0.2 | 0.7
514 | ONC Infusion Center | bevacizumab | First Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:15 | 01:30 | 03:00 | 01:30 | mg/kg | 5 | 15 |  | mg/mL |  | 1 | 25
515 | ONC Infusion Center | bevacizumab | Second Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 01:30 | 01:00 | mg/kg | 5 | 15 |  | mg/mL |  | 1 | 25
516 | ONC Infusion Center | bevacizumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 01:30 | 00:30 | mg/kg | 5 | 15 |  | mg/mL |  | 1 | 25
517 | ONC Infusion Center | bleomycin | BSA Based Dose | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:05 | 00:15 | 00:10 | unit/m2 | 2 | 30 |  | unit/mL |  | 0.1 | 3
518 | ONC Infusion Center | bleomycin | Non-Based Based Dose | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:08 | 00:20 | 00:10 | unit | 1 | 60 | 61 | unit/mL |  | 0.1 | 3
519 | ONC Infusion Center | brentuximab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 00:40 | 00:30 | mg/kg | 1 | 1.8 | 1.81 | mg/mL |  | 0.4 | 1.8 | MAX DOSE 180 mg
520 | ONC Infusion Center | cabazitaxel |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:10 | 01:00 | mg/m2 | 15 | 25 |  | mg/mL |  | 0.1 | 0.26 | 0.22 Filter Needed
521 | ONC Infusion Center | CARBOplatin | Based on AUC | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 24:00 | 00:30 | mg | 100 | 800 |  | mg/mL |  | 0.4 | 5
522 | ONC Infusion Center | CARBOplatin | Based on BSA | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 | 00:30 | mg/m2 | 60 | 700 |  | mg/mL |  | 0.5 | 5
523 | ONC Infusion Center | carfilzomib |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:35 | 00:30 | mg/m2 | 15 | 70 |  | mg/mL |  | 0.3 | 1.6
524 | ONC Infusion Center | cemiplimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 00:35 | 00:30 | mg | 349 | 351 |  | mg/mL |  | 1 | 20 | In-line filter
525 | ONC Infusion Center | cetuximab | Loading Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:45 |  | 04:00 | 02:00 | mg/m2 | 400 | 500 |  | No |  |  |  | Low Binding Filter
526 | ONC Infusion Center | cetuximab | Maintenance Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg/m2 | 150 | 250 |  | No |  |  |  | Low Binding Filter
527 | ONC Infusion Center | CISplatin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 24:00 |  | mg/m2 | 10 | 100 |  | mg/mL |  | 0.04 | 0.5
528 | ONC Infusion Center | cyclophosphamide | BSA Based Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:30 | 01:00 | mg/m2 | 100 | 2000 |  | mg/mL | 1.9 | 2 | 20
529 | ONC Infusion Center | cyclophosphamide | Weight Based | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:30 |  | mg/kg | 1 | 50 | 51 | mg/mL | 1.9 | 2 | 20
530 | ONC Infusion Center | dacarbazine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:40 |  | mg/m2 | 175 | 850 |  | mg/mL |  | 0.19 | 20.1
531 | ONC Infusion Center | daratumumab | First Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 02:20 | 02:30 | 20:00 |  | mg/kg | 7 | 16 | 16.1 | mg/mL |  | 0.4 | 4 | 0.22 Filter Needed
532 | ONC Infusion Center | daratumumab | Second Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 02:30 | 04:00 | 10:00 |  | mg/kg | 15 | 16 | 16.1 | mg/mL |  | 0.9 | 5 | 0.22 Filter Needed
533 | ONC Infusion Center | daratumumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 02:30 | 03:50 | 10:00 |  | mg/kg | 15 | 16 | 16.1 | mg/mL |  | 0.9 | 5 | 0.22 Filter Needed
534 | ONC Infusion Center | decitabine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 |  | mg/m2 | 10 | 22 |  | mg/mL | 0.089 | 0.09 | 1.1
535 | ONC Infusion Center | dexamethasone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:30 | 00:15 | mg | 1 | 21 |  | No
536 | ONC Infusion Center | DOCEtaxel |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg/m2 | 16 | 100 |  | mg/mL |  | 0.3 | 0.74 | DEHP
537 | ONC Infusion Center | DOXOrubicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 24:00 |  | mg/m2 | 3 | 90 |  | mg/mL |  | 0.01 | 5
538 | ONC Infusion Center | DOXOrubicin, liposo |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg/m2 | 2.5 | 55 |  | mg/mL |  | 0.1 | 0.36 | NO FILTER
539 | ONC Infusion Center | durvalumab | .Weight >= 30 kg | 1500 | mg | 280 | mL | 5.3571 | mg/mL | Pump | Primary Only |  | 00:55 | 08:00 | 01:00 | mg | 1499 |  | 1501 | n/a |  |  |  | Low Binding Filter
540 | ONC Infusion Center | durvalumab | Weight < 30 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only |  | 00:55 | 08:00 | 01:00 | mg/kg | 9 | 21 |  | mg/mL | 0.9 | 1 | 15 | Low Binding Filter
541 | ONC Infusion Center | eriBULin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:02 | 00:20 | 00:10 | mg/m2 | 0.7 | 1.4 |  | mg/mL |  | 0.001 | 0.5 | No Dextrose
542 | ONC Infusion Center | etoposide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 03:00 |  | mg/m2 | 35 | 750 |  | mg/mL |  | 0.2 | 0.4 | Non-PVC
543 | ONC Infusion Center | ferumoxytol | .MG Dosing | 510 | mg | 117 | mL | 4.359 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:30 | mg | 509 | 1021 | 1022 | n/a
544 | ONC Infusion Center | ferumoxytol | .MG Dosing | 1020 | mg | 284 | mL | 3.5915 | mg/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | mg | 509 | 1021 | 1022 | n/a
545 | ONC Infusion Center | ferumoxytol | GM dosing | 0.51 | gram | 117 | mL | 0.0044 | gram/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:30 | gram | 0.509 | 1.021 | 1.022 | n/a
546 | ONC Infusion Center | ferumoxytol | GM dosing | 1.02 | gram | 284 | mL | 0.0036 | gram/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | gram | 0.509 | 1.021 | 1.022 | n/a
547 | ONC Infusion Center | fluorouracil |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 24:00 |  | mg/m2 | 200 | 3000 |  | mg/mL |  | 0.1 | 50
548 | ONC Infusion Center | fosaprepitant |  | 150 | mg | 150 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 00:40 | 00:30 | mg | 149 | 151 |  | n/a
549 | ONC Infusion Center | gemcitabine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg/m2 | 300 | 1250 |  | mg/mL |  | 0.1 | 38
550 | ONC Infusion Center | golimumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:00 | 00:30 | mg/kg | 1.8 |  | 2.2 | mg/mL | 0.49 | 0.5 | 4 | Low Binding Filter
551 | ONC Infusion Center | ifosfamide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 25:00 |  | mg/m2 | 250 | 5000 |  | mg/mL |  | 0.6 | 80
552 | ONC Infusion Center | ipilimumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:25 | 01:35 | 01:30 | mg/kg | 1 | 10 |  | mg/mL |  | 1 | 2
553 | ONC Infusion Center | irinotecan |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:20 |  | 02:30 | 01:30 | mg/m2 | 20 | 600 |  | mg/mL |  | 0.12 | 2.8
554 | ONC Infusion Center | irinotecan liposomal |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:20 | 02:00 | 01:30 | mg/m2 | 34 | 71 |  | mg/mL | 0.05 | 0.1 | 0.6 | NO FILTER
555 | ONC Infusion Center | leucovorin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mg/m2 | 10 | 1000 |  | mg/mL |  | 0.05 | 20
556 | ONC Infusion Center | leucovorin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mg | 1 | 100 |  | mg/mL |  | 0.05 | 20
557 | ONC Infusion Center | mag sulfate |  | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 02:30 | 01:00 | gram | 0.9 | 1.1 |  | n/a
558 | ONC Infusion Center | mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:45 | 08:00 |  | gram | 0.9 | 1.1 |  | No
559 | ONC Infusion Center | mesna |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 25:00 |  | mg/m2 | 300 | 5000 |  | mg/mL |  | 0.01 | 20
560 | ONC Infusion Center | mesna |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 25:00 |  | mg | 1 | 1000 |  | mg/mL |  | 0.01 | 20
561 | ONC Infusion Center | mitoMYcin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg/m2 | 1 | 20 |  | mg/mL |  | 0.02 | 0.04
562 | ONC Infusion Center | mitoMYcin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg | 1 | 20 |  | mg/mL |  | 0.02 | 0.04
563 | ONC Infusion Center | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
564 | ONC Infusion Center | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
565 | ONC Infusion Center | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
566 | ONC Infusion Center | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
567 | ONC Infusion Center | nivolumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | mg/mL |  | 1 | 10 | In-line filter
568 | ONC Infusion Center | obinutuzumab | First Dose | 100 | mg | 104 | mL | 0.9615 | mg/mL | Pump | Primary and Secondary |  | 04:00 | 08:00 | 04:00 | mg | 99 | 100.1 |  | n/a
569 | ONC Infusion Center | obinutuzumab | Subsequent Doses | 900 | mg | 286 | mL | 3.1469 | mg/mL | Pump | Primary and Secondary |  | 02:00 | 60:00 |  | mg | 900 | 1000 |  | n/a
570 | ONC Infusion Center | obinutuzumab | Subsequent Doses | 1000 | mg | 290 | mL | 3.4483 | mg/mL | Pump | Primary and Secondary |  | 02:00 | 60:00 |  | mg | 900 | 1000 |  | n/a
571 | ONC Infusion Center | obinutuzumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 02:00 | 60:00 |  | mg | 900 | 1000 |  | mg/mL |  | 0.9 | 3.5
572 | ONC Infusion Center | ondansetron |  | 8 | mg | 54 | mL | 0.1481 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 4 | 16 |  | n/a
573 | ONC Infusion Center | ondansetron |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 |  | mg | 4 | 16 |  | mg/mL |  | 0.01 | 0.64
574 | ONC Infusion Center | oxaliplatin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:00 | 06:30 | 02:00 | mg/m2 | 20 | 130 |  | mg/mL |  | 0.2 | 1.2
575 | ONC Infusion Center | PACLitaxel |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 24:00 |  | mg/m2 | 40 | 225 |  | mg/mL |  | 0.3 | 1.2 | Filter and DEHP Free
576 | ONC Infusion Center | PACLitaxel protein-b |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/m2 | 49 | 261 |  | mg/mL |  | 4.9 | 5.1 | NO FILTER
577 | ONC Infusion Center | panitumumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg/kg | 3 | 6 |  | mg/mL |  | 2 | 10 | Low Binding Filter
578 | ONC Infusion Center | pembrolizumab |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 00:35 | 00:30 | mg | 100 | 250 |  | n/a |  |  |  | In-line filter
579 | ONC Infusion Center | pembrolizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 00:35 | 00:30 | mg | 100 | 250 |  | mg/mL |  | 1 | 10 | In-line filter
580 | ONC Infusion Center | PEMEtrexed |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:10 | mg/m2 | 250 | 900 |  | mg/mL |  | 3 | 13
581 | ONC Infusion Center | pertuzumab | Loading Dose | 840 | mg | 278 | mL | 3.0216 | mg/mL | Pump | Primary and Secondary | 01:00 |  | 02:00 | 01:00 | mg | 839 |  | 841 | n/a
582 | ONC Infusion Center | pertuzumab | Maintenance Dose | 420 | mg | 264 | mL | 1.5909 | mg/mL | Pump | Primary and Secondary | 00:30 |  | 02:00 | 00:30 | mg | 419 |  | 421 | n/a
583 | ONC Infusion Center | ramucirumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 |  | mg/kg | 5 | 10 |  | mg/mL |  | 0.8 | 12 | Low Binding Filter
584 | ONC Infusion Center | raNITIdine |  | 50 | mg | 102 | mL | 0.4902 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:20 | mg | 25 | 50 |  | n/a
585 | ONC Infusion Center | riTUXimab | Based on BSA | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg/m2 | 250 | 500 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
586 | ONC Infusion Center | riTUXimab | Rheumatoid Arthritis | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg | 90 | 1200 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
587 | ONC Infusion Center | temozolomide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:20 | 02:00 | 01:30 | mg/m2 | 74 | 201 |  | mg/mL |  | 2.4 | 2.6 | Flush Before/AfterNS
588 | ONC Infusion Center | topotecan |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg/m2 | 0.2 | 1.5 | 4 | mg/mL |  | 0.01 | 0.5
589 | ONC Infusion Center | trastuzumab | Loading Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:00 | 03:00 | 01:30 | mg/kg | 4 | 8 |  | mg/mL |  | 0.4 | 8
590 | ONC Infusion Center | trastuzumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 2 | 6 |  | mg/mL |  | 0.4 | 8
591 | ONC Infusion Center | vinBLAStine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 01:00 | 00:15 | mg/m2 | 0.75 | 18.5 | 18.6 | mg/mL |  | 0.001 | 0.6 | Vesicant
592 | ONC Infusion Center | vinCRIStine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:20 | 00:10 | mg | 0.4 | 2 | 2.1 | mg/mL |  | 0.001 | 0.1 | Vesicant
593 | ONC Infusion Center | vinorelbine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:06 | 00:30 | 00:10 | mg/m2 | 5 | 35 |  | mg/mL |  | 0.5 | 2
594 | ONC Infusion Center | ziv-aflibercept |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:55 | 01:05 | 01:00 | mg/kg | 2 | 4 |  | mg/mL |  | 0.6 | 8 | 0.22 Filter Needed
595 | ONC Infusion Center | zoledronic acid |  | 4 | mg | 100 | mL | 0.04 | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 02:00 | 00:30 | mg | 3 | 4 |  | n/a
596 | OutPt InfusionCenter | abatacept |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
597 | OutPt InfusionCenter | abatacept |  | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
598 | OutPt InfusionCenter | abatacept |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
599 | OutPt InfusionCenter | abatacept |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg/kg | 9 |  | 11 | mg/mL | 2 | 3.8 | 10.2 | Low Binding 0.2or1.2
600 | OutPt InfusionCenter | albumin 25% |  | 5 | gram | 20 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:08 | 01:00 | 00:20 | gram | 4.9 | 25.1 |  | n/a
601 | OutPt InfusionCenter | albumin 25% |  | 12.5 | gram | 50 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:16 | 02:00 | 00:50 | gram | 4.9 | 25.1 |  | n/a
602 | OutPt InfusionCenter | albumin 25% |  | 25 | gram | 100 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:32 | 04:00 | 01:40 | gram | 4.9 | 25.1 |  | n/a
603 | OutPt InfusionCenter | amikacin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 |  | mg/kg | 1 | 21 | 39 | mg/mL | 0.2 |  | 5.1
604 | OutPt InfusionCenter | ampho B (liposomal) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only |  | 00:55 | 04:00 | 02:00 | mg/kg | 1.9 | 10.5 |  | mg/mL | 0.8 | 0.9 | 2.1 | AmBisome
605 | OutPt InfusionCenter | belimumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 00:55 |  | 08:00 | 01:00 | mg/kg | 9 |  | 11 | mg/mL | 1.6 |  | 8.5
606 | OutPt InfusionCenter | bezlotoxumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 | 00:51 | 03:00 | 01:00 | mg/kg | 9 | 11 | 11.1 | mg/mL | 0.8 | 0.9 | 10.1 | 0.2 to 5 micron filt
607 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
608 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
609 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
610 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
611 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | mg/mL | 5 | 10 | 40
612 | OutPt InfusionCenter | ceFAZolin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
613 | OutPt InfusionCenter | ceFAZolin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
614 | OutPt InfusionCenter | ceFAZolin | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
615 | OutPt InfusionCenter | ceFAZolin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
616 | OutPt InfusionCenter | ceFAZolin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | gram/mL | 0.005 | 0.01 | 0.04
617 | OutPt InfusionCenter | cefepime | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
618 | OutPt InfusionCenter | cefepime | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
619 | OutPt InfusionCenter | cefepime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
620 | OutPt InfusionCenter | cefepime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
621 | OutPt InfusionCenter | cefepime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
622 | OutPt InfusionCenter | cefepime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
623 | OutPt InfusionCenter | cefepime | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 4 | 4.9 | 40
624 | OutPt InfusionCenter | cefepime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
625 | OutPt InfusionCenter | cefepime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
626 | OutPt InfusionCenter | cefepime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
627 | OutPt InfusionCenter | cefepime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
628 | OutPt InfusionCenter | cefepime | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.004 | 0.009 | 0.04
629 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
630 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
631 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
632 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
633 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
634 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
635 | OutPt InfusionCenter | cefTAZidime | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
636 | OutPt InfusionCenter | cefTAZidime | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
637 | OutPt InfusionCenter | cefTAZidime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
638 | OutPt InfusionCenter | cefTAZidime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
639 | OutPt InfusionCenter | cefTAZidime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
640 | OutPt InfusionCenter | cefTAZidime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
641 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
642 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
643 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
644 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
645 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | mg/mL | 4.9 | 9 | 40
646 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
647 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
648 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
649 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
650 | OutPt InfusionCenter | cefTRIAXone | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | gram/mL | 0.004 | 0.009 | 0.04
651 | OutPt InfusionCenter | ciprofloxacin |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
652 | OutPt InfusionCenter | ciprofloxacin |  | 400 | mg | 200 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
653 | OutPt InfusionCenter | clindamycin |  | 300 | mg | 50 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
654 | OutPt InfusionCenter | clindamycin |  | 300 | mg | 100 | mL | 3 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
655 | OutPt InfusionCenter | clindamycin |  | 600 | mg | 50 | mL | 12 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
656 | OutPt InfusionCenter | clindamycin |  | 600 | mg | 100 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
657 | OutPt InfusionCenter | clindamycin |  | 900 | mg | 50 | mL | 18 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
658 | OutPt InfusionCenter | clindamycin |  | 900 | mg | 100 | mL | 9 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
659 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 375 | mg | 100 | mL | 3.75 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
660 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 375 | mg | 250 | mL | 1.5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
661 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
662 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 500 | mg | 250 | mL | 2 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
663 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
664 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
665 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 1125 | mg | 250 | mL | 4.5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
666 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
667 | OutPt InfusionCenter | dalbavancin | GM dosing | 0.375 | gram | 100 | mL | 0.0038 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
668 | OutPt InfusionCenter | dalbavancin | GM dosing | 0.375 | gram | 250 | mL | 0.0015 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
669 | OutPt InfusionCenter | dalbavancin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
670 | OutPt InfusionCenter | dalbavancin | GM dosing | 0.5 | gram | 250 | mL | 0.002 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
671 | OutPt InfusionCenter | dalbavancin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
672 | OutPt InfusionCenter | dalbavancin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
673 | OutPt InfusionCenter | dalbavancin | GM dosing | 1.125 | gram | 250 | mL | 0.0045 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
674 | OutPt InfusionCenter | dalbavancin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
675 | OutPt InfusionCenter | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg | 100 | 2000 |  | mg/mL | 1 | 2 | 20
676 | OutPt InfusionCenter | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg/kg | 4 | 12.1 |  | mg/mL | 1 | 2 | 20
677 | OutPt InfusionCenter | dexamethasone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:30 | 00:15 | mg | 1 | 21 |  | No
678 | OutPt InfusionCenter | doxycycline |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
679 | OutPt InfusionCenter | doxycycline |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
680 | OutPt InfusionCenter | ertapenem | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
681 | OutPt InfusionCenter | ertapenem | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
682 | OutPt InfusionCenter | ertapenem | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
683 | OutPt InfusionCenter | ertapenem | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
684 | OutPt InfusionCenter | ertapenem | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
685 | OutPt InfusionCenter | ertapenem | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
686 | OutPt InfusionCenter | ertapenem | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
687 | OutPt InfusionCenter | ertapenem | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
688 | OutPt InfusionCenter | erythromycin |  | 250 | mg | 100 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
689 | OutPt InfusionCenter | erythromycin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
690 | OutPt InfusionCenter | erythromycin |  | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:30 | 01:00 | mg | 100 | 1000 |  | n/a
691 | OutPt InfusionCenter | erythromycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | mg/mL | 0.05 | 1 | 5
692 | OutPt InfusionCenter | famotidine |  | 20 | mg | 50 | mL | 0.4 | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg | 8 | 44 | 50 | n/a
693 | OutPt InfusionCenter | famotidine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 02:00 |  | mg | 8 | 44 | 50 | mg/mL | 0.09 | 0.1 | 3.5
694 | OutPt InfusionCenter | ferumoxytol | .MG Dosing | 510 | mg | 117 | mL | 4.359 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 509 | 1021 | 1022 | n/a |  |  |  | Hypersensitivity
695 | OutPt InfusionCenter | ferumoxytol | .MG Dosing | 1020 | mg | 284 | mL | 3.5915 | mg/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | mg | 509 | 1021 | 1022 | n/a |  |  |  | Hypersensitivity
696 | OutPt InfusionCenter | ferumoxytol | GM dosing | 0.51 | gram | 117 | mL | 0.0044 | gram/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | gram | 0.509 | 1.021 | 1.022 | n/a |  |  |  | Hypersensitivity
697 | OutPt InfusionCenter | ferumoxytol | GM dosing | 1.02 | gram | 284 | mL | 0.0036 | gram/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | gram | 0.509 | 1.021 | 1.022 | n/a |  |  |  | Hypersensitivity
698 | OutPt InfusionCenter | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 |  | mg | 10 | 1000 |  | mg/mL | 0.05 | 0.1 | 10
699 | OutPt InfusionCenter | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 |  | mg/kg | 0.5 | 10 |  | mg/mL | 0.05 | 0.1 | 10
700 | OutPt InfusionCenter | golimumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:00 | 00:30 | mg/kg | 1.8 |  | 2.2 | mg/mL | 0.49 | 0.5 | 4 | Low Binding Filter
701 | OutPt InfusionCenter | immune glob (IVIG) |  | 5 | gram | 50 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:00 |  | gram | 1 | 10.1 |  | n/a |  |  |  | Hypersensitivity
702 | OutPt InfusionCenter | immune glob (IVIG) |  | 10 | gram | 100 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:40 | 08:00 |  | gram | 1 | 10.1 |  | n/a |  |  |  | Hypersensitivity
703 | OutPt InfusionCenter | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg | 50 | 2000 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
704 | OutPt InfusionCenter | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg/kg | 2 | 10 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
705 | OutPt InfusionCenter | iron sucrose |  | 100 | mg | 105 | mL | 0.9524 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 08:00 | 00:30 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
706 | OutPt InfusionCenter | iron sucrose |  | 200 | mg | 110 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 08:00 | 00:30 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
707 | OutPt InfusionCenter | iron sucrose |  | 300 | mg | 265 | mL | 1.1321 | mg/mL | Pump | Primary and Secondary | 01:25 |  | 08:00 | 01:30 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
708 | OutPt InfusionCenter | iron sucrose |  | 400 | mg | 270 | mL | 1.4815 | mg/mL | Pump | Primary and Secondary | 02:25 |  | 08:00 | 02:30 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
709 | OutPt InfusionCenter | iron sucrose |  | 500 | mg | 275 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 03:25 |  | 08:00 | 04:00 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
710 | OutPt InfusionCenter | levoFLOXacin |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
711 | OutPt InfusionCenter | levoFLOXacin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
712 | OutPt InfusionCenter | levoFLOXacin |  | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 01:15 | 02:00 | 01:30 | mg | 250 | 750 |  | n/a
713 | OutPt InfusionCenter | mag sulfate |  | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 02:00 | gram | 0.9 | 7 | 8 | n/a
714 | OutPt InfusionCenter | mag sulfate |  | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:30 | 06:00 | 04:00 | gram | 0.9 | 7 | 8 | n/a
715 | OutPt InfusionCenter | mag sulfate |  | 4 | gram | 100 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 01:00 | 09:00 | 08:00 | gram | 0.9 | 7 | 8 | n/a
716 | OutPt InfusionCenter | mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:30 | 12:00 |  | gram | 0.9 | 7 | 8 | gram/mL | 0.005 | 0.009 | 0.12
717 | OutPt InfusionCenter | meropenem | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
718 | OutPt InfusionCenter | meropenem | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
719 | OutPt InfusionCenter | meropenem | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
720 | OutPt InfusionCenter | meropenem | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
721 | OutPt InfusionCenter | meropenem | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
722 | OutPt InfusionCenter | meropenem | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
723 | OutPt InfusionCenter | meropenem | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
724 | OutPt InfusionCenter | meropenem | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
725 | OutPt InfusionCenter | meropenem | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
726 | OutPt InfusionCenter | meropenem | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
727 | OutPt InfusionCenter | methylPREDNISolone |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | n/a
728 | OutPt InfusionCenter | methylPREDNISolone |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | n/a
729 | OutPt InfusionCenter | methylPREDNISolone |  | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 125 | 2000 |  | n/a
730 | OutPt InfusionCenter | methylPREDNISolone |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 125 | 2000 |  | n/a
731 | OutPt InfusionCenter | methylPREDNISolone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 |  | mg | 125 | 2000 |  | mg/mL | 0.5 | 1 | 25
732 | OutPt InfusionCenter | micafungin |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
733 | OutPt InfusionCenter | micafungin |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
734 | OutPt InfusionCenter | micafungin |  | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
735 | OutPt InfusionCenter | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
736 | OutPt InfusionCenter | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
737 | OutPt InfusionCenter | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
738 | OutPt InfusionCenter | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
739 | OutPt InfusionCenter | nivolumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | mg/mL |  | 1 | 10 | In-line filter
740 | OutPt InfusionCenter | ocrelizumab |  | 300 | mg | 260 | mL | 1.1538 | mg/mL | Pump | Primary and Secondary | 02:20 | 02:25 | 24:00 |  | mg | 290 | 610 | 611 | n/a |  |  |  | 0.22 Filter Needed
741 | OutPt InfusionCenter | ocrelizumab |  | 600 | mg | 520 | mL | 1.1538 | mg/mL | Pump | Primary and Secondary | 01:50 | 01:55 | 24:00 |  | mg | 290 | 610 | 611 | n/a |  |  |  | 0.22 Filter Needed
742 | OutPt InfusionCenter | ondansetron |  | 8 | mg | 54 | mL | 0.1481 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 4 | 16 |  | n/a
743 | OutPt InfusionCenter | ondansetron |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 |  | mg | 4 | 16 |  | mg/mL | 0.02 | 0.04 | 0.64
744 | OutPt InfusionCenter | oritavacin | Kimsyra gram | 1.2 | gram | 250 | mL | 0.0048 | gram/mL | Pump | Primary and Secondary | 00:55 |  | 01:15 | 01:00 | gram | 1.199 |  | 1.201 | n/a |  |  |  | Flush Before/After
745 | OutPt InfusionCenter | oritavacin | Kimyrsa mg | 1200 | mg | 250 | mL | 4.8 | mg/mL | Pump | Primary and Secondary | 00:55 |  | 01:15 | 01:00 | mg | 1199 |  | 1201 | n/a |  |  |  | Flush Before/After
746 | OutPt InfusionCenter | oritavacin | Orbactiv gram | 1.2 | gram | 250 | mL | 0.0048 | gram/mL | Pump | Primary and Secondary | 02:45 |  | 03:30 | 03:00 | gram | 1.199 |  | 1.201 | n/a |  |  |  | Flush Before/AfterD5
747 | OutPt InfusionCenter | oritavacin | Orbactiv mg | 1200 | mg | 1000 | mL | 1.2 | mg/mL | Pump | Primary and Secondary | 02:45 |  | 03:30 | 03:00 | mg | 1199 |  | 1201 | n/a |  |  |  | Flush Before/AfterD5
748 | OutPt InfusionCenter | piperacillin/tazo |  | 3.375 | gram | 50 | mL | 0.0675 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
749 | OutPt InfusionCenter | piperacillin/tazo |  | 3.375 | gram | 100 | mL | 0.0338 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
750 | OutPt InfusionCenter | piperacillin/tazo |  | 4.5 | gram | 50 | mL | 0.09 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
751 | OutPt InfusionCenter | piperacillin/tazo |  | 4.5 | gram | 100 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
752 | OutPt InfusionCenter | potassium chloride |  | 10 | mEq | 100 | mL | 0.1 | mEq/mL | Pump | Primary and Secondary | 00:25 | 00:45 | 06:00 | 01:00 | mEq | 5 |  | 41 | n/a
753 | OutPt InfusionCenter | potassium chloride |  | 40 | mEq | 520 | mL | 0.0769 | mEq/mL | Pump | Primary and Secondary | 03:15 |  | 12:00 | 04:00 | mEq | 5 |  | 41 | n/a
754 | OutPt InfusionCenter | potassium chloride |  | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:50 | 01:45 | 12:00 |  | mEq | 5 |  | 41 | mEq/mL | 0.04 | 0.07 | 0.11
755 | OutPt InfusionCenter | promethazine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | mg | 6.25 | 25 |  | mg/mL | 1 | 1.25 | 25
756 | OutPt InfusionCenter | raNITIdine |  | 50 | mg | 102 | mL | 0.4902 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:20 | mg | 25 | 50 |  | n/a
757 | OutPt InfusionCenter | rifAMPin |  | 300 | mg | 250 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
758 | OutPt InfusionCenter | rifAMPin |  | 600 | mg | 500 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
759 | OutPt InfusionCenter | rifAMPin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | mg/mL | 0.1 | 0.5 | 6
760 | OutPt InfusionCenter | riTUXimab | Based on BSA | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg/m2 | 250 | 500 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
761 | OutPt InfusionCenter | riTUXimab | Rheumatoid Arthritis | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg | 90 | 1200 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
762 | OutPt InfusionCenter | spesolimab |  | 900 | mg | 100 | mL | 9 | mg/mL | Pump | Primary Only |  | 01:25 | 03:00 | 01:30 | mg | 899 |  | 901 | n/a |  |  |  | Low Binding Filter
763 | OutPt InfusionCenter | teprotumumab | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 | 00:55 | 08:00 |  | mg/kg | 7 | 21 | 22 | mg/mL | 2.9 | 3 | 18.1
764 | OutPt InfusionCenter | teprotumumab | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 | 00:55 | 08:00 |  | gram/kg | 0.007 | 0.021 | 0.022 | gram/mL | 0.002 | 0.003 | 0.019
765 | OutPt InfusionCenter | tocilizumab |  | 800 | mg | 100 | mL | 8 | mg/mL | Pump | Primary Only | 00:55 |  | 04:00 | 01:00 | mg | 161 |  | 801 | n/a
766 | OutPt InfusionCenter | tocilizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 00:55 |  | 04:00 | 01:00 | mg/kg | 3.9 |  | 8.1 | mg/mL | 1 | 1.1 | 8.1
767 | OutPt InfusionCenter | ustekinumab |  | 260 | mg | 250 | mL | 1.04 | mg/mL | Pump | Primary Only | 00:55 |  | 03:00 | 01:00 | mg | 259 |  | 521 | n/a |  |  |  | Low Binding Filter
768 | OutPt InfusionCenter | ustekinumab |  | 390 | mg | 250 | mL | 1.56 | mg/mL | Pump | Primary Only | 00:55 |  | 03:00 | 01:00 | mg | 259 |  | 521 | n/a |  |  |  | Low Binding Filter
769 | OutPt InfusionCenter | ustekinumab |  | 520 | mg | 250 | mL | 2.08 | mg/mL | Pump | Primary Only | 00:55 |  | 03:00 | 01:00 | mg | 259 |  | 521 | n/a |  |  |  | Low Binding Filter
770 | OutPt InfusionCenter | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 500 | 2000 |  | n/a
771 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
772 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
773 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
774 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
775 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | mg | 500 | 2000 |  | n/a
776 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg | 500 | 2000 |  | n/a
777 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg | 500 | 2000 |  | n/a
778 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
779 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
780 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
781 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
782 | OutPt InfusionCenter | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | mg | 500 | 2000 |  | mg/mL | 1 | 2.9 | 5.1
783 | OutPt InfusionCenter | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg/kg | 5 |  | 30 | n/a
784 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
785 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
786 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
787 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
788 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
789 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
790 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
791 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
792 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
793 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
794 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
795 | OutPt InfusionCenter | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 |  | mg/kg | 5 |  | 30 | mg/mL | 1 | 3 | 5
796 | OutPt InfusionCenter | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.5 | 2 |  | n/a
797 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
798 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
799 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
800 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
801 | OutPt InfusionCenter | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | gram | 0.5 | 2 |  | n/a
802 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram | 0.5 | 2 |  | n/a
803 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram | 0.5 | 2 |  | n/a
804 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
805 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
806 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
807 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
808 | OutPt InfusionCenter | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.002 | 0.006
809 | OutPt InfusionCenter | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram/kg | 0.005 |  | 0.03 | n/a
810 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
811 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
812 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
813 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
814 | OutPt InfusionCenter | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
815 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
816 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
817 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
818 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
819 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
820 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
821 | OutPt InfusionCenter | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | gram/kg | 0.005 |  | 0.03 | gram/mL | 0.001 | 0.002 | 0.006
822 | OutPt InfusionCenter | vedolizumab |  | 300 | mg | 250 | mL | 1.2 | mg/mL | Pump | Primary and Secondary | 00:28 |  | 02:00 | 00:30 | mg | 299 |  | 301 | n/a |  |  |  | Flush After NS-LR
823 | OutPt InfusionCenter | zoledronic acid |  | 4 | mg | 100 | mL | 0.04 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | n/a
824 | OutPt InfusionCenter | zoledronic acid |  | 5 | mg | 100 | mL | 0.05 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | n/a
825 | OutPt InfusionCenter | zoledronic acid |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | mg/mL | 0.028 | 0.029 | 0.051
826 | Pediatrics | acetaminophen |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 999 | 1001 |  | n/a
827 | Pediatrics | acetaminophen |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg/kg | 7 |  | 16 | mg/mL | 9 |  | 11
828 | Pediatrics | acetylcysteine | 1st Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg/kg | 142.5 |  | 157.5 | mg/mL |  | 1 | 75
829 | Pediatrics | acetylcysteine | 1st Dose >100kg | 15000 | mg | 275 | mL | 54.5455 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 14999 |  | 15001 | n/a
830 | Pediatrics | acetylcysteine | 2nd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg/kg | 47.5 |  | 52.5 | mg/mL |  | 1 | 75
831 | Pediatrics | acetylcysteine | 2nd Dose >100kg | 5000 | mg | 525 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg | 4999 |  | 5001 | n/a
832 | Pediatrics | acetylcysteine | 3rd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg/kg | 95 |  | 105 | mg/mL |  | 1 | 75
833 | Pediatrics | acetylcysteine | 3rd Dose >100kg | 10000 | mg | 1050 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg | 9999 |  | 10001 | n/a
834 | Pediatrics | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 4 | 25 |  | mg/mL | 0.1 | 1 | 7
835 | Pediatrics | albumin 25% |  | 5 | gram | 20 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:08 | 01:00 | 00:20 | gram | 4.9 | 25.1 |  | n/a
836 | Pediatrics | albumin 25% |  | 12.5 | gram | 50 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:16 | 02:00 | 00:50 | gram | 4.9 | 25.1 |  | n/a
837 | Pediatrics | albumin 25% |  | 25 | gram | 100 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:32 | 04:00 | 01:40 | gram | 4.9 | 25.1 |  | n/a
838 | Pediatrics | ampicillin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:30 |  | mg/kg | 4 | 100 |  | mg/mL | 1 | 5 | 35
839 | Pediatrics | ampicillin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:30 | 00:30 | mg | 500 | 3000 |  | mg/mL | 1 | 5 | 35
840 | Pediatrics | ampicillin/sulbact | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 |  | mg/kg | 4 | 80 |  | mg/mL | 5 | 15 | 35
841 | Pediatrics | ampicillin/sulbact | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 |  | mg | 1000 | 3000 |  | mg/mL | 5 | 15 | 35
842 | Pediatrics | antivenin crotalidae |  | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 01:00 | 11:30 | 10:00 | unit | 1 | 6 |  | unit/mL | 0.004 |  | 0.024 | Antivenin
843 | Pediatrics | azithromycin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 3 | 14 |  | mg/mL | 0.1 | 1 | 2.2
844 | Pediatrics | azithromycin | pt weight >= 40 kg | 250 | mg | 250 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:40 | 02:00 | 01:00 | mg | 100 | 600 |  | n/a
845 | Pediatrics | azithromycin | pt weight >= 40 kg | 500 | mg | 250 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:40 | 02:00 | 01:00 | mg | 100 | 600 |  | n/a
846 | Pediatrics | aztreonam | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:15 | 02:00 |  | mg/kg | 25.5 | 50 |  | mg/mL | 4 | 5 | 40
847 | Pediatrics | aztreonam | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:15 | 02:00 |  | mg | 500 | 2000 |  | mg/mL | 4 | 5 | 40
848 | Pediatrics | caffeine citrate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 06:00 |  | mg | 1 | 1200 |  | mg/mL |  | 0.01 | 10
849 | Pediatrics | caffeine citrate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 06:00 |  | mg/kg | 1 | 22 |  | mg/mL |  | 0.01 | 10
850 | Pediatrics | caffeine/sod benzoa |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 04:00 | 01:00 | mg | 50 | 501 | 1001 | mg/mL |  | 0.24 | 6 | 0.22 Filter Needed
851 | Pediatrics | caffeine/sod benzoa |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 04:00 | 01:00 | mg/kg | 1 | 22 |  | mg/mL |  | 0.24 | 6 | 0.22 Filter Needed
852 | Pediatrics | ceFAZolin | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg/kg | 9 | 55 |  | mg/mL | 1 | 5 | 40
853 | Pediatrics | ceFAZolin | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
854 | Pediatrics | ceFAZolin | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram/kg | 0.009 | 0.055 |  | gram/mL | 0.001 | 0.005 | 0.04
855 | Pediatrics | ceFAZolin | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
856 | Pediatrics | cefepime | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 15 | 55 |  | mg/mL | 1 | 5 | 40
857 | Pediatrics | cefepime | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
858 | Pediatrics | cefepime | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | gram/kg | 0.015 | 0.055 |  | gram/mL | 0.001 | 0.005 | 0.04
859 | Pediatrics | cefepime | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
860 | Pediatrics | cefotaxime | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 22 | 80 |  | mg/mL | 5 | 10 | 55
861 | Pediatrics | cefotaxime | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 55
862 | Pediatrics | cefotaxime | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram/kg | 0.022 | 0.08 |  | gram/mL | 0.005 | 0.01 | 0.055
863 | Pediatrics | cefotaxime | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.055
864 | Pediatrics | cefOXitin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg/kg | 20 | 55 |  | mg/mL | 5 | 10 | 40
865 | Pediatrics | cefOXitin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 40
866 | Pediatrics | cefTAZidime | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | mg/kg | 25 | 67 | 135 | mg/mL | 1 | 5 | 50
867 | Pediatrics | cefTAZidime | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | mg | 250 | 2000 |  | mg/mL | 1 | 5 | 50
868 | Pediatrics | cefTAZidime | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | gram/kg | 0.025 | 0.067 | 0.135 | gram/mL | 0.001 | 0.005 | 0.05
869 | Pediatrics | cefTAZidime | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | gram | 0.25 | 2 |  | gram/mL | 0.001 | 0.005 | 0.05
870 | Pediatrics | cefTRIAXone | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg/kg | 21 | 105 |  | mg/mL | 5 | 10 | 50
871 | Pediatrics | cefTRIAXone | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 50
872 | Pediatrics | cefTRIAXone | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | gram/kg | 0.021 | 0.105 |  | gram/mL | 0.005 | 0.01 | 0.05
873 | Pediatrics | cefTRIAXone | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.05
874 | Pediatrics | cefuroxime | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg/kg | 18 | 55 |  | mg/mL | 5 | 7.5 | 30
875 | Pediatrics | cefuroxime | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg | 750 | 1500 |  | mg/mL | 5 | 7.5 | 30
876 | Pediatrics | chlorothiazide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 250 | 1000 |  | mg/mL | 0.5 | 1 | 30
877 | Pediatrics | chlorothiazide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg/kg | 0.4 | 11 |  | mg/mL | 0.5 | 1 | 30
878 | Pediatrics | clindamycin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 4 | 23 |  | mg/mL | 1 | 2 | 20
879 | Pediatrics | clindamycin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 300 | 2400 |  | mg/mL | 1 | 2 | 20
880 | Pediatrics | erythromycin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg/kg | 2.5 | 15 |  | mg/mL | 1 | 2 | 5.5
881 | Pediatrics | erythromycin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg | 100 | 1000 |  | mg/mL | 1 | 2 | 5.5
882 | Pediatrics | fluconazole | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 02:00 | 01:00 | mg/kg | 3 | 12 | 14 | mg/mL | 1.5 |  | 2.5
883 | Pediatrics | fluconazole | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg | 100 | 800 |  | mg/mL | 1.5 |  | 2.5
884 | Pediatrics | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:08 | 01:00 | 00:30 | mg | 100 | 1500 |  | mg/mL | 0.9 | 1 | 30
885 | Pediatrics | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:08 | 01:00 | 00:30 | mg/kg | 2 |  | 29 | mg/mL | 0.9 | 1 | 30
886 | Pediatrics | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:15 | 00:30 | mg/kg | 1 | 10 |  | mg/mL | 0.1 | 0.5 | 11
887 | Pediatrics | levETIRAcetam | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:08 | 00:10 | 01:00 |  | mg/kg | 4.3 | 57 | 115 | mg/mL | 0.5 | 1 | 30
888 | Pediatrics | levETIRAcetam | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:08 | 00:10 | 01:00 |  | mg | 200 | 3000 | 5000 | mg/mL | 0.5 | 1 | 30
889 | Pediatrics | mannitol 20% |  | 100 | gram | 500 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 00:31 | 02:00 |  | gram | 1 | 101 |  | n/a
890 | Pediatrics | mannitol 20% |  | 100 | gram | 500 | mL | 0.2 | gram/mL | Pump | Primary and Secondary | 00:25 | 00:31 | 02:00 |  | gram/kg | 0.2 | 2 |  | n/a |  |  |  | In-line filter
891 | Pediatrics | mannitol 20% |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:10 | 02:00 |  | gram/kg | 0.2 | 2 |  | gram/mL |  | 0.1 | 0.5 | In-line filter
892 | Pediatrics | meropenem | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | mg/kg | 10 | 40 |  | mg/mL | 1 | 5 | 40
893 | Pediatrics | meropenem | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
894 | Pediatrics | meropenem | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | gram/kg | 0.01 | 0.04 |  | gram/mL | 0.001 | 0.005 | 0.04
895 | Pediatrics | meropenem | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
896 | Pediatrics | metroNIDAZOLE |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 03:00 | 01:00 | mg | 250 | 500 |  | n/a
897 | Pediatrics | metroNIDAZOLE |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 03:00 | 01:00 | mg | 250 | 500 |  | n/a
898 | Pediatrics | metroNIDAZOLE |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 03:00 | 01:00 | mg/kg | 7 | 16 |  | mg/mL | 4 |  | 6
899 | Pediatrics | micafungin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 0.85 | 7 |  | mg/mL | 0.4 |  | 4
900 | Pediatrics | micafungin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg | 40 | 150 |  | mg/mL | 0.4 |  | 4
901 | Pediatrics | nafcillin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 10.6 | 57.5 |  | mg/mL | 1 | 5 | 40
902 | Pediatrics | nafcillin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
903 | Pediatrics | oxacillin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:08 | 01:30 |  | mg/kg | 5 | 50 |  | mg/mL | 5 | 10 | 40
904 | Pediatrics | oxacillin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:08 | 01:30 |  | mg | 500 | 2000 |  | mg/mL | 5 | 10 | 40
905 | Pediatrics | penG x Million Units |  | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | unit | 1 | 6 |  | unit/mL | 0.01 | 0.02 | 0.12 | 1unit=1million
906 | Pediatrics | pentamidine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 04:30 | 02:00 | mg/kg | 2.9 |  | 4.1 | mg/mL |  | 0.6 | 30
907 | Pediatrics | PHENobarbital |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:08 | 02:00 |  | mg/kg | 0.5 | 7.5 |  | mg/mL | 0.1 | 0.5 | 135 | High Risk
908 | Pediatrics | phenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:15 | 02:00 |  | mg/kg | 1 |  | 29 | mg/mL | 1 | 4 | 55 | In-line filter
909 | Pediatrics | piperacillin/tazo | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | mg/kg | 20 | 140 |  | mg/mL | 0.5 | 10 | 100
910 | Pediatrics | piperacillin/tazo | GM dosing | 2.25 | gram | 50 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
911 | Pediatrics | piperacillin/tazo | GM dosing | 2.25 | gram | 100 | mL | 0.0225 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
912 | Pediatrics | piperacillin/tazo | GM dosing | 3.375 | gram | 50 | mL | 0.0675 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
913 | Pediatrics | piperacillin/tazo | GM dosing | 3.375 | gram | 100 | mL | 0.0338 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
914 | Pediatrics | piperacillin/tazo | GM dosing | 4.5 | gram | 50 | mL | 0.09 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
915 | Pediatrics | piperacillin/tazo | GM dosing | 4.5 | gram | 100 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
916 | Pediatrics | piperacillin/tazo | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 04:30 |  | gram/kg | 0.025 | 0.14 |  | gram/mL | 0.005 | 0.01 | 0.1
917 | Pediatrics | potassium chloride |  | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:45 | 00:50 | 06:00 |  | mEq | 1 |  | 41 | mEq/mL | 0.04 | 0.05 | 0.2
918 | Pediatrics | potassium chloride |  | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:45 | 00:50 | 06:00 |  | mEq/kg | 0.2 | 1.5 | 2.5 | mEq/mL | 0.04 | 0.05 | 0.2 | High Risk
919 | Pediatrics | potassium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 01:30 | 03:00 | 12:30 |  | mmol | 5 | 20 | 40 | mmol/mL | 0.009 | 0.01 | 0.3
920 | Pediatrics | raNITIdine | pt weight < 50 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:10 | 01:00 | 00:30 | mg/kg | 0.4 | 2.5 |  | mg/mL | 0.1 | 0.5 | 3
921 | Pediatrics | raNITIdine | pt weight >= 50 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:10 | 01:00 |  | mg | 25 | 50 | 150 | mg/mL | 0.1 | 0.5 | 3
922 | Pediatrics | rifAMPin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 03:00 |  | mg/kg | 4.3 | 23 |  | mg/mL | 0.5 | 1 | 7
923 | Pediatrics | rifAMPin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 03:00 |  | mg | 300 | 600 |  | mg/mL | 0.5 | 1 | 7
924 | Pediatrics | SMX/TMP |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:30 | mg/kg | 1.7 | 8 | 10 | mg/mL | 0.01 | 0.04 | 1 | 48 Hour Tubing
925 | Pediatrics | sodium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 00:45 | 01:00 | 10:30 |  | mmol | 1 |  | 31 | mmol/mL | 0.009 | 0.01 | 0.3 | Calcium Precipitate
926 | Pediatrics | sodium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 00:45 | 01:00 | 10:30 |  | mmol/kg | 0.04 | 0.4 | 0.6 | mmol/mL | 0.009 | 0.01 | 0.3 | Calcium Precipitate
927 | Pediatrics | tobramycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:15 |  | mg/kg | 0.8 | 7 | 12 | mg/mL | 0.01 | 0.1 | 11
928 | Pediatrics | tocilizumab |  | 800 | mg | 100 | mL | 8 | mg/mL | Pump | Primary and Secondary | 00:55 |  | 04:00 | 01:00 | mg | 161 |  | 801 | n/a
929 | Pediatrics | tocilizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 00:55 |  | 04:00 | 01:00 | mg/kg | 3.9 |  | 12.1 | mg/mL | 1 | 1.1 | 8.1
930 | Pediatrics | valproate sodium |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 |  | mg/kg | 5 | 60 |  | mg/mL | 1 |  | 25
931 | Pediatrics | vancomycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:30 | 00:50 | 04:30 |  | mg/kg | 5 | 30 |  | mg/mL | 0.5 | 1 | 5
932 | Pediatrics | voriconazole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:10 |  | mg/kg | 2.6 | 9 |  | mg/mL | 0.1 | 0.4 | 5.5
933 | Pediatrics | zidovudine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 0.8 | 9 |  | mg/mL | 0.5 | 1 | 4
Note on Y26: Existing IV lines must be flushed with D5w prior to infusion. If unable, must start a new line. **An in-fine filter >= 1 micron may be used with infusion.
Note on Y46: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y47: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y48: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y49: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y50: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y51: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y71: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y73: Use a 0.22 micron in-line filter when administering this product.
Note on Y74: Use a 0.22 micron in-line filter when administering this product.
Note on Y93: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y94: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y95: Do not infuse any other drugs through this line.
Note on Y96: Do not infuse any other drugs through this line.
Note on Y97: Do not infuse any other drugs through this line.
Note on Y200: Advise patient to lie down during and 30 minutes after infusion to prevent hypotension.
Note on Y212: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y213: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y214: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y215: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y216: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y217: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y218: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y219: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y220: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y221: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y222: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y223: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y236: Protect this drug from light.
Note on Y237: Protect this drug from light.
Note on Y254: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y255: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y256: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y257: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y258: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y259: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y260: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y261: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y270: Infusion requires the use of an in-line filter.
Note on Y284: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y285: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y286: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y287: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y288: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y289: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y290: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y291: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y292: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y293: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y294: Infusion requires the use of an in-line filter.
Note on Y295: Infusion requires the use of an in-line filter.
Note on Y296: INR 2 - 4 ** Max dose of Kcentra should not exceed 2500 units **
Note on Y297: INR 4 - 6 ** Max dose of Kcentra should not exceed 3500 units **
Note on Y298: INR > 6 ** Max dose of Kcentra should not exceed 5000 units **
Note on Y311: Infusion requires the use of an in-line filter.
Note on Y312: Infusion requires the use of an in-line filter.
Note on Y313: Infusion requires the use of an in-line filter.
Note on Y359: Do not infuse any other drugs through this line.
Note on Y361: 1 unit = 1 million units
Note on Y362: 1 unit = 1 million units
Note on Y363: 1 unit = 1 million units
Note on Y364: 1 unit = 1 million units
Note on Y365: 1 unit = 1 million units
Note on Y366: 1 unit = 1 million units
Note on Y372: Infusion requires the use of an in-line filter.
Note on Y407: Premedicate before each infusion.
Note on Y408: Change tubing every 48 hours.
Note on Y410: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y411: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y489: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y490: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y491: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y492: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y493: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y494: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y499: Adminster over 60 minutes to decrease risk of bilirubin encephalopathy.
Note on Y503: 1 unit in pump = 1,000 units on label. Eg. 200 units in pump = 200,000 units on drug label.
Note on Y506: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y507: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y508: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y509: This medication must be infused using DEHP FREE tubing.
Note on Y510: This medication must be infused using DEHP FREE tubing.
Note on Y521: MAXIMUM dose of brentuximab is 180 mg.
Note on Y522: Use a 0.22 micron in-line filter when administering this product.
Note on Y526: Infusion requires the use of an in-line filter.
Note on Y527: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y528: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y533: Use a 0.22 micron in-line filter when administering this product.
Note on Y534: Use a 0.22 micron in-line filter when administering this product.
Note on Y535: Use a 0.22 micron in-line filter when administering this product.
Note on Y538: This medication must be infused using DEHP FREE tubing.
Note on Y540: Do NOT use an in-line filter for administration.
Note on Y541: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y542: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y543: Do NOT administer through lines containing dextrose.
Note on Y544: This medication must be infused using Non-PVC tubing.
Note on Y552: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y556: Do NOT use an in-line filter for administration.
Note on Y565: Infusion requires the use of an in-line filter.
Note on Y566: Infusion requires the use of an in-line filter.
Note on Y567: Infusion requires the use of an in-line filter.
Note on Y568: Infusion requires the use of an in-line filter.
Note on Y569: Infusion requires the use of an in-line filter.
Note on Y577: Requires Low protein binding 0.2 micron filter and DEHP free bag
Note on Y578: Do NOT use an in-line filter for administration.
Note on Y579: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y580: Infusion requires the use of an in-line filter.
Note on Y581: Infusion requires the use of an in-line filter.
Note on Y585: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y587: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y588: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y589: Flush line before and after administration. Can be given on same line as sodium chloride 0.9%.
Note on Y593: Drug is a vesicant. Check blood return.
Note on Y594: Drug is a vesicant. Check blood return.
Note on Y596: Use a 0.22 micron in-line filter when administering this product.
Note on Y598: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y599: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y600: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y601: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y606: Existing IV lines must be flushed with D5w prior to infusion. If unable, must start a new line. **An in-fine filter >= 1 micron may be used with infusion.
Note on Y608: Use a low-protein binding in-line or add-on filter between 0.2 and 5 microns.
Note on Y661: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y662: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y663: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y664: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y665: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y666: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y667: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y668: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y680: Protect this drug from light.
Note on Y681: Protect this drug from light.
Note on Y696: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y697: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y698: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y699: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y702: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y703: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y704: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y705: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y706: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y707: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y708: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y709: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y710: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y711: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y737: Infusion requires the use of an in-line filter.
Note on Y738: Infusion requires the use of an in-line filter.
Note on Y739: Infusion requires the use of an in-line filter.
Note on Y740: Infusion requires the use of an in-line filter.
Note on Y741: Infusion requires the use of an in-line filter.
Note on Y742: Use a 0.22 micron in-line filter when administering this product.
Note on Y743: Use a 0.22 micron in-line filter when administering this product.
Note on Y746: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w or NS.
Note on Y747: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w or NS.
Note on Y748: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y749: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y762: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y763: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y764: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y769: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y770: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y771: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y824: Following infusion, flush with 30 mL of sterile normal saline or lactated ringers.
Note on Y844: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y852: Use a 0.22 micron in-line filter when administering this product.
Note on Y853: Use a 0.22 micron in-line filter when administering this product.
Note on Y892: Infusion requires the use of an in-line filter.
Note on Y893: Infusion requires the use of an in-line filter.
Note on Y907: 1 unit = 1 million units
Note on Y909: This is a HIGH RISK medication. 2nd RN to verify prior to initiating therapy.
Note on Y910: Infusion requires the use of an in-line filter.
Note on Y920: This is a HIGH RISK medication. 2nd RN to verify prior to initiating therapy.
Note on Y926: Change tubing every 48 hours.
Note on Y927: Do not infuse Calcium and Phosphate though the same line.
Note on Y928: Do not infuse Calcium and Phosphate though the same line.

## Fluids
Fluids
Line Item | Profile | Fluid Name | Therapy Name | Module Availability | Supports Secondary (Y/N) | Rate Units | Rate Limits (mL/h) Soft Min | Rate Limits (mL/h) Soft Max | Rate Limits (mL/h) Hard Max | Clinical Advisory
1 | Adult | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 999
2 | Adult | .IVF + KCL | 10 mEq/L | Pump | Y | mL/h | 1 | 999
3 | Adult | .IVF + KCL | 15 mEq/L | Pump | Y | mL/h | 1 | 666 | 700
4 | Adult | .IVF + KCL | 20 mEq/L | Pump | Y | mL/h | 1 | 500 | 550
5 | Adult | .IVF + KCL | 30 mEq/L | Pump | Y | mL/h | 1 | 333 | 366
6 | Adult | .IVF + KCL | 40 mEq/L | Pump | Y | mL/h | 1 | 250 | 275
7 | Adult | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 1 | 999
8 | Adult | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 500 | 550
9 | Adult | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 1 | 250 | 275
10 | Adult | .Maint IVF+Additives |  | Pump | Y | mL/h | 1 |  | 999
11 | Adult | 3% Sodium Chloride |  | Pump | Y | mL/h | 1 | 50 | 100
12 | Adult | Bicarb Drip |  | Pump | Y | mL/h | 1 |  | 999
13 | Adult | Blood (whole) |  | Pump | N | mL/h | 1 | 301 |  | Dedicated IV Line
14 | Adult | Dextrose 10% |  | Pump | Y | mL/h | 1 | 150
15 | Adult | Fresh Frozen Plasma |  | Pump | N | mL/h | 1 | 201 |  | Dedicated IV Line
16 | Adult | Insulin Maternity |  | Pump | Y | mL/h | 70 | 160 | 200
17 | Adult | Parenteral Nutrition |  | Pump | N | mL/h | 19 | 95 |  | Midline Warning
18 | Adult | PRBCs |  | Pump | N | mL/h | 1 | 301 |  | Dedicated IV Line
19 | Nursery | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 15 | 25
20 | Nursery | Dextrose 10% |  | Pump | Y | mL/h | 1 | 15 | 25
21 | ONC Infusion Center | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 999
22 | ONC Infusion Center | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 1 | 999
23 | ONC Infusion Center | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 667
24 | ONC Infusion Center | .IVF + KCL 30 mEq/L |  | Pump | Y | mL/h | 1 | 334 | 667
25 | ONC Infusion Center | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 1 | 250 | 500
26 | ONC Infusion Center | .Maint IVF+Additives |  | Pump | Y | mL/h | 1 | 250
27 | ONC Infusion Center | Dextrose 10% |  | Pump | Y | mL/h | 15 | 200
28 | ONC Infusion Center | Rituximab (Rituxan) |  | Pump | Y | mL/h | 0.1 | 250
29 | OutPt InfusionCenter | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 999
30 | OutPt InfusionCenter | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 1 | 999
31 | OutPt InfusionCenter | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 667
32 | OutPt InfusionCenter | .IVF + KCL 30 mEq/L |  | Pump | Y | mL/h | 1 | 334 | 667
33 | OutPt InfusionCenter | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 1 | 250 | 500
34 | OutPt InfusionCenter | .Maint IVF+Additives |  | Pump | Y | mL/h | 1 | 250
35 | OutPt InfusionCenter | Dextrose 10% |  | Pump | Y | mL/h | 15 | 200
36 | OutPt InfusionCenter | PRBCs |  | Pump and Syringe | N | mL/h | 1 | 150 |  | Dedicated IV Line
37 | OutPt InfusionCenter | Rituximab (Rituxan) |  | Pump | Y | mL/h | 0.1 | 250
38 | Pediatrics | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 300
39 | Pediatrics | .IVF + KCL | 10 mEq/L | Pump | Y | mL/h | 1 | 999
40 | Pediatrics | .IVF + KCL | 15 mEq/L | Pump | Y | mL/h | 1 | 666 | 700
41 | Pediatrics | .IVF + KCL | 20 mEq/L | Pump | Y | mL/h | 1 | 500 | 550
42 | Pediatrics | .IVF + KCL | 30 mEq/L | Pump | Y | mL/h | 1 | 333 | 366
43 | Pediatrics | .IVF + KCL | 40 mEq/L | Pump | Y | mL/h | 1 | 250 | 275
44 | Pediatrics | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 0.1 | 200 | 999
45 | Pediatrics | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 200 | 500
46 | Pediatrics | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 0.1 | 200 | 275
47 | Pediatrics | Cryoprecipitate |  | Pump | N | mL/h | 0.1 | 300
48 | Pediatrics | Dextrose 10% |  | Pump | Y | mL/h | 1 | 50 | 150
49 | Pediatrics | Fresh Frozen Plasma |  | Pump | N | mL/h | 0.1 | 150 |  | Dedicated IV Line
50 | Pediatrics | Parenteral Nutrition |  | Pump | N | mL/h | 35 | 95 |  | Midline Warning
51 | Pediatrics | PRBCs |  | Pump | N | mL/h | 0.1 | 150 |  | Dedicated IV Line
Note on K15: Do not infuse any other drugs through this line.
Note on K17: Do not infuse any other drugs through this line.
Note on K19: Do NOT infuse through a midline catheter.
Note on K20: Do not infuse any other drugs through this line.
Note on K38: Do not infuse any other drugs through this line.
Note on K51: Do not infuse any other drugs through this line.
Note on K52: Do NOT infuse through a midline catheter.
Note on K53: Do not infuse any other drugs through this line.

## PCA Drugs
PCA Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount | Drug Amount Units | Diluent Volume | Diluent Volume Units | Final Concentration | Final Concentration Units | Max Accumulated Lockout | Dosing Units | Max Accumulated Dose Range Soft Min | Max Accumulated Dose Range Soft Max | Max Accumulated Dose Range Hard Max | Max Accumulated Dose Range Initial Value | PCA Dose Limits Soft Min | PCA Dose Limits Soft Max | PCA Dose Limits Hard Max | PCA Dose Limits Initial Value | Lock Out Interval (minutes) Hard Min | Lock Out Interval (minutes) Soft Min | Lock Out Interval (minutes) Soft Max | Lock Out Interval (minutes) Initial Value | Continuous Dose Limits Soft Min | Continuous Dose Limits Soft Max | Continuous Dose Limits Hard Max | Continuous Dose Limits Initial Value | Bolus Dose Limits Soft Min | Bolus Dose Limits Soft Max | Bolus Dose Limits Hard Max | Bolus Dose included in Accum Max Dose (Y/N) | Loading Dose Limits Soft Min | Loading Dose Limits Soft Max | Loading Dose Limits Hard Max | Conc. Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max | Supports PCA Pause Protocol (Y/N) | Clinical Advisory
1 | Adult | No drugs exist in this section
2 | Nursery | No drugs exist in this section
3 | ONC Infusion Center | No drugs exist in this section
4 | OutPt InfusionCenter | No drugs exist in this section
5 | Pediatrics | No drugs exist in this section

## PC Unit Config
| PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit
Line Item | Profile | Alarm Audio Profile | Minimum Audio Volume | Default Audio Volume | Anesthesia Mode | Battery Meter Display | Color Display | Key Click Audio | Limit Checking | Max Patient BSA (m2) | Max Patient Weight (kg) | Patient ID Entry | Preventive Maintenance Reminder | Tamper Resistant | Authorized User Mode | No Guardrails Basic Infusion Clinical Advisory | Clinical Advisory Message
1 | Adult | Profile 1 | 3 | 4 | Enabled | Enabled | Enabled | Enabled | Always | 3 | 500 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
2 | Nursery | Profile 1 | 2 | 3 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 7 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
3 | ONC Infusion Center | Profile 1 | 3 | 4 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 500 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
4 | OutPt InfusionCenter | Profile 1 | 3 | 4 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 500 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
5 | Pediatrics | Profile 1 | 3 | 4 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 250 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection

## Pump Config
| Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module
Line Item | Profile | Accumulated Air-In-Line | Air In Line (microliter) | Auto-Restart Attempts | KVO Rate Adjust (mL/h) | Max Occlusion Pressure (mmHg) | Default Pressure (mmHg) | Maximum Rate (mL/h) | Maximum VTBI (mL) | Pressure Mode Lock Status | Pressure Mode Selection | Secondary Infusion | SEC-PRI Audio Alert
1 | Adult | Enabled | 250 | 9 | 20 | 525 | 525 | 999 | 3000 | Unlocked | Pump | Enabled | Enabled
2 | Nursery | Enabled | 50 | 1 | 2 | 525 | 525 | 25 | 1200 | Unlocked | Pump | Enabled | Enabled
3 | ONC Infusion Center | Enabled | 250 | 9 | 20 | 525 | 525 | 999 | 3000 | Unlocked | Pump | Enabled | Enabled
4 | OutPt InfusionCenter | Enabled | 250 | 9 | 20 | 525 | 525 | 999 | 3000 | Unlocked | Pump | Enabled | Enabled
5 | Pediatrics | Enabled | 75 | 9 | 5 | 525 | 525 | 999 | 2200 | Unlocked | Pump | Enabled | Enabled

## Syringe Config
| Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module
Line Item | Profile | All Mode | Auto Occlusion Pressure | Back-Off After Occlusion | Fast-Start | KVO Option | KVO Rate Adjust (mL/h) | KVO Volume Adjust (%) | Near End of Infusion (Continuous, Fluids, Basic) | Alert Time (Continuous, Fluids, Basic) (minutes) | Near End of Infusion (Intermittent) | Alert Time (Intermittent) (minutes) | Occlusion Pressure Limit with Disc (mmHg) | Occlusion Pressure Limit No Disc | Priming
1 | Adult | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled
2 | Nursery | Enabled | Enabled | Enabled | Enabled | Enabled | 2 | 5 | Enabled | 2 | Enabled | 2 | 1000 | High | Enabled
3 | ONC Infusion Center | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled
4 | OutPt InfusionCenter | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled
5 | Pediatrics | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled

## Pump & Syringe Shared Config
| Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe)
Line Item | Profile | Delay Start Options | Delay Start Callback | Drug Calculation | Drug Calculation Bolus Dose | Pressure Dynamic Display | Volume/Duration
1 | Adult | Disabled | n/a | Enabled | Enabled | Enabled | Enabled
2 | Nursery | Disabled | n/a | Enabled | Enabled | Enabled | Enabled
3 | ONC Infusion Center | Disabled | n/a | Disabled | n/a | Enabled | Enabled
4 | OutPt InfusionCenter | Disabled | n/a | Disabled | n/a | Enabled | Enabled
5 | Pediatrics | Disabled | n/a | Enabled | Enabled | Enabled | Enabled

## PCA Config
| PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module
Line Item | Profile | Bolus Delivery Rate (mL/h) | Dose Request Cord (Profile) | Max Dose Limit | Time Window (h) | Module Location Enforcement | Near End of Infusion | Near End of Infusion Alert Time (%) | Pressure Limit | Priming | Security Access Level | Security Code
1 | Adult | 150 | 2 | Enabled | 1 | Enabled | Enabled | 5 | High | Enabled | 1 | n/a
2 | Nursery | 150 | 2 | Enabled | 4 | Enabled | Disabled | n/a | High | Enabled | 1 | n/a
3 | ONC Infusion Center | 150 | 2 | Enabled | 4 | Enabled | Disabled | n/a | High | Enabled | 1 | n/a
4 | OutPt InfusionCenter | 150 | 2 | Enabled | 4 | Enabled | Disabled | n/a | High | Enabled | 1 | n/a
5 | Pediatrics | 150 | 2 | Enabled | 1 | Enabled | Enabled | 5 | High | Enabled | 1 | n/a

## PCA Pause Protocol
| PCA Pause Protocol | PCA Pause Protocol | PCA Pause Protocol | PCA Pause Protocol | PCA Pause Protocol
Line Item | Profile | PCA Pause Protocol Enabled | Available For | Monitoring Module Attach Enforcement | PCA Pause Protocol
1 | Adult | Enabled | Etco2 | Disabled | PCA Infusion has paused due to a decline in respiratory status. Check Patient!
2 | Nursery | Disabled | n/a | n/a | n/a
3 | ONC Infusion Center | Disabled | n/a | n/a | n/a
4 | OutPt InfusionCenter | Disabled | n/a | n/a | n/a
5 | Pediatrics | Disabled | n/a | n/a | n/a

## PCA & Syringe Shared Config
| Shared Infusion Settings (PCA & Syringe) | Shared Infusion Settings (PCA & Syringe)
Line Item | Profile | Maximum Rate (mL/h) | NEOI Snooze | NEOI Snooze Time (minutes)
1 | Adult | 999 | Disabled | n/a
2 | Nursery | 25 | Enabled | 5
3 | ONC Infusion Center | 999 | Disabled | n/a
4 | OutPt InfusionCenter | 999 | Disabled | n/a
5 | Pediatrics | 999 | Disabled | n/a

## EtCO2 Config
| EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module
Line Item | Profile | Alarm Limit Type | EtCO2 High (mmHg) | EtCO2 Low (mmHg) | Respiratory Rate High (bpm) | Respiratory Rate Low (bpm) | No Breath Alarm (Seconds) | FiCO2 High (mmHg) | Waveform Time Scale (Seconds) | PCA Pause Resp. Rate Lower Limit (bpm) | PCA Pause Resp. Rate Initial Value (bpm)
1 | Adult | Adult | 60 | 10 | 35 | 7 | 30 | 8 | 10 | 5 | 6
2 | Nursery | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a
3 | ONC Infusion Center | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a
4 | OutPt InfusionCenter | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a
5 | Pediatrics | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a

## Channel Labels
Channel Labels | Channel Labels | Channel Labels | Channel Labels | Channel Labels
Line Item | Channel Labels are not used | Channel Labels are not used | Channel Labels are not used | Channel Labels are not used