## Data Set Properties
Data Set Properties
Data Set Name: | Conway2024v64is
Hospital Name: | Conway Medical Center
Status: | Released
Data Set ID: | 02d4a96b8-R

## ContBolus Non-Anesth Drugs
Continuous/Bolus Non-Anesthesia Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount (- - = custom concentration) | Drug Amount Units | Diluent Volume (- - = custom concentration) | Diluent Volume Units | Final Concentration | Final Concentration Units | Module Availability | Dosing Units | Dose Limits Soft Min | Dose Limits Soft Max | Dose Limits Hard Max | Dose Initial Value | Bolus Dose Units | Bolus Dose Limits Soft Min | Bolus Dose Limits Soft Max | Bolus Dose Limits Hard Max | Bolus Dose Limits Initial Value | Bolus Dose Admin. Rate Units | Bolus Dose Admin. Rate Limits Soft Min | Bolus Dose Admin. Rate Limits Soft Max | Bolus Dose Admin. Rate Limits Hard Max | Bolus Dose Admin. Rate Initial Value | Conc. Limit Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max | Clinical Advisory
1 | Adult | adenosine (CATH LAB) | .Standard wt < 115kg | 90 | mg | 90 | mL | 1 | mg/mL | Pump | mcg/kg/min | 135 | 145 |  | 140 | [10 empty] | n/a
2 | Adult | adenosine (CATH LAB) | Weight >= 115 kg | 180 | mg | 90 | mL | 2 | mg/mL | Pump | mcg/kg/min | 135 | 145 |  | 140 | [10 empty] | n/a |  |  |  | Double Concentrated
3 | Adult | alteplase -stroke | Stroke < 101 kg | - - | mg | - - | mL |  | mg/mL | Pump | mg/kg/h | 0.7 | 0.9 |  | 0.81 | [10 empty] | mg/mL | 0.85 | 0.99 | 1.01 | Alteplase wt based
4 | Adult | alteplase -stroke | Stroke 101kg or more | 81 | mg | 81 | mL | 1 | mg/mL | Pump | mg/h | 71 | 90 |  | 81 | [10 empty] | n/a
5 | Adult | amiodarone Drip |  | 360 | mg | 200 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.51 | [11 empty] | n/a |  |  |  | In-line filter
6 | Adult | amiodarone Drip |  | 450 | mg | 250 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.51 | [11 empty] | n/a |  |  |  | In-line filter
7 | Adult | amiodarone Drip |  | 900 | mg | 500 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.51 | [11 empty] | n/a |  |  |  | In-line filter
8 | Adult | argatroban | .Standard Dose | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 10 | 11.5 | 2 | [10 empty] | n/a
9 | Adult | argatroban | .Standard Dose | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 10 | 11.5 | 2 | [10 empty] | n/a
10 | Adult | argatroban | Hepatic Dose | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.1 | 5 | 10.1 | 0.5 | [10 empty] | n/a
11 | Adult | argatroban | Hepatic Dose | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.1 | 5 | 10.1 | 0.5 | [10 empty] | n/a
12 | Adult | argatroban | Undergoing PCI | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 15 | 30 | 34 | 25 | mcg/kg | 149 | 351 |  |  | mcg/kg/min | 29 | 117 |  |  | n/a
13 | Adult | argatroban | Undergoing PCI | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 15 | 30 | 34 | 25 | mcg/kg | 149 | 351 |  |  | mcg/kg/min | 29 | 117 |  |  | n/a
14 | Adult | atracurium |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 0.5 | 29.5 | 39 |  | mg/kg | 0.07 | 0.5 | 0.58 |  | mg/kg/min | 0.07 | 0.25 | 0.29 |  | mcg/mL | 200 | 5000 | 5100 | Neuromuscular BA
15 | Adult | bivalirudin | .Standard Dose | 250 | mg | 50 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.1 | 1.8 | 2 | 1.75 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.76 |  | 0.75 | n/a
16 | Adult | bivalirudin | .Standard Dose | 500 | mg | 100 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.1 | 1.8 | 2 | 1.75 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.76 |  | 0.75 | n/a
17 | Adult | bivalirudin | .Standard Dose | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.1 | 1.8 | 2 | 1.75 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.76 |  | 0.75 | n/a
18 | Adult | bivalirudin | ESRD | 250 | mg | 50 | mL | 5 | mg/mL | Pump | mcg/kg/h | 0.1 | 1.2 | 1.4 | 0.25 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
19 | Adult | bivalirudin | ESRD | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/h | 0.1 | 1.2 | 1.4 | 0.25 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
20 | Adult | bivalirudin | Renal Dose CrCl < 30 | 250 | mg | 50 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.24 | 1.2 | 1.4 | 1 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
21 | Adult | bivalirudin | Renal Dose CrCl < 30 | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.24 | 1.2 | 1.4 | 1 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
22 | Adult | bivalirudin | Renal Dose CrCl < 30 | 500 | mg | 100 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.24 | 1.2 | 1.4 | 1 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
23 | Adult | bumetanide continuos |  | 24 | mg | 96 | mL | 0.25 | mg/mL | Pump | mg/h | 0.1 | 2 |  | 0.5 | mg | 0.25 | 4 |  |  | mg/min | 0.25 |  | 0.5 |  | n/a |  |  |  | Protect From Light
24 | Adult | cangrelor |  | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 2 | 4.2 | 4.6 | 4 | mcg/kg | 26 |  | 35 | 30 | mcg/kg/min | 26 |  | 35 |  | n/a
25 | Adult | cisatracurium |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | mcg/kg/min | 0.5 | 5 | 7 | 3 | [10 empty] | n/a |  |  |  | Neuromuscular BA
26 | Adult | conivaptan |  | 20 | mg | 100 | mL | 0.2 | mg/mL | Pump | mg/h | 0.6 | 1.7 | 2 | 0.8 | mg | 10 | 20 |  | 20 | mg/min | 0.333 | 0.7 |  | 0.667 | n/a
27 | Adult | conivaptan |  | 20 | mg | 250 | mL | 0.08 | mg/mL | Pump | mg/h | 0.6 | 1.7 | 2 | 0.8 | mg | 10 | 20 |  | 20 | mg/min | 0.333 | 0.7 |  | 0.667 | n/a
28 | Adult | dexmedetomidine |  | 200 | mcg | 50 | mL | 4 | mcg/mL | Pump | mcg/kg/h | 0.05 | 0.7 | 1.5 | 0.2 | mcg/kg | 0.9 | 1.1 | 2 | 1 | mcg/kg/min | 0.01 | 0.11 | 0.2 | 0.1 | n/a
29 | Adult | dexmedetomidine |  | 400 | mcg | 100 | mL | 4 | mcg/mL | Pump | mcg/kg/h | 0.05 | 0.7 | 1.5 | 0.2 | mcg/kg | 0.9 | 1.1 | 2 | 1 | mcg/kg/min | 0.01 | 0.11 | 0.2 | 0.1 | n/a
30 | Adult | dilTIAZem |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | mg/h | 0.5 | 15 | 25 | 5 | [10 empty] | n/a
31 | Adult | dilTIAZem |  | 125 | mg | 125 | mL | 1 | mg/mL | Pump | mg/h | 0.5 | 15 | 25 | 5 | [10 empty] | n/a
32 | Adult | DOBUTamine | ..Standard | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a
33 | Adult | DOBUTamine | ..Standard | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a
34 | Adult | DOBUTamine | .Double Concentrated | 500 | mg | 250 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a |  |  |  | Double Concentrated
35 | Adult | DOBUTamine | .Double Concentrated | 1000 | mg | 500 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a |  |  |  | Double Concentrated
36 | Adult | DOBUTamine | Echo Stress Test | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 4.9 | 50 |  | 5 | [10 empty] | mcg/mL | 900 |  | 5000
37 | Adult | DOBUTamine | Stress Test | 500 | mg | 250 | mL | 2 | mg/mL | Pump | mcg/kg/min | 5 | 50 |  | 10 | [10 empty] | n/a
38 | Adult | DOPamine | ..Standard | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 0.9 | 20 | 30 | [11 empty] | n/a
39 | Adult | DOPamine | .Double Concentrated | 800 | mg | 250 | mL | 3.2 | mg/mL | Pump | mcg/kg/min | 0.9 | 20 | 30 | [11 empty] | n/a |  |  |  | Double Concentrated
40 | Adult | DOPamine | Renal Dose | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 0.9 | 5 | 6 | 1 | [10 empty] | n/a
41 | Adult | doxapram |  | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mg/min | 0.1 | 6 | [12 empty] | n/a
42 | Adult | doxapram |  | - - | mg | - - | mL |  | mg/mL | Pump | mg/min | 0.1 | 6 | [12 empty] | mg/mL | 0.89 | 0.9 | 2.1
43 | Adult | epidural w/ fentanyl |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 0.01 | 12 | 15 |  | unit | 2 | 7 | 8 | 5 | unit/min | 4 | 14 | 16 | 10 | n/a |  |  |  | 1 unit = 1 mL
44 | Adult | epidural w/o fentany |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 0.01 | 12 | 15 |  | unit | 2 | 7 | 8 | 5 | unit/min | 4 | 14 | 16 | 10 | n/a |  |  |  | 1 unit = 1 mL
45 | Adult | EPINEPHrine | ..Standard | 5 | mg | 250 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 40 | [12 empty] | n/a |  |  |  | Dedicated IV Line
46 | Adult | EPINEPHrine | ..Standard | 10 | mg | 500 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 40 | [12 empty] | n/a |  |  |  | Dedicated IV Line
47 | Adult | EPINEPHrine | Cardiac Arrest | 30 | mg | 250 | mL | 0.12 | mg/mL | Pump | mcg/min | 0.33 | 40 |  | 33.33 | [10 empty] | n/a |  |  |  | Dedicated IV Line
48 | Adult | epoprostenol |  | 1.5 | mg | 50 | mL | 0.03 | mg/mL | Syringe | nanogram/kg/min | 5 | 51 |  | 50 | [10 empty] | n/a
49 | Adult | eptifibatide | .Standard & wt>121kg | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mg/h | 13 | 17 |  | 15 | [10 empty] | n/a |  |  |  | SCr<2.0
50 | Adult | eptifibatide | .Standard & wt>121kg | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mg/h | 13 | 17 |  | 15 | [10 empty] | n/a |  |  |  | SCr<2.0
51 | Adult | eptifibatide | .Standard Dose | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mcg/kg/min | 0.5 | 2.2 | 2.5 | 2 | [10 empty] | n/a |  |  |  | SCr<2.0
52 | Adult | eptifibatide | .Standard Dose | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 2.2 | 2.5 | 2 | [10 empty] | n/a |  |  |  | SCr<2.0
53 | Adult | eptifibatide | Renal Dose | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mcg/kg/min | 0.25 | 1.2 | 1.4 | 1 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
54 | Adult | eptifibatide | Renal Dose | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.25 | 1.2 | 1.4 | 1 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
55 | Adult | eptifibatide | Renal Dose wt >121kg | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mg/h | 6 | 8 | 9 | 7.5 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
56 | Adult | eptifibatide | Renal Dose wt >121kg | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mg/h | 6 | 8 | 9 | 7.5 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
57 | Adult | esmolol |  | 2500 | mg | 250 | mL | 10 | mg/mL | Pump | mcg/kg/min | 25 | 300 |  | 50 | mcg/kg | 100 | 500 |  |  | mcg/kg/min | 100 | 500 |  |  | n/a
58 | Adult | fenoldopam |  | 10 | mg | 250 | mL | 0.04 | mg/mL | Pump | mcg/kg/min | 0.03 | 1.6 | 1.8 | 0.1 | [10 empty] | n/a
59 | Adult | fentaNYL |  | 2500 | mcg | 100 | mL | 25 | mcg/mL | Pump | mcg/h | 0.0001 | 250 | 750 |  | mcg | 0.0001 | 250 | 500 |  | mcg/min | 0.0001 | 250 | 500 |  | n/a
60 | Adult | furosemide |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mg/h | 1 | 40 | [12 empty] | n/a |  |  |  | Protect From Light
61 | Adult | glucagon |  | 10 | mg | 50 | mL | 0.2 | mg/mL | Pump | mg/h | 0.1 | 10 | [12 empty] | n/a
62 | Adult | glucagon |  | 10 | mg | 100 | mL | 0.1 | mg/mL | Pump | mg/h | 0.1 | 10 | [12 empty] | n/a
63 | Adult | heparin | ..Std. Cardiac/PVD | 25000 | unit | 500 | mL | 50 | unit/mL | Pump | unit/kg/h | 1 | 100 |  | 12 | unit/kg | 24 |  | 62 |  | unit/kg/min | 12 |  | 31 |  | n/a |  |  |  | Heparin
64 | Adult | heparin | .VTE (DVT, PE) | 25000 | unit | 500 | mL | 50 | unit/mL | Pump | unit/kg/h | 1 | 100 |  | 18 | unit/kg | 24 |  | 82 |  | unit/kg/min | 12 |  | 41 |  | n/a |  |  |  | Heparin
65 | Adult | heparin | Cerebrovascular | 25000 | unit | 500 | mL | 50 | unit/mL | Pump | unit/kg/h | 1 | 100 |  | 10 | [10 empty] | n/a |  |  |  | Heparin
66 | Adult | insulin |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 1 | 20 | 30 |  | unit | 1 | 11 |  |  | unit/min | 0.5 | 15 |  | 10 | n/a |  |  |  | 2nd RN to verify
67 | Adult | insulin |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/kg/h | 0.001 | 0.3 | 0.8 |  | unit/kg | 0.001 | 0.3 |  | 0.1 | unit/min | 0.5 | 15 |  | 10 | n/a |  |  |  | 2nd RN to verify
68 | Adult | isoproterenol |  | 1 | mg | 250 | mL | 0.004 | mg/mL | Pump | mcg/min | 0.15 | 4 | 30 | 1 | [10 empty] | n/a
69 | Adult | isoproterenol |  | 2 | mg | 500 | mL | 0.004 | mg/mL | Pump | mcg/min | 0.15 | 4 | 30 | 1 | [10 empty] | n/a
70 | Adult | ketamine ANES |  | - - | mg | - - | mL |  | mg/mL | Pump | mg/kg/h | 0.05 | 0.5 | 1 | 0.1 | [10 empty] | mg/mL | 0.1 | 0.5 | 1.5
71 | Adult | labetalol |  | 200 | mg | 200 | mL | 1 | mg/mL | Pump | mg/min | 0.05 | 2 |  | 2 | [10 empty] | n/a
72 | Adult | lidocaine |  | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | mg/min | 0.5 | 4 | 4.6 | 2 | [10 empty] | n/a
73 | Adult | mag OB maintenance |  | 20 | gram | 500 | mL | 0.04 | gram/mL | Pump | gram/h | 0.5 | 3 | 5 | [11 empty] | n/a
74 | Adult | midazolam | Anti-Convulsant | 100 | mg | 100 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.0001 | 2 |  | 0.1 | mg/kg | 0.0001 | 0.21 |  | 0.2 | mg/min | 1.9 | 2.1 |  | 2 | n/a
75 | Adult | midazolam | Anti-Convulsant | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.0001 | 2 |  | 0.1 | mg/kg | 0.0001 | 0.21 |  | 0.2 | mg/min | 1.9 | 2.1 |  | 2 | n/a
76 | Adult | midazolam | Sedation | 100 | mg | 100 | mL | 1 | mg/mL | Pump | mg/h | 0.0001 | 5 | 10 |  | mg | 0.0001 | 1 | 5 |  | mg/min | 0.0001 | 1 | 5 |  | n/a
77 | Adult | milrinone |  | 40 | mg | 200 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 0.19 | 0.75 |  | 0.5 | mcg/kg | 25 | 50.1 |  | 50 | mcg/kg/min | 1 | 5.01 |  | 5 | n/a
78 | Adult | naloxone |  | 4 | mg | 100 | mL | 0.04 | mg/mL | Pump | mg/h | 0.25 | 2 | [12 empty] | n/a
79 | Adult | nesiritide |  | 1.5 | mg | 250 | mL | 0.006 | mg/mL | Pump | mcg/kg/min | 0.005 | 0.03 | 0.035 | 0.01 | [10 empty] | n/a
80 | Adult | niCARdipine |  | 20 | mg | 200 | mL | 0.1 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
81 | Adult | niCARdipine |  | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
82 | Adult | niCARdipine |  | 40 | mg | 200 | mL | 0.2 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
83 | Adult | niCARdipine |  | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
84 | Adult | niCARdipine |  | 50 | mg | 500 | mL | 0.1 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
85 | Adult | nitroglycerin | ..Standard | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mcg/min | 5 | 200 | 400 | 5 | [10 empty] | n/a
86 | Adult | nitroglycerin | .Double Concentrated | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mcg/min | 5 | 200 | 400 | 5 | [10 empty] | n/a
87 | Adult | nitroprusside |  | 50 | mg | 500 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.1 |  | 10 | 0.3 | [10 empty] | n/a |  |  |  | Protect From Light
88 | Adult | norepinephrine | ..Standard | 4 | mg | 250 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Levophed
89 | Adult | norepinephrine | ..Standard | 8 | mg | 500 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Levophed
90 | Adult | norepinephrine | .Double Concentrated | 16 | mg | 500 | mL | 0.032 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Double Concentrated
91 | Adult | octreotide |  | 600 | mcg | 240 | mL | 2.5 | mcg/mL | Pump | mcg/h | 25 | 50 |  |  | mcg | 25 | 150 |  | 50 | mcg/min | 1 | 50 |  | 10 | n/a
92 | Adult | oxytocin |  | 30 | unit | 500 | mL | 0.06 | unit/mL | Pump | milliunit/min | 1 | 41 |  |  | unit | 0.9 | 30.1 |  | 10 | unit/min | 0.1 | 1.01 |  | 0.999 | n/a
93 | Adult | pancuronium |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 0.01 | 2 | 3 |  | mcg/kg | 0.06 | 0.1 |  |  | mcg/kg/min | 0.06 | 1 |  |  | mcg/mL | 99 | 100 | 2001 | Neuromuscular BA
94 | Adult | pantoprazole drip |  | 40 | mg | 100 | mL | 0.4 | mg/mL | Pump | mg/h | 7 | 9 | 11 | 8 | [10 empty] | n/a
95 | Adult | pantoprazole drip |  | 96 | mg | 240 | mL | 0.4 | mg/mL | Pump | mg/h | 7 | 9 | 11 | 8 | [10 empty] | n/a
96 | Adult | phenylephrine |  | 10 | mg | 250 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Phenylephrine
97 | Adult | phenylephrine |  | 20 | mg | 500 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Phenylephrine
98 | Adult | phenylephrine |  | 40 | mg | 1000 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Phenylephrine
99 | Adult | phenylephrine ANES |  | 10 | mg | 101 | mL | 0.099 | mg/mL | Pump | mcg/min | 5 | 120 | 180 | [11 empty] | n/a
100 | Adult | phenylephrine ANES |  | 10 | mg | 251 | mL | 0.0398 | mg/mL | Pump | mcg/min | 5 | 120 | 180 | [11 empty] | n/a
101 | Adult | procainamide | Loading Dose | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | mg/min | 15 | 50 |  | 33.33 | [10 empty] | n/a
102 | Adult | procainamide | Maintenance Dose | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | mg/kg/min | 0.01 | 0.06 |  | 0.05 | [10 empty] | n/a
103 | Adult | propofol |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | mcg/kg/min | 1 | 55 | 100 | 5 | [10 empty] | n/a
104 | Adult | rocuronium |  | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 4 | 12 | 16 | 8 | [10 empty] | n/a |  |  |  | Neuromuscular BA
105 | Adult | tirofiban | CrCl <= 60 | 12.5 | mg | 250 | mL | 0.05 | mg/mL | Pump | mcg/kg/min | 0.065 |  | 0.085 | 0.075 | mcg/kg | 22 |  | 28 | 25 | mcg/kg/min | 4 |  | 6 | 5 | n/a |  |  |  | Bleeding Precautions
106 | Adult | tirofiban | CrCl > 60 | 12.5 | mg | 250 | mL | 0.05 | mg/mL | Pump | mcg/kg/min | 0.13 |  | 0.17 | 0.15 | mcg/kg | 22 |  | 28 | 25 | mcg/kg/min | 4 |  | 6 | 5 | n/a |  |  |  | Bleeding Precautions
107 | Adult | vasopressin | ..Standard | 40 | unit | 100 | mL | 0.4 | unit/mL | Pump | unit/min | 0.01 | 0.05 | 0.06 | 0.04 | [10 empty] | n/a
108 | Adult | vasopressin | .Double Concentrated | 80 | unit | 100 | mL | 0.8 | unit/mL | Pump | unit/min | 0.01 | 0.05 | 0.06 | 0.04 | [10 empty] | n/a |  |  |  | Double Concentrated
109 | Adult | vecuronium |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | mcg/kg/min | 0.01 | 1.3 | 1.67 | 1 | mg | 5 | 10 |  |  | mg/min | 5 | 20 |  |  | n/a |  |  |  | Neuromuscular BA
110 | Adult | zz amiodarone DRIP |  | 360 | mg | 200 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.41 | [11 empty] | n/a |  |  |  | Code only drugs
111 | Adult | zz amiodarone DRIP |  | 900 | mg | 500 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.41 | [11 empty] | n/a |  |  |  | Code only drugs
112 | Adult | zz DOBUTamine | ..Standard | 500 | mg | 250 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a |  |  |  | Code only drugs
113 | Adult | zz DOBUTamine | .Double Concentrated | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a |  |  |  | Code only drugs
114 | Adult | zz DOPamine | ..Standard | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 0.9 | 20 | 30 | [11 empty] | n/a |  |  |  | Code only drugs
115 | Adult | zz DOPamine | .Double Concentrated | 800 | mg | 250 | mL | 3.2 | mg/mL | Pump | mcg/kg/min | 0.9 | 20 | 30 | [11 empty] | n/a |  |  |  | Code only drugs
116 | Adult | zz EPINEPHrine |  | 5 | mg | 250 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 40 | [12 empty] | n/a |  |  |  | Code only drugs
117 | Adult | zz EPINEPHrine |  | 10 | mg | 500 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 40 | [12 empty] | n/a |  |  |  | Code only drugs
118 | Adult | zz lidocaine |  | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | mg/min | 0.5 | 4 | 4.6 | 2 | [10 empty] | n/a |  |  |  | Code only drugs
119 | Adult | zz nitroglycerin | ..Standard | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mcg/min | 5 | 200 | 400 | 5 | [10 empty] | n/a |  |  |  | Code only drugs
120 | Adult | zz nitroglycerin | .Double Concentrated | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mcg/min | 5 | 200 | 400 | 5 | [10 empty] | n/a
121 | Adult | zz nitroprusside |  | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 0.1 |  | 10 | 0.3 | [10 empty] | n/a |  |  |  | Code only drugs
122 | Adult | zz norepinephrine | ..Standard | 4 | mg | 250 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Code only drugs
123 | Adult | zz norepinephrine | ..Standard | 8 | mg | 500 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Code only drugs
124 | Adult | zz norepinephrine | .Double Concentrated | 16 | mg | 500 | mL | 0.032 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Code only drugs
125 | Adult | zz phenylephrine |  | 10 | mg | 250 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Code only drugs
126 | Adult | zz phenylephrine |  | 20 | mg | 500 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Code only drugs
127 | Adult | zz phenylephrine |  | 40 | mg | 1000 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Code only drugs
128 | Nursery | No drugs exist in this section
129 | ONC Infusion Center | No drugs exist in this section
130 | OutPt InfusionCenter | No drugs exist in this section
131 | Pediatrics | alprostadil |  | 500 | mcg | 250 | mL | 2 | mcg/mL | Pump | mcg/kg/min | 0.05 | 0.1 | 0.4 | 0.05 | [10 empty] | n/a
132 | Pediatrics | alprostadil |  | - - | mcg | - - | mL |  | mcg/mL | Pump | mcg/kg/min | 0.05 | 0.1 | 0.4 | 0.05 | [10 empty] | mcg/mL | 0.5 | 1 | 20
133 | Pediatrics | aminophylline |  | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.2 | 1.1 |  |  | mg/kg | 0.001 |  | 6 |  | mg/kg/min | 0.19 |  | 0.36 | 0.2 | n/a
134 | Pediatrics | DOBUTamine |  | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | 2 | [10 empty] | n/a
135 | Pediatrics | DOPamine |  | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 2 |  | 20 | 2 | [10 empty] | n/a
136 | Pediatrics | EPINEPHrine |  | 5 | mg | 250 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 20 | 30 | [11 empty] | n/a
137 | Pediatrics | insulin |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 1 | 20 | 30 |  | unit | 1 | 11 |  |  | unit/min | 0.5 | 15 |  | 10 | n/a |  |  |  | 2nd RN to verify
138 | Pediatrics | insulin |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/kg/h | 0.001 | 0.3 | 0.8 |  | unit/kg | 0.001 | 0.3 |  | 0.1 | unit/min | 0.5 | 15 |  | 10 | n/a |  |  |  | 2nd RN to verify
139 | Pediatrics | lidocaine |  | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | mg/min | 0.5 | 4 | 4.6 | 2 | [10 empty] | n/a
140 | Pediatrics | milrinone |  | 20 | mg | 100 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 0.125 | 0.75 |  | 0.5 | mcg/kg | 25 | 50 |  |  | mcg/kg/min | 1 | 5 |  |  | n/a
141 | Pediatrics | naloxone |  | 2 | mg | 500 | mL | 0.004 | mg/mL | Pump | mg/h | 0.25 | 2 | [12 empty] | n/a
142 | Pediatrics | niCARdipine |  | 20 | mg | 200 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.01 | 4 | 5 | 0.5 | mcg/kg | 29 |  | 50 | 30 | mcg/kg/min | 29 |  | 50 | 30 | n/a
143 | Pediatrics | niCARdipine |  | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.01 | 4 | 5 | 0.5 | mcg/kg | 29 |  | 50 | 30 | mcg/kg/min | 29 |  | 50 | 30 | n/a
144 | Pediatrics | niCARdipine |  | 40 | mg | 200 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 0.01 | 4 | 5 | 0.5 | mcg/kg | 29 |  | 50 | 30 | mcg/kg/min | 29 |  | 50 | 30 | n/a
145 | Pediatrics | niCARdipine |  | 50 | mg | 500 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.01 | 4 | 5 | 0.5 | mcg/kg | 29 |  | 50 | 30 | mcg/kg/min | 29 |  | 50 | 30 | n/a
146 | Pediatrics | nitroglycerin |  | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mcg/min | 5 | 150 | 400 | 5 | [10 empty] | n/a
147 | Pediatrics | nitroprusside |  | 50 | mg | 500 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.1 | 8 | 8.1 | 1 | [10 empty] | n/a |  |  |  | Protect From Light
148 | Pediatrics | norepinephrine |  | 4 | mg | 250 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 45 | [12 empty] | n/a
149 | Pediatrics | norepinephrine |  | 8 | mg | 500 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 45 | [12 empty] | n/a
150 | Pediatrics | propofol |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | mcg/kg/min | 1 | 155 | 310 | [11 empty] | n/a
Note on AE4: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on AE5: Dose and volume to be infused need to be set according to weight based protocol to avoid potential over infusion of entire contents of bag.
Note on AE7: Infusion requires the use of an in-line filter.
Note on AE8: Infusion requires the use of an in-line filter.
Note on AE9: Infusion requires the use of an in-line filter.
Note on J16: Units Only concentration and Dosing Units have different units of measure
Note on AA16: Units Only concentration and Dosing Units have different units of measure
Note on AE16: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE25: Protect this drug from light.
Note on AE27: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE36: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on AE37: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on J38: Units Only concentration and Dosing Units have different units of measure
Note on AA38: Units Only concentration and Dosing Units have different units of measure
Note on AE41: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on AE45: 1 unit = 1 mL
Note on AE46: 1 unit = 1 mL
Note on AE47: Do not infuse any other drugs through this line.
Note on AE48: Do not infuse any other drugs through this line.
Note on AE49: Do not infuse any other drugs through this line.
Note on AE51: SCr < 2.0
Note on AE52: SCr < 2.0
Note on AE53: SCr < 2.0
Note on AE54: SCr < 2.0
Note on AE55: SCr 2.0 - 4.0
Note on AE56: SCr 2.0 - 4.0
Note on AE57: SCr 2.0 - 4.0
Note on AE58: SCr 2.0 - 4.0
Note on AE62: Protect this drug from light.
Note on AE65: Monitor PTT. Do not administer other drugs through same IV line.
Note on AE66: Monitor PTT. Do not administer other drugs through same IV line.
Note on AE67: Monitor PTT. Do not administer other drugs through same IV line.
Note on AE68: 2 RNs to verify and document concentration and programming prior to infusion.
Note on AE69: 2 RNs to verify and document concentration and programming prior to infusion.
Note on AE89: Protect this drug from light.
Note on AE90: *****Must have physician order for rates greater than 60 mcg/min*****
Note on AE91: *****Must have physician order for rates greater than 60 mcg/min*****
Note on AE92: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on J95: Units Only concentration and Dosing Units have different units of measure
Note on AA95: Units Only concentration and Dosing Units have different units of measure
Note on AE95: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE98: *****Must have physician order for rates greater than 300 mcg/min*****
Note on AE99: *****Must have physician order for rates greater than 300 mcg/min*****
Note on AE100: *****Must have physician order for rates greater than 300 mcg/min*****
Note on AE106: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE107: Monitor for signs of bleeding.
Note on AE108: Monitor for signs of bleeding.
Note on AE110: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on AE111: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE112: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE113: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE114: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE115: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE116: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE117: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE118: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE119: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE120: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE121: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE123: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE124: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE125: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE126: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE127: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE128: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE129: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE139: 2 RNs to verify and document concentration and programming prior to infusion.
Note on AE140: 2 RNs to verify and document concentration and programming prior to infusion.
Note on AE149: Protect this drug from light.

## ContBolus Anesth Drugs
Continuous/Bolus Anesthesia Only Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount (- - = custom concentration) | Drug Amount Units | Diluent Volume (- - = custom concentration) | Diluent Volume Units | Final Concentration | Final Concentration Units | Module Availability | Dosing Units | Dose Limits Soft Min | Dose Limits Soft Max | Dose Limits Hard Max | Dose Initial Value | Bolus Dose Units | Bolus Dose Limits Soft Min | Bolus Dose Limits Soft Max | Bolus Dose Limits Hard Max | Bolus Dose Limits Initial Value | Bolus Dose Admin. Rate Units | Bolus Dose Admin. Rate Limits Soft Min | Bolus Dose Admin. Rate Limits Soft Max | Bolus Dose Admin. Rate Limits Hard Max | Bolus Dose Admin. Rate Initial Value | Conc. Limit Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max
1 | Adult | ketamine ANES |  | 50 | mg | 50 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.05 | 0.5 | 1 | 0.1 | [10 empty] | n/a
2 | Adult | ketamine ANES |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.05 | 0.5 | 1 | 0.1 | [10 empty] | n/a
3 | Adult | phenylephrine ANES |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/min | 5 | 120 | 180 | [11 empty] | mcg/mL | 10 | 38 | 110
4 | Adult | propofol ANES |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | mcg/kg/min | 1 |  | 300 | 5 | mg | 10 | 70 |  | 30 | mg/min | 50 | 140 |  | 120 | n/a
5 | Adult | propofol ANES |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | mcg/kg/min | 1 |  | 300 | 5 | mg | 10 | 70 |  | 30 | mg/min | 50 | 140 |  | 120 | n/a
6 | Adult | remifentanil ANES |  | 1000 | mcg | 40 | mL | 25 | mcg/mL | Pump | mcg/kg/min | 0.025 | 2 | 4 | 0.05 | mcg/kg | 0.1 |  | 2 | 1 | mcg/min | 1 | 2000 | 4000 | 1000 | n/a
7 | Adult | remifentanil ANES |  | 2000 | mcg | 80 | mL | 25 | mcg/mL | Pump | mcg/kg/min | 0.025 | 2 | 4 | 0.05 | mcg/kg | 0.1 |  | 2 | 1 | mcg/min | 1 | 2000 | 4000 | 1000 | n/a
8 | Adult | remifentanil ANES |  | - - | mcg | - - | mL |  | mcg/mL | Pump | mcg/kg/min | 0.025 | 2 | 4 | 0.05 | mcg/kg | 0.1 |  | 2 | 1 | mcg/min | 1 | 2000 | 4000 | 1000 | mcg/mL | 8.9 | 9 | 250
9 | Adult | rocuronium ANES |  | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 4 | 12 | 16 | 8 | [10 empty] | n/a
10 | Adult | rocuronium ANES |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 4 | 12 | 16 | 8 | [10 empty] | mcg/mL | 890 | 900 | 1100
11 | Adult | SUFentanil ANES |  | 100 | mcg | 100 | mL | 1 | mcg/mL | Pump | mcg/kg/h | 0.1 | 1 | [12 empty] | n/a
12 | Adult | SUFentanil ANES |  | - - | mcg | - - | mL |  | mcg/mL | Pump | mcg/kg/h | 0.1 | 1 | [12 empty] | mcg/mL | 0.1 | 0.2 | 4
13 | Adult | vecuronium |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 0.01 | 1.3 | 1.67 | 1 | mg | 5 | 10 |  |  | mg/min | 5 | 20 |  |  | mcg/mL | 440 | 450 | 550
14 | Nursery | No drugs exist in this section
15 | ONC Infusion Center | No drugs exist in this section
16 | OutPt InfusionCenter | No drugs exist in this section
17 | Pediatrics | No drugs exist in this section
Note on J5: Units Only concentration and Dosing Units have different units of measure
Note on AA5: Units Only concentration and Dosing Units have different units of measure
Note on J12: Units Only concentration and Dosing Units have different units of measure
Note on AA12: Units Only concentration and Dosing Units have different units of measure
Note on J15: Units Only concentration and Dosing Units have different units of measure
Note on AA15: Units Only concentration and Dosing Units have different units of measure

## Intermittent Drugs
Intermittent Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount (- - = custom concentration) | Drug Amount Units | Diluent Volume (- - = custom concentration) | Diluent Volume Units | Final Concentration | Final Concentration Units | Module Availability | Available As | Duration Limits (hh:mm) Hard Min | Duration Limits (hh:mm) Soft Min | Duration Limits (hh:mm) Soft Max | Duration (hh:mm) Initial Value | Dosing Units | Dose Limits Soft Min | Dose Limits Soft Max | Dose Limits Hard Max | Conc. Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max | Clinical Advisory
1 | Adult | acetaminophen | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 500 | 1000 |  | n/a
2 | Adult | acetaminophen | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | gram | 0.5 | 1 |  | n/a
3 | Adult | acetaZOLAMIDE |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 00:40 | 00:30 | mg | 100 | 1000 |  | n/a
4 | Adult | acetylcysteine | 1st Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg/kg | 142.5 |  | 157.5 | mg/mL | 1 |  | 75
5 | Adult | acetylcysteine | 1st Dose >100kg | 15000 | mg | 275 | mL | 54.5455 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 13500 | 16500 | 18000 | n/a
6 | Adult | acetylcysteine | 2nd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg/kg | 47.5 |  | 52.5 | mg/mL | 1 |  | 10
7 | Adult | acetylcysteine | 2nd Dose >100kg | 5000 | mg | 525 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg | 4000 | 5500 | 6500 | n/a
8 | Adult | acetylcysteine | 3rd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg/kg | 95 |  | 105 | mg/mL | 1 |  | 10
9 | Adult | acetylcysteine | 3rd Dose >100kg | 10000 | mg | 1050 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg | 9000 | 11000 | 12500 | n/a
10 | Adult | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 100 | 2000 |  | mg/mL | 0.1 | 1 | 7
11 | Adult | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg/kg | 1 | 10.1 |  | mg/mL | 0.1 | 1 | 7
12 | Adult | albumin 25% |  | 5 | gram | 20 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:08 | 01:00 | 00:20 | gram | 4.9 | 25.1 |  | n/a
13 | Adult | albumin 25% |  | 12.5 | gram | 50 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:16 | 02:00 | 00:50 | gram | 4.9 | 25.1 |  | n/a
14 | Adult | albumin 25% |  | 25 | gram | 100 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:32 | 04:00 | 01:40 | gram | 4.9 | 25.1 |  | n/a
15 | Adult | alteplase-non-stroke | Central IV Occlusion | 2 | mg | 50 | mL | 0.04 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:00 | 00:15 | mg | 2 | 4 |  | n/a
16 | Adult | alteplase-non-stroke | Central IV Occlusion | 2 | mg | 100 | mL | 0.02 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:00 | 00:15 | mg | 2 | 4 |  | n/a
17 | Adult | alteplase-non-stroke | DVT / Ischemic Limb | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:00 |  | 24:00 |  | mg | 0.9 |  | 24 | mg/mL | 0.004 | 0.008 | 0.51
18 | Adult | alteplase-non-stroke | Pulmonary Embolism | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary Only |  | 01:45 | 02:15 | 02:00 | mg | 49 | 115 | 130 | n/a
19 | Adult | amikacin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 |  | mg/kg | 1 | 21 | 39 | mg/mL | 0.2 |  | 5.1
20 | Adult | aminocaproic acid | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 500 | 5000 |  | mg/mL | 1 | 5 | 50
21 | Adult | aminocaproic acid | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | gram | 0.5 | 5.7 |  | gram/mL | 0.001 | 0.005 | 0.05
22 | Adult | amiodarone *Loading* |  | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:08 | 00:15 | 00:10 | mg | 149 | 151 |  | n/a
23 | Adult | amiodarone *Loading* |  | 150 | mg | 103 | mL | 1.4563 | mg/mL | Pump | Primary and Secondary |  | 00:08 | 00:15 | 00:10 | mg | 149 | 151 |  | n/a
24 | Adult | ampho B (liposomal) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only |  | 00:55 | 04:00 | 02:00 | mg/kg | 1.9 | 10.5 |  | mg/mL | 0.8 | 0.9 | 2.1 | AmBisome
25 | Adult | amphotericin B |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 02:00 | 03:00 | 10:00 | 06:00 | mg/kg | 0.05 | 1 | 1.5 | mg/mL | 0.01 |  | 0.1
26 | Adult | ampicillin | .MG Dosing | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
27 | Adult | ampicillin | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
28 | Adult | ampicillin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
29 | Adult | ampicillin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
30 | Adult | ampicillin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
31 | Adult | ampicillin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
32 | Adult | ampicillin | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.1 | 2.4 | n/a
33 | Adult | ampicillin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.1 | 2.4 | n/a
34 | Adult | ampicillin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.1 | 2.4 | n/a
35 | Adult | ampicillin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.1 | 2.4 | n/a
36 | Adult | ampicillin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.1 | 2.4 | n/a
37 | Adult | ampicillin/sulbact | .MG Dosing | 1500 | mg | 50 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | mg | 1200 | 3400 |  | n/a
38 | Adult | ampicillin/sulbact | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | mg | 1200 | 3400 |  | n/a
39 | Adult | ampicillin/sulbact | .MG Dosing | 3000 | mg | 100 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 04:00 | 01:00 | mg | 1200 | 3400 |  | n/a
40 | Adult | ampicillin/sulbact | GM dosing | 1.5 | gram | 50 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | gram | 1.2 | 3.4 |  | n/a
41 | Adult | ampicillin/sulbact | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | gram | 1.2 | 3.4 |  | n/a
42 | Adult | ampicillin/sulbact | GM dosing | 3 | gram | 100 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 04:00 | 01:00 | gram | 1.2 | 3.4 |  | n/a
43 | Adult | andulafungin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 04:00 | 02:00 | mg | 50 | 200 |  | mg/mL | 0.1 | 0.35 | 0.5
44 | Adult | antivenin crotalidae | Anavip | 4 | unit | 250 | mL | 0.016 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 10.1 | 11 | n/a |  |  |  | Antivenin
45 | Adult | antivenin crotalidae | Anavip | 10 | unit | 250 | mL | 0.04 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 10.1 | 11 | n/a |  |  |  | Antivenin
46 | Adult | antivenin crotalidae | Anavip | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 10.1 | 11 | unit/mL | 0.003 |  | 0.045 | Antivenin
47 | Adult | antivenin crotalidae | Crofab | 4 | unit | 250 | mL | 0.016 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 6.1 | 13 | n/a |  |  |  | Antivenin
48 | Adult | antivenin crotalidae | Crofab | 6 | unit | 250 | mL | 0.024 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 6.1 | 13 | n/a |  |  |  | Antivenin
49 | Adult | antivenin crotalidae | Crofab | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 6.1 | 13 | unit/mL | 0.004 |  | 0.025 | Antivenin
50 | Adult | avibactam/ceftazidim |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 01:42 | 03:20 | 02:00 | gram | 0.93 |  | 2.8 | gram/mL | 0.009 |  | 0.051
51 | Adult | azithromycin |  | 250 | mg | 250 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:10 | 01:00 | mg | 210 | 575 |  | n/a
52 | Adult | azithromycin |  | 500 | mg | 250 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:10 | 01:00 | mg | 210 | 575 |  | n/a
53 | Adult | aztreonam | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
54 | Adult | aztreonam | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
55 | Adult | aztreonam | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg | 250 | 2000 |  | n/a
56 | Adult | aztreonam | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg | 250 | 2000 |  | n/a
57 | Adult | aztreonam | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.25 | 2 |  | n/a
58 | Adult | aztreonam | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.25 | 2 |  | n/a
59 | Adult | aztreonam | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | gram | 0.25 | 2 |  | n/a
60 | Adult | aztreonam | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 02:00 | gram | 0.25 | 2 |  | n/a
61 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 110 | mL | 6.3636 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:20 | 06:00 | 00:22 | mg | 699 | 701 |  | n/a
62 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 160 | mL | 4.375 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 06:00 | 00:31 | mg | 699 | 701 |  | n/a
63 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 210 | mL | 3.3333 | mg/mL | Pump | Primary and Secondary | 00:36 | 00:37 | 06:00 | 00:41 | mg | 699 | 701 |  | n/a
64 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 310 | mL | 2.2581 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 06:00 | 01:00 | mg | 699 | 701 |  | n/a
65 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 110 | mL | 6.3636 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:20 | 06:00 | 00:22 | mg | 699 | 701 |  | n/a
66 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 160 | mL | 4.375 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 06:00 | 00:31 | mg | 699 | 701 |  | n/a
67 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 210 | mL | 3.3333 | mg/mL | Pump | Primary and Secondary | 00:36 | 00:37 | 06:00 | 00:41 | mg | 699 | 701 |  | n/a
68 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 310 | mL | 2.2581 | mg/mL | Pump | Primary and Secondary | 00:59 | 01:00 | 06:00 | 01:10 | mg | 699 | 701 |  | n/a
69 | Adult | bamlanivimab |  | 700 | mg | 270 | mL | 2.5926 | mg/mL | Pump | Primary and Secondary | 01:00 |  | 10:00 | 01:00 | mg | 650 | 750 |  | n/a |  |  |  | Low Binding Filter
70 | Adult | caffeine citrate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 06:00 |  | mg | 1 | 1200 |  | mg/mL |  | 0.01 | 10
71 | Adult | caffeine/sod benzoa |  | 500 | mg | 1000 | mL | 0.5 | mg/mL | Pump | Primary and Secondary | 01:01 |  | 08:00 | 01:02 | mg | 50 | 501 | 1001 | n/a |  |  |  | 0.22 Filter Needed
72 | Adult | caffeine/sod benzoa |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 04:00 | 01:00 | mg | 50 | 501 | 1001 | mg/mL |  | 0.24 | 6 | 0.22 Filter Needed
73 | Adult | calcium chloride | .MG Dosing | 1000 | mg | 60 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | mg | 100 | 2000 |  | n/a
74 | Adult | calcium chloride | .MG Dosing | 1000 | mg | 110 | mL | 9.0909 | mg/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | mg | 100 | 2000 |  | n/a
75 | Adult | calcium chloride | .MG Dosing | 2000 | mg | 120 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 | 02:00 | mg | 100 | 2000 |  | n/a
76 | Adult | calcium chloride | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 |  | mg | 100 | 2000 |  | mg/mL |  | 0.1 | 20
77 | Adult | calcium chloride | GM dosing | 1 | gram | 60 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | gram | 0.1 | 10 |  | n/a
78 | Adult | calcium chloride | GM dosing | 1 | gram | 110 | mL | 0.0091 | gram/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | gram | 0.1 | 10 |  | n/a
79 | Adult | calcium chloride | GM dosing | 2 | gram | 120 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 | 02:00 | gram | 0.1 | 10 |  | n/a
80 | Adult | calcium chloride | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 |  | gram | 0.1 | 10 |  | gram/mL |  | 0.001 | 0.04
81 | Adult | calcium gluconate | .MG Dosing | 1000 | mg | 60 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | mg | 500 | 2400 |  | n/a
82 | Adult | calcium gluconate | .MG Dosing | 1000 | mg | 110 | mL | 9.0909 | mg/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | mg | 500 | 2400 |  | n/a
83 | Adult | calcium gluconate | .MG Dosing | 2000 | mg | 70 | mL | 28.5714 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | mg | 500 | 2400 |  | n/a
84 | Adult | calcium gluconate | .MG Dosing | 2000 | mg | 120 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | mg | 500 | 2400 |  | n/a
85 | Adult | calcium gluconate | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:25 | 06:00 |  | mg | 500 | 2400 |  | mg/mL | 1 |  | 40
86 | Adult | calcium gluconate | GM dosing | 1 | gram | 60 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
87 | Adult | calcium gluconate | GM dosing | 1 | gram | 110 | mL | 0.0091 | gram/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
88 | Adult | calcium gluconate | GM dosing | 2 | gram | 70 | mL | 0.0286 | gram/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
89 | Adult | calcium gluconate | GM dosing | 2 | gram | 120 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
90 | Adult | calcium gluconate | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 | 00:25 | 06:00 |  | gram | 0.5 | 2.4 |  | gram/mL | 0.001 |  | 0.04
91 | Adult | casirivimab/imdevima |  | 1200 | mg | 260 | mL | 4.6154 | mg/mL | Pump | Primary and Secondary | 00:48 | 00:49 | 04:00 | 01:00 | mg | 1190 | 1201 | 1210 | n/a |  |  |  | Low Binding Filter
92 | Adult | casirivimab/imdevima |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:18 | 00:19 | 04:00 |  | mg | 1190 | 1201 | 1210 | No |  |  |  | Low Binding Filter
93 | Adult | caspofungin |  | 35 | mg | 100 | mL | 0.35 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg | 35 | 90 |  | n/a |  |  |  | Dedicated IV Line
94 | Adult | caspofungin |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg | 35 | 90 |  | n/a |  |  |  | Dedicated IV Line
95 | Adult | caspofungin |  | 70 | mg | 250 | mL | 0.28 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg | 35 | 90 |  | n/a |  |  |  | Dedicated IV Line
96 | Adult | ceFAZolin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 |  | mg | 480 | 3100 |  | n/a
97 | Adult | ceFAZolin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 3100 |  | n/a
98 | Adult | ceFAZolin | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 3100 |  | n/a
99 | Adult | ceFAZolin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 3100 |  | n/a
100 | Adult | ceFAZolin | .MG Dosing | 3000 | mg | 100 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 3100 |  | n/a
101 | Adult | ceFAZolin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 3.1 |  | n/a
102 | Adult | ceFAZolin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 3.1 |  | n/a
103 | Adult | ceFAZolin | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 3.1 |  | n/a
104 | Adult | ceFAZolin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 3.1 |  | n/a
105 | Adult | ceFAZolin | GM dosing | 3 | gram | 100 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 3.1 |  | n/a
106 | Adult | cefepime | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
107 | Adult | cefepime | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
108 | Adult | cefepime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
109 | Adult | cefepime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
110 | Adult | cefepime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
111 | Adult | cefepime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
112 | Adult | cefepime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
113 | Adult | cefepime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
114 | Adult | cefepime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
115 | Adult | cefepime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
116 | Adult | cefiderocol | .MG Dosing | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
117 | Adult | cefiderocol | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
118 | Adult | cefiderocol | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
119 | Adult | cefiderocol | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
120 | Adult | cefiderocol | GM dosing | 0.75 | gram | 100 | mL | 0.0075 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
121 | Adult | cefiderocol | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
122 | Adult | cefiderocol | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
123 | Adult | cefiderocol | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
124 | Adult | cefotaxime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
125 | Adult | cefotaxime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
126 | Adult | cefotaxime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
127 | Adult | cefotaxime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
128 | Adult | cefotaxime | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 40
129 | Adult | cefotaxime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
130 | Adult | cefotaxime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
131 | Adult | cefotaxime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
132 | Adult | cefotaxime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
133 | Adult | cefotaxime | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.04
134 | Adult | cefoTEtan | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
135 | Adult | cefoTEtan | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
136 | Adult | cefoTEtan | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
137 | Adult | cefoTEtan | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
138 | Adult | cefoTEtan | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
139 | Adult | cefoTEtan | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
140 | Adult | cefoTEtan | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
141 | Adult | cefoTEtan | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
142 | Adult | cefOXitin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
143 | Adult | cefOXitin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
144 | Adult | cefOXitin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 20
145 | Adult | cefOXitin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
146 | Adult | cefOXitin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
147 | Adult | cefOXitin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.02
148 | Adult | ceftaroline |  | 200 | mg | 50 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
149 | Adult | ceftaroline |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
150 | Adult | ceftaroline |  | 300 | mg | 50 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
151 | Adult | ceftaroline |  | 300 | mg | 100 | mL | 3 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
152 | Adult | ceftaroline |  | 400 | mg | 50 | mL | 8 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
153 | Adult | ceftaroline |  | 400 | mg | 100 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
154 | Adult | ceftaroline |  | 600 | mg | 50 | mL | 12 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
155 | Adult | ceftaroline |  | 600 | mg | 100 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
156 | Adult | ceftaroline |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:07 | 04:00 | 01:00 | mg | 200 | 600 |  | mg/mL |  | 0.1 | 14
157 | Adult | cefTAZidime | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
158 | Adult | cefTAZidime | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
159 | Adult | cefTAZidime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
160 | Adult | cefTAZidime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
161 | Adult | cefTAZidime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
162 | Adult | cefTAZidime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
163 | Adult | cefTAZidime | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
164 | Adult | cefTAZidime | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
165 | Adult | cefTAZidime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
166 | Adult | cefTAZidime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
167 | Adult | cefTAZidime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
168 | Adult | cefTAZidime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
169 | Adult | ceftolozane/tazo | .MG Dosing | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 149 | 3010 |  | n/a
170 | Adult | ceftolozane/tazo | .MG Dosing | 375 | mg | 100 | mL | 3.75 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 149 | 3010 |  | n/a
171 | Adult | ceftolozane/tazo | .MG Dosing | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 149 | 3010 |  | n/a
172 | Adult | ceftolozane/tazo | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 149 | 3010 |  | n/a
173 | Adult | ceftolozane/tazo | .MG Dosing | 3000 | mg | 100 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 06:00 | 03:00 | mg | 149 | 3010 |  | n/a
174 | Adult | ceftolozane/tazo | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 149 | 3010 |  | mg/mL | 1 | 1.4 | 30.1
175 | Adult | ceftolozane/tazo | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | gram | 0.149 | 3.1 |  | n/a
176 | Adult | ceftolozane/tazo | GM dosing | 3 | gram | 100 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 06:00 | 03:00 | gram | 0.149 | 3.1 |  | n/a
177 | Adult | ceftolozane/tazo | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | gram | 0.149 | 3.1 |  | gram/mL | 0.001 |  | 0.031
178 | Adult | cefTRIAXone | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
179 | Adult | cefTRIAXone | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
180 | Adult | cefTRIAXone | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
181 | Adult | cefTRIAXone | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
182 | Adult | cefTRIAXone | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | mg/mL | 4.9 | 9 | 40
183 | Adult | cefTRIAXone | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
184 | Adult | cefTRIAXone | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
185 | Adult | cefTRIAXone | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
186 | Adult | cefTRIAXone | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
187 | Adult | cefTRIAXone | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | gram/mL | 0.004 | 0.009 | 0.04
188 | Adult | cefuroxime | .MG Dosing | 750 | mg | 50 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
189 | Adult | cefuroxime | .MG Dosing | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
190 | Adult | cefuroxime | .MG Dosing | 1500 | mg | 50 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
191 | Adult | cefuroxime | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
192 | Adult | cefuroxime | GM dosing | 0.75 | gram | 50 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
193 | Adult | cefuroxime | GM dosing | 0.75 | gram | 100 | mL | 0.0075 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
194 | Adult | cefuroxime | GM dosing | 1.5 | gram | 50 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
195 | Adult | cefuroxime | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
196 | Adult | chloramphenicol |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg/kg | 3 | 25 | 30 | mg/mL | 5 | 10 | 110
197 | Adult | chlorothiazide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 1000 | 2300 | mg/mL | 1 | 2 | 20
198 | Adult | chlorproMAZINE |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 02:00 | 00:15 | mg | 1 | 51 |  | mg/mL |  | 0.001 | 1.1 | Patient Lie Down
199 | Adult | ciprofloxacin |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
200 | Adult | ciprofloxacin |  | 400 | mg | 200 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
201 | Adult | clindamycin |  | 300 | mg | 50 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
202 | Adult | clindamycin |  | 300 | mg | 100 | mL | 3 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
203 | Adult | clindamycin |  | 600 | mg | 50 | mL | 12 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
204 | Adult | clindamycin |  | 600 | mg | 100 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
205 | Adult | clindamycin |  | 900 | mg | 50 | mL | 18 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
206 | Adult | clindamycin |  | 900 | mg | 100 | mL | 9 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
207 | Adult | colistimethate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 10 | 300 |  | mg/mL | 0.1 | 0.5 | 5
208 | Adult | colistimethate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 0.3 |  | 5 | mg/mL | 0.1 | 0.5 | 5
209 | Adult | conivaptan Loading |  | 20 | mg | 100 | mL | 0.2 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 15 | 25 |  | n/a
210 | Adult | dalbavancin | .MG Dosing | 375 | mg | 250 | mL | 1.5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
211 | Adult | dalbavancin | .MG Dosing | 500 | mg | 250 | mL | 2 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
212 | Adult | dalbavancin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
213 | Adult | dalbavancin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
214 | Adult | dalbavancin | .MG Dosing | 1125 | mg | 250 | mL | 4.5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
215 | Adult | dalbavancin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
216 | Adult | dalbavancin | GM dosing | 0.375 | gram | 250 | mL | 0.0015 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
217 | Adult | dalbavancin | GM dosing | 0.5 | gram | 250 | mL | 0.002 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
218 | Adult | dalbavancin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
219 | Adult | dalbavancin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
220 | Adult | dalbavancin | GM dosing | 1.125 | gram | 250 | mL | 0.0045 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
221 | Adult | dalbavancin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
222 | Adult | dantrolene |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg/kg | 1 | 2.5 |  | mg/mL | 0.05 | 0.1 | 100
223 | Adult | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg | 100 | 2000 |  | mg/mL | 1 | 2 | 20
224 | Adult | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg/kg | 4 | 12.1 |  | mg/mL | 1 | 2 | 20
225 | Adult | deferoxamine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 04:00 | 08:00 | 24:00 |  | mg | 500 | 6000 |  | mg/mL | 1 |  | 95
226 | Adult | deferoxamine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 04:00 | 08:00 | 24:00 |  | mg/kg | 30 | 60 |  | mg/mL | 1 |  | 95
227 | Adult | desmopressin |  | - - | mcg | - - | mL |  | mcg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mcg | 0.1 | 22 |  | mcg/mL | 0.005 | 0.01 | 4
228 | Adult | desmopressin |  | - - | mcg | - - | mL |  | mcg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mcg/kg | 0.25 | 0.41 |  | mcg/mL | 0.005 | 0.01 | 4
229 | Adult | dexamethasone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:07 | 02:00 | 00:30 | mg | 1 | 20 |  | mg/mL | 0.01 |  | 10
230 | Adult | digoxin immune fab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 10 | 800 |  | mg/mL | 0.05 | 0.1 | 1
231 | Adult | doripenem |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 100 | 500 |  | n/a
232 | Adult | doripenem |  | 250 | mg | 100 | mL | 2.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 100 | 500 |  | n/a
233 | Adult | doripenem |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 100 | 500 |  | n/a
234 | Adult | doxycycline |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
235 | Adult | doxycycline |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
236 | Adult | ertapenem | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
237 | Adult | ertapenem | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
238 | Adult | ertapenem | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
239 | Adult | ertapenem | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
240 | Adult | ertapenem | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
241 | Adult | ertapenem | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
242 | Adult | ertapenem | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
243 | Adult | ertapenem | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
244 | Adult | erythromycin |  | 250 | mg | 100 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
245 | Adult | erythromycin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
246 | Adult | erythromycin |  | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:30 | 01:00 | mg | 100 | 1000 |  | n/a
247 | Adult | erythromycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | mg/mL | 0.05 | 1 | 5
248 | Adult | famotidine |  | 20 | mg | 50 | mL | 0.4 | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg | 8 | 44 | 50 | n/a
249 | Adult | famotidine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 02:00 |  | mg | 8 | 44 | 50 | mg/mL | 0.09 | 0.1 | 3.5
250 | Adult | fat emulsion 20% |  | 25 | gram | 125 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 11:00 | 24:00 |  | gram | 24 | 51 |  | n/a
251 | Adult | fat emulsion 20% |  | 50 | gram | 250 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 11:00 | 24:00 |  | gram | 24 | 51 |  | n/a
252 | Adult | ferric carboxymaltos |  | 750 | mg | 265 | mL | 2.8302 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg | 250 | 750 | 1501 | n/a |  |  |  | Hypersensitivity
253 | Adult | ferric carboxymaltos |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg | 250 | 750 | 1501 | mg/mL | 1.9 | 2 | 4 | Hypersensitivity
254 | Adult | ferric carboxymaltos |  | 750 | mg | 265 | mL | 2.8302 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg/kg | 5 | 16 |  | n/a |  |  |  | Hypersensitivity
255 | Adult | ferric carboxymaltos |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg/kg | 5 | 16 |  | mg/mL | 1.9 | 2 | 4 | Hypersensitivity
256 | Adult | ferumoxytol | .MG Dosing | 510 | mg | 117 | mL | 4.359 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 509 | 1021 | 1022 | n/a |  |  |  | Hypersensitivity
257 | Adult | ferumoxytol | .MG Dosing | 1020 | mg | 284 | mL | 3.5915 | mg/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | mg | 509 | 1021 | 1022 | n/a |  |  |  | Hypersensitivity
258 | Adult | ferumoxytol | GM dosing | 0.51 | gram | 117 | mL | 0.0044 | gram/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | gram | 0.509 | 1.021 | 1.022 | n/a |  |  |  | Hypersensitivity
259 | Adult | ferumoxytol | GM dosing | 1.02 | gram | 284 | mL | 0.0036 | gram/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | gram | 0.509 | 1.021 | 1.022 | n/a |  |  |  | Hypersensitivity
260 | Adult | fluconazole |  | 100 | mg | 50 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 04:00 | 00:30 | mg | 50 | 801 |  | n/a
261 | Adult | fluconazole |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 04:00 | 01:00 | mg | 50 | 801 |  | n/a
262 | Adult | fluconazole |  | 400 | mg | 200 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 01:00 | 04:00 | 02:00 | mg | 50 | 801 |  | n/a
263 | Adult | fluconazole |  | 800 | mg | 400 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 02:00 | 08:00 | 04:00 | mg | 50 | 801 |  | n/a
264 | Adult | folic acid |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 0.1 | 5 |  | mg/mL | 0.001 | 0.009 | 0.06
265 | Adult | fomepizole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 9 | 16 |  | mg/mL | 0.5 | 1 | 20
266 | Adult | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:11 | 01:00 | 00:20 | mg | 50 | 1500 |  | mg/mL | 0.5 | 1 | 30
267 | Adult | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:11 | 01:00 | 00:20 | mg/kg | 1 | 20 |  | mg/mL | 0.5 | 1 | 30
268 | Adult | ganciclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg/kg | 0.625 | 8 |  | mg/mL | 0.5 | 1 | 10 | In-line filter
269 | Adult | gentamicin |  | 80 | mg | 100 | mL | 0.8 | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 | 00:30 | mg | 10 | 1000 |  | n/a
270 | Adult | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 | 01:00 | mg | 10 | 1000 |  | mg/mL | 0.05 | 0.1 | 10
271 | Adult | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 |  | 01:20 | 01:00 | mg/kg | 0.5 | 10 |  | mg/mL | 0.05 | 0.1 | 10
272 | Adult | hetastarch 6% |  | 30 | gram | 500 | mL | 0.06 | gram/mL | Pump | Primary and Secondary |  | 00:31 | 24:00 |  | gram | 29 | 31 |  | n/a
273 | Adult | hydrocortisone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:30 | mg | 100 | 1000 |  | mg/mL | 0.5 | 1 | 20
274 | Adult | ibutilide | pt weight < 60 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:07 |  | 00:30 | 00:10 | mg/kg | 0.009 | 0.011 |  | mg/mL | 0.001 |  | 0.02
275 | Adult | ibutilide | pt weight >= 60 kg | 1 | mg | 60 | mL | 0.0167 | mg/mL | Pump | Primary and Secondary | 00:07 |  | 00:30 | 00:10 | mg | 0.4 | 1.25 |  | n/a
276 | Adult | imipenem/cilastatin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 199 | 1000 |  | n/a
277 | Adult | imipenem/cilastatin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 199 | 1000 |  | n/a
278 | Adult | imipenem/cilastatin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 01:00 | mg | 199 | 1000 |  | mg/mL | 1 | 2 | 5
279 | Adult | imipenem/cilastatin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 0.5 | 1 |  | n/a
280 | Adult | imipenem/cilastatin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | gram | 0.5 | 1 |  | n/a
281 | Adult | imipenem/cilastatin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 01:00 | gram | 0.5 | 1 |  | gram/mL | 0.001 | 0.003 | 0.006
282 | Adult | immune glob (IVIG) |  | 5 | gram | 50 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:00 |  | gram | 1 | 10.1 |  | n/a |  |  |  | Hypersensitivity
283 | Adult | immune glob (IVIG) |  | 10 | gram | 100 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:40 | 08:00 |  | gram | 1 | 10.1 |  | n/a |  |  |  | Hypersensitivity
284 | Adult | immune glob (IVIG) |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 02:00 | 36:00 |  | gram/kg | 1 | 2.4 |  | gram/mL | 0.01 | 0.05 | 0.1 | Hypersensitivity
285 | Adult | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg | 50 | 2000 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
286 | Adult | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg/kg | 2 | 10 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
287 | Adult | iron sucrose |  | 100 | mg | 105 | mL | 0.9524 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 08:00 | 00:30 | mg | 99 |  | 510 | n/a |  |  |  | Hypersensitivity
288 | Adult | iron sucrose |  | 200 | mg | 110 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 08:00 | 00:30 | mg | 99 |  | 510 | n/a |  |  |  | Hypersensitivity
289 | Adult | iron sucrose |  | 300 | mg | 265 | mL | 1.1321 | mg/mL | Pump | Primary and Secondary | 01:25 |  | 08:00 | 01:30 | mg | 99 |  | 510 | n/a |  |  |  | Hypersensitivity
290 | Adult | iron sucrose |  | 400 | mg | 270 | mL | 1.4815 | mg/mL | Pump | Primary and Secondary | 02:25 |  | 08:00 | 02:30 | mg | 99 |  | 510 | n/a |  |  |  | Hypersensitivity
291 | Adult | iron sucrose |  | 500 | mg | 275 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 03:25 |  | 08:00 | 04:00 | mg | 99 |  | 510 | n/a |  |  |  | Hypersensitivity
292 | Adult | itraconazole |  | 200 | mg | 60 | mL | 3.3333 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 100 | 400 |  | n/a |  |  |  | In-line filter
293 | Adult | itraconazole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 100 | 400 |  | mg/mL | 3.29 | 3.3 | 3.4 | In-line filter
294 | Adult | Kcentra (Factor IX) | .INR 2 to < 4 | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | unit/kg | 24 | 26 |  | unit/mL | 19 | 20 | 30 | INR 2-4
295 | Adult | Kcentra (Factor IX) | .INR 4 - 6 | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | unit/kg | 34 | 36 |  | unit/mL | 19 | 20 | 30 | INR 4-6
296 | Adult | Kcentra (Factor IX) | INR > 6 | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | unit/kg | 49 | 53 |  | unit/mL | 19 | 20 | 30 | INR > 6
297 | Adult | Kinevac (sincalide) |  | - - | mcg | - - | mL |  | mcg/mL | Pump | Primary and Secondary |  | 00:12 | 02:00 | 00:16 | mcg/kg | 0.018 | 0.13 |  | mcg/mL | 0.01 |  | 0.5
298 | Adult | lacosamide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 25 | 200 |  | mg/mL | 0.25 | 0.5 | 4.1
299 | Adult | levETIRAcetam |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:06 | 02:00 | 00:15 | mg | 250 | 2001 |  | mg/mL | 2 | 2.5 | 20.1
300 | Adult | levETIRAcetam |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:06 | 02:00 | 00:15 | mg/kg | 1 | 60 |  | mg/mL | 2 | 2.5 | 20.1
301 | Adult | levoFLOXacin |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
302 | Adult | levoFLOXacin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
303 | Adult | levoFLOXacin |  | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 01:15 | 02:00 | 01:30 | mg | 250 | 750 |  | n/a
304 | Adult | linezolid |  | 600 | mg | 300 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 04:00 | 01:00 | mg | 400 | 800 |  | n/a
305 | Adult | mag sulfate |  | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 02:00 | gram | 0.5 | 7 | 8 | n/a
306 | Adult | mag sulfate |  | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:30 | 06:00 | 04:00 | gram | 0.5 | 7 | 8 | n/a
307 | Adult | mag sulfate |  | 4 | gram | 100 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 09:00 |  | gram | 0.5 | 7 | 8 | n/a
308 | Adult | mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:30 | 12:00 |  | gram | 0.5 | 7 | 8 | gram/mL | 0.005 | 0.009 | 0.12
309 | Adult | mannitol 20% |  | 100 | gram | 500 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 00:31 | 04:00 |  | gram | 10 | 100 |  | n/a |  |  |  | In-line filter
310 | Adult | mannitol 20% |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:31 | 08:00 |  | gram | 10 | 100 |  | gram/mL | 0.15 | 0.18 | 0.21 | In-line filter
311 | Adult | mannitol 20% |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:31 | 08:00 |  | gram/kg | 0.1 | 2.1 |  | gram/mL | 0.15 | 0.18 | 0.21 | In-line filter
312 | Adult | meropenem | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
313 | Adult | meropenem | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
314 | Adult | meropenem | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
315 | Adult | meropenem | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
316 | Adult | meropenem | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
317 | Adult | meropenem | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
318 | Adult | meropenem | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
319 | Adult | meropenem | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
320 | Adult | meropenem | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
321 | Adult | meropenem | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
322 | Adult | meropenem/vaborbacta |  | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
323 | Adult | meropenem/vaborbacta |  | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
324 | Adult | meropenem/vaborbacta |  | 2 | gram | 250 | mL | 0.008 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
325 | Adult | meropenem/vaborbacta |  | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
326 | Adult | meropenem/vaborbacta |  | 4 | gram | 250 | mL | 0.016 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
327 | Adult | meropenem/vaborbacta |  | 4 | gram | 500 | mL | 0.008 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
328 | Adult | meropenem/vaborbacta |  | 4 | gram | 1000 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
329 | Adult | methocarbamol |  | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:10 |  | 04:00 |  | mg | 249 |  | 1001 | n/a
330 | Adult | methocarbamol |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 04:00 |  | mg | 249 |  | 1001 | mg/mL | 0.5 | 1 | 15
331 | Adult | methylPREDNISolone |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | n/a
332 | Adult | methylPREDNISolone |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | n/a
333 | Adult | methylPREDNISolone |  | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 125 | 2000 |  | n/a
334 | Adult | methylPREDNISolone |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 125 | 2000 |  | n/a
335 | Adult | methylPREDNISolone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | mg/mL | 0.5 | 1 | 25
336 | Adult | metroNIDAZOLE |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 250 | 1000 |  | n/a
337 | Adult | metroNIDAZOLE |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 250 | 1000 |  | n/a
338 | Adult | micafungin |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
339 | Adult | micafungin |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
340 | Adult | micafungin |  | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
341 | Adult | MVI + Thiame + Folic |  | 10 | unit | 1011 | mL | 0.0099 | unit/mL | Pump | Primary and Secondary | 01:01 |  | 24:00 | 08:06 | unit | 9.9 | 10 |  | n/a
342 | Adult | MVI Only |  | 10 | unit | 510 | mL | 0.0196 | unit/mL | Pump | Primary and Secondary | 00:31 |  | 24:00 | 04:05 | unit | 9.9 | 10.1 |  | n/a
343 | Adult | MVI Only |  | 10 | unit | 1010 | mL | 0.0099 | unit/mL | Pump | Primary and Secondary | 01:01 |  | 24:00 | 08:05 | unit | 9.9 | 10.1 |  | n/a
344 | Adult | nafcillin | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
345 | Adult | nafcillin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
346 | Adult | nafcillin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
347 | Adult | nafcillin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
348 | Adult | nafcillin | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | mg | 200 | 2000 |  | n/a
349 | Adult | nafcillin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | mg | 200 | 2000 |  | n/a
350 | Adult | nafcillin | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
351 | Adult | nafcillin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
352 | Adult | nafcillin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
353 | Adult | nafcillin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
354 | Adult | nafcillin | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | gram | 0.2 | 2 |  | n/a
355 | Adult | nafcillin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | gram | 0.2 | 2 |  | n/a
356 | Adult | ondansetron |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:30 | mg | 2 | 32 |  | mg/mL | 0.02 | 0.04 | 0.64
357 | Adult | pamidronate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 02:00 | 20:00 | 04:00 | mg | 29 | 99 | 110 | mg/mL | 0.001 |  | 0.4 | Dedicated IV Line
358 | Adult | pantoprazole bolus |  | 80 | mg | 100 | mL | 0.8 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:15 | mg | 40 | 80 |  | n/a
359 | Adult | penG 2.5 Mill Units |  | 2.5 | unit | 50 | mL | 0.05 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | unit | 2 | 3 |  | n/a |  |  |  | 1unit=1million
360 | Adult | penG 2.5 Mill Units |  | 2.5 | unit | 100 | mL | 0.025 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | unit | 2 | 3 |  | n/a |  |  |  | 1unit=1million
361 | Adult | penG 3 Million Units |  | 3 | unit | 50 | mL | 0.06 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | unit | 2.8 | 3.2 | 3.3 | n/a |  |  |  | 1unit=1million
362 | Adult | penG 5 Million Units |  | 5 | unit | 50 | mL | 0.1 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | unit | 4.5 | 5.5 |  | n/a |  |  |  | 1unit=1million
363 | Adult | penG 5 Million Units |  | 5 | unit | 100 | mL | 0.05 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | unit | 4.5 | 5.5 |  | n/a |  |  |  | 1unit=1million
364 | Adult | penG x Million Units |  | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | unit | 1 | 6 |  | unit/mL | 0.01 | 0.02 | 0.12 | 1unit=1million
365 | Adult | PHENobarbital |  | 130 | mg | 52 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:04 |  | 04:00 |  | mg | 5 | 270 |  | n/a
366 | Adult | PHENobarbital |  | 130 | mg | 102 | mL | 1.2745 | mg/mL | Pump | Primary and Secondary | 00:07 |  | 04:00 |  | mg | 5 | 270 |  | n/a
367 | Adult | PHENobarbital |  | 260 | mg | 54 | mL | 4.8148 | mg/mL | Pump | Primary and Secondary | 00:04 |  | 04:00 |  | mg | 5 | 270 |  | n/a
368 | Adult | PHENobarbital |  | 260 | mg | 104 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:07 |  | 04:00 |  | mg | 5 | 270 |  | n/a
369 | Adult | PHENobarbital |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:04 |  | 04:00 |  | mg | 5 | 270 |  | No
370 | Adult | phenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:30 | mg/kg | 0.5 | 20 |  | mg/mL | 0.5 | 1 | 10 | In-line filter
371 | Adult | phytonadione |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 01:00 | 00:30 | mg | 1 | 10 |  | mg/mL | 0.005 | 0.01 | 0.2
372 | Adult | piperacillin/tazo |  | 2.25 | gram | 50 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:30 | 00:30 | gram | 3 |  | 4.6 | n/a
373 | Adult | piperacillin/tazo |  | 2.25 | gram | 100 | mL | 0.0225 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:30 | 00:30 | gram | 3 |  | 4.6 | n/a
374 | Adult | piperacillin/tazo |  | 3.375 | gram | 50 | mL | 0.0675 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 | 04:00 | gram | 3 |  | 4.6 | n/a
375 | Adult | piperacillin/tazo |  | 3.375 | gram | 100 | mL | 0.0338 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 | 04:00 | gram | 3 |  | 4.6 | n/a
376 | Adult | piperacillin/tazo |  | 4.5 | gram | 50 | mL | 0.09 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
377 | Adult | piperacillin/tazo |  | 4.5 | gram | 100 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
378 | Adult | potassium acetate |  | 40 | mEq | 520 | mL | 0.0769 | mEq/mL | Pump | Primary and Secondary | 02:00 | 03:30 | 12:00 | 04:00 | mEq | 1 | 46 | 48 | n/a
379 | Adult | potassium acetate |  | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:45 |  | 12:00 |  | mEq | 1 | 46 | 48 | mEq/mL |  | 0.01 | 0.46
380 | Adult | potassium chloride | Central Line | 10 | mEq | 100 | mL | 0.1 | mEq/mL | Pump | Primary and Secondary | 00:14 | 00:45 | 06:00 | 01:00 | mEq | 5 |  | 41 | n/a
381 | Adult | potassium chloride | Central Line | 20 | mEq | 100 | mL | 0.2 | mEq/mL | Pump | Primary and Secondary | 00:25 | 01:30 | 08:00 | 02:00 | mEq | 5 |  | 41 | n/a
382 | Adult | potassium chloride | Central Line | 40 | mEq | 520 | mL | 0.0769 | mEq/mL | Pump | Primary and Secondary | 00:55 | 03:15 | 12:00 | 04:00 | mEq | 5 |  | 41 | n/a
383 | Adult | potassium chloride | Central Line | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:55 | 02:15 | 12:00 |  | mEq | 5 |  | 41 | mEq/mL | 0.01 | 0.02 | 0.21
384 | Adult | potassium chloride | Peripheral Line | 10 | mEq | 100 | mL | 0.1 | mEq/mL | Pump | Primary and Secondary | 00:25 | 00:45 | 06:00 | 01:00 | mEq | 5 |  | 41 | n/a
385 | Adult | potassium chloride | Peripheral Line | 40 | mEq | 520 | mL | 0.0769 | mEq/mL | Pump | Primary and Secondary | 03:15 |  | 12:00 | 04:00 | mEq | 5 |  | 41 | n/a
386 | Adult | potassium chloride | Peripheral Line | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:50 | 01:45 | 12:00 |  | mEq | 5 |  | 41 | mEq/mL | 0.04 | 0.07 | 0.11
387 | Adult | potassium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 01:00 | 02:00 | 09:00 |  | mmol | 15 | 40 |  | mmol/mL | 0.005 | 0.01 | 0.24
388 | Adult | pralidoxime | Load <67kg WtBasedGM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:15 | gram/kg | 0.001 |  | 0.034 | gram/mL |  | 0.009 | 0.051
389 | Adult | pralidoxime | Load <67kg WtBasedMG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:15 | mg/kg | 1 |  | 34 | mg/mL |  | 9 | 51
390 | Adult | pralidoxime | Load >67kg WtBasedGM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:20 | gram | 1 | 21 |  | gram/mL |  | 0.009 | 0.051
391 | Adult | pralidoxime | Load >67kg WtBasedMG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:20 | mg | 500 |  | 2300 | mg/mL |  | 9 | 51
392 | Adult | pralidoxime | Load NonWtBased MG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:20 | mg | 900 |  | 2300 | mg/mL |  | 9 | 51
393 | Adult | pralidoxime | Load NonWtBasedGM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:15 | gram | 0.9 |  | 2.3 | gram/mL |  | 0.009 | 0.051
394 | Adult | pralidoxime | Maintenance Dose GM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 |  | 24:00 |  | gram | 0.9 |  | 2.3 | gram/mL |  | 0.009 | 0.051
395 | Adult | pralidoxime | Maintenance Dose MG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 24:00 |  | mg | 900 |  | 2300 | mg/mL | 5 | 9 | 51
396 | Adult | promethazine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | mg | 6.25 | 25 |  | mg/mL | 0.031 |  | 1
397 | Adult | protamine sulfate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 |  | 01:00 | 00:10 | mg | 0.4 |  | 50 | mg/mL | 0.004 |  | 1.11
398 | Adult | raNITIdine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg | 25 | 51 |  | mg/mL | 0.1 | 0.5 | 2.5
399 | Adult | remdesivir |  | 100 | mg | 230 | mL | 0.4348 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
400 | Adult | remdesivir |  | 100 | mg | 250 | mL | 0.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
401 | Adult | remdesivir |  | 200 | mg | 210 | mL | 0.9524 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
402 | Adult | remdesivir |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
403 | Adult | rifAMPin |  | 300 | mg | 250 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
404 | Adult | rifAMPin |  | 600 | mg | 500 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
405 | Adult | rifAMPin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | mg/mL | 0.1 | 0.5 | 6
406 | Adult | riTUXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:15 | 02:45 | 25:00 |  | mg | 250 | 1100 |  | mg/mL | 0.1 | 0.2 | 4 | Premedicate
407 | Adult | SMX/TMP |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:00 | 02:00 | 01:30 | mg/kg | 1.5 | 7 |  | mg/mL | 0.05 | 0.1 | 1.2 | 48 Hour Tubing
408 | Adult | sodium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 00:45 |  | 12:00 |  | mmol | 15 | 60 |  | mmol/mL | 0.005 | 0.01 | 0.25
409 | Adult | sotrovimab |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:22 | 00:23 | 04:00 | 00:30 | mg | 490 | 510 | 511 | n/a |  |  |  | Low Binding Filter
410 | Adult | sotrovimab |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:22 | 00:23 | 04:00 | 00:30 | mg | 490 | 510 | 511 | n/a |  |  |  | Low Binding Filter
411 | Adult | tedizolid phosphate |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 01:15 | 01:00 | mg | 175 | 225 | 250 | n/a
412 | Adult | thiamine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:15 | 00:30 | mg | 90 | 550 |  | mg/mL | 0.001 | 0.7 | 10.1
413 | Adult | ticarcillin/clav |  | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
414 | Adult | ticarcillin/clav |  | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
415 | Adult | ticarcillin/clav |  | 3.1 | gram | 50 | mL | 0.062 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
416 | Adult | ticarcillin/clav |  | 3.1 | gram | 100 | mL | 0.031 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
417 | Adult | tigecycline |  | 50 | mg | 50 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:15 | 00:30 | mg | 25 | 100 |  | n/a
418 | Adult | tigecycline |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:15 | 00:30 | mg | 25 | 100 |  | n/a
419 | Adult | tigecycline |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:15 | 00:30 | mg | 25 | 100 |  | n/a
420 | Adult | tobramycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 |  | 01:00 | 00:30 | mg/kg | 0.5 | 9 |  | mg/mL | 0.1 | 0.5 | 10
421 | Adult | tocilizumab |  | 800 | mg | 100 | mL | 8 | mg/mL | Pump | Primary and Secondary | 00:55 |  | 04:00 | 01:00 | mg | 161 |  | 801 | n/a |  |  |  | Dedicated IV Line
422 | Adult | tocilizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:55 |  | 04:00 | 01:00 | mg/kg | 3.9 |  | 8.1 | mg/mL | 1 | 1.1 | 8.1 | Dedicated IV Line
423 | Adult | tranexamic BOLUS | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:09 |  | 00:40 | 00:15 | mg | 500 | 2000 |  | n/a
424 | Adult | tranexamic BOLUS | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:07 | 00:20 | 00:40 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
425 | Adult | tranexamic BOLUS | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary | 00:10 |  | 00:40 | 00:15 | gram | 0.5 | 2 |  | n/a
426 | Adult | tranexamic BOLUS | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:07 | 00:20 | 00:40 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
427 | Adult | valproate sodium |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg | 250 | 2000 |  | mg/mL | 0.1 | 1 | 11
428 | Adult | valproate sodium |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg/kg | 1 | 40 |  | mg/mL | 0.1 | 1 | 11
429 | Adult | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 500 | 2000 |  | n/a
430 | Adult | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
431 | Adult | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
432 | Adult | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
433 | Adult | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
434 | Adult | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | mg | 500 | 2000 |  | n/a
435 | Adult | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg | 500 | 2000 |  | n/a
436 | Adult | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg | 500 | 2000 |  | n/a
437 | Adult | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
438 | Adult | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
439 | Adult | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
440 | Adult | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
441 | Adult | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | mg | 500 | 2000 |  | mg/mL | 1 | 2.9 | 5.1
442 | Adult | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg/kg | 5 |  | 30 | n/a
443 | Adult | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
444 | Adult | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
445 | Adult | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
446 | Adult | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
447 | Adult | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
448 | Adult | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
449 | Adult | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
450 | Adult | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
451 | Adult | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
452 | Adult | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
453 | Adult | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
454 | Adult | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | mg/kg | 5 |  | 30 | mg/mL | 1 | 2.9 | 5.1
455 | Adult | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.5 | 2 |  | n/a
456 | Adult | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
457 | Adult | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
458 | Adult | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
459 | Adult | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
460 | Adult | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | gram | 0.5 | 2 |  | n/a
461 | Adult | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram | 0.5 | 2 |  | n/a
462 | Adult | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram | 0.5 | 2 |  | n/a
463 | Adult | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
464 | Adult | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
465 | Adult | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
466 | Adult | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
467 | Adult | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.002 | 0.006
468 | Adult | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram/kg | 0.005 |  | 0.03 | n/a
469 | Adult | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
470 | Adult | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
471 | Adult | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
472 | Adult | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
473 | Adult | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
474 | Adult | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
475 | Adult | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
476 | Adult | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
477 | Adult | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
478 | Adult | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
479 | Adult | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
480 | Adult | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | gram/kg | 0.005 |  | 0.03 | gram/mL | 0.001 | 0.002 | 0.006
481 | Adult | voriconazole |  | 200 | mg | 50 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 02:00 | 01:00 | mg | 200 | 400 |  | n/a
482 | Adult | voriconazole |  | 400 | mg | 100 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 01:40 | 04:00 | 02:00 | mg | 200 | 400 |  | n/a
483 | Adult | voriconazole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 | 01:00 | 04:00 | 02:00 | mg/kg | 2 | 7 |  | mg/mL |  | 0.5 | 5
484 | Adult | zidovudine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 24:00 |  | mg/kg | 1 | 2.2 | 3 | mg/mL | 0.001 | 0.01 | 4.2
485 | Adult | zoledronic acid |  | 4 | mg | 100 | mL | 0.04 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | n/a
486 | Adult | zoledronic acid |  | 5 | mg | 100 | mL | 0.05 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | n/a
487 | Adult | zoledronic acid |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | mg/mL | 0.028 | 0.029 | 0.051
488 | Adult | zz amiodarone BOLUS |  | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:08 | 00:15 | 00:10 | mg | 149 | 151 |  | n/a |  |  |  | Code only drugs
489 | Adult | zz mag sulfate |  | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary | 00:07 | 00:14 | 02:00 | 00:20 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
490 | Adult | zz mag sulfate |  | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary | 00:07 | 00:14 | 02:00 | 00:20 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
491 | Adult | zz mag sulfate |  | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary | 00:14 | 00:28 | 04:00 | 00:40 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
492 | Adult | zz mag sulfate |  | 4 | gram | 100 | mL | 0.04 | gram/mL | Pump | Primary and Secondary | 00:28 | 00:56 | 08:00 | 01:20 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
493 | Adult | zz mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:14 | 00:28 | 08:00 |  | gram | 1 | 6 |  | gram/mL | 0.005 | 0.01 | 0.08 | Code only drugs
494 | Nursery | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:50 | 00:55 | 02:00 | 01:00 | mg/kg | 4 | 21 | 22 | mg/mL | 1 | 4 | 7
495 | Nursery | ampicillin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:07 | 00:08 | 01:10 | 00:30 | mg/kg | 24 | 101 | 105 | mg/mL |  | 10 | 105
496 | Nursery | ceFAZolin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:07 | 00:08 | 01:05 | 00:30 | mg/kg | 8 | 55 | 60 | mg/mL |  | 45 | 55
497 | Nursery | cefotaxime |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:12 | 00:13 | 01:05 | 00:30 | mg/kg | 12.5 | 80 | 81 | mg/mL |  | 45 | 55
498 | Nursery | cefTRIAXone |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:45 | 01:10 | 01:00 | mg/kg | 20 | 135 | 140 | mg/mL |  | 35 | 45 | CeftriSyringe
499 | Nursery | clindamycin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:08 | 00:09 | 01:05 | 00:30 | mg/kg | 3 | 7.1 | 7.2 | mg/mL |  | 13 | 17
500 | Nursery | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:25 | 02:15 | 00:30 | mg/kg | 0.5 | 7 | 7.1 | mg/mL |  | 3 | 7
501 | Nursery | nafcillin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:25 | 01:10 | 00:30 | mg/kg | 10 | 53 | 55 | mg/mL |  | 35 | 45
502 | Nursery | PenG Potass Syringe |  | - - | unit | - - | mL |  | unit/mL | Syringe | Primary Only | 00:15 |  | 02:00 | 00:30 | unit/kg | 20 | 52.5 | 152 | unit/mL | 10 |  | 150 | PenG-Syringe
503 | Nursery | piperacillin/tazo |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:25 | 04:15 | 00:30 | mg/kg | 49 | 125 | 130 | mg/mL |  | 15 | 85
504 | Nursery | vancomycin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:52 | 00:55 | 03:10 | 01:00 | mg/kg | 4 | 22 | 23 | mg/mL |  | 4.5 | 6.5
505 | ONC Infusion Center | abatacept |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
506 | ONC Infusion Center | abatacept |  | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
507 | ONC Infusion Center | abatacept |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
508 | ONC Infusion Center | ado-trastuzumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:28 | 00:29 | 02:00 |  | mg/kg | 2.25 | 3.9 | 4 | mg/mL | 0.3 | 0.4 | 3.6 | 0.22 Filter Needed
509 | ONC Infusion Center | aprepitant |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:30 | mg | 99 | 131 |  | n/a |  |  |  | DEHP
510 | ONC Infusion Center | aprepitant |  | 130 | mg | 100 | mL | 1.3 | mg/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:30 | mg | 99 | 131 |  | n/a |  |  |  | DEHP
511 | ONC Infusion Center | atezolizumab |  | 840 | mg | 250 | mL | 3.36 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 799 | 1681 |  | n/a
512 | ONC Infusion Center | atezolizumab |  | 1200 | mg | 250 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 799 | 1681 |  | n/a
513 | ONC Infusion Center | atezolizumab |  | 1680 | mg | 250 | mL | 6.72 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 799 | 1681 |  | n/a
514 | ONC Infusion Center | Bendamusti (Bendeka) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:08 | 00:15 | 00:10 | mg/m2 | 25 | 120 |  | mg/mL |  | 1.85 | 5.6
515 | ONC Infusion Center | Bendamusti (Treanda) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:20 | 00:30 | mg/m2 | 25 | 120 |  | mg/mL |  | 0.2 | 0.7
516 | ONC Infusion Center | bevacizumab | First Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:15 | 01:30 | 03:00 | 01:30 | mg/kg | 5 | 15 |  | mg/mL |  | 1 | 25
517 | ONC Infusion Center | bevacizumab | Second Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 01:30 | 01:00 | mg/kg | 5 | 15 |  | mg/mL |  | 1 | 25
518 | ONC Infusion Center | bevacizumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 01:30 | 00:30 | mg/kg | 5 | 15 |  | mg/mL |  | 1 | 25
519 | ONC Infusion Center | bleomycin | BSA Based Dose | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:05 | 00:15 | 00:10 | unit/m2 | 2 | 30 |  | unit/mL |  | 0.1 | 3
520 | ONC Infusion Center | bleomycin | Non-Based Based Dose | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:08 | 00:20 | 00:10 | unit | 1 | 60 | 61 | unit/mL |  | 0.1 | 3
521 | ONC Infusion Center | brentuximab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 00:40 | 00:30 | mg/kg | 1 | 1.8 | 1.81 | mg/mL |  | 0.4 | 1.8 | MAX DOSE 180 mg
522 | ONC Infusion Center | cabazitaxel |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:10 | 01:00 | mg/m2 | 15 | 25 |  | mg/mL |  | 0.1 | 0.26 | 0.22 Filter Needed
523 | ONC Infusion Center | CARBOplatin | Based on AUC | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 24:00 | 00:30 | mg | 100 | 800 |  | mg/mL |  | 0.4 | 5
524 | ONC Infusion Center | CARBOplatin | Based on BSA | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 | 00:30 | mg/m2 | 60 | 700 |  | mg/mL |  | 0.5 | 5
525 | ONC Infusion Center | carfilzomib |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:35 | 00:30 | mg/m2 | 15 | 70 |  | mg/mL |  | 0.3 | 1.6
526 | ONC Infusion Center | cemiplimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 00:35 | 00:30 | mg | 349 | 351 |  | mg/mL |  | 1 | 20 | In-line filter
527 | ONC Infusion Center | cetuximab | Loading Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:45 |  | 04:00 | 02:00 | mg/m2 | 400 | 500 |  | No |  |  |  | Low Binding Filter
528 | ONC Infusion Center | cetuximab | Maintenance Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg/m2 | 150 | 250 |  | No |  |  |  | Low Binding Filter
529 | ONC Infusion Center | CISplatin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 24:00 |  | mg/m2 | 10 | 100 |  | mg/mL |  | 0.04 | 0.5
530 | ONC Infusion Center | crizanlizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 04:30 | 00:30 | mg/kg | 4.5 | 5.5 | 5.51 | mg/mL | 1 | 1.6 | 5 | 0.22 Filter Needed
531 | ONC Infusion Center | cyclophosphamide | BSA Based Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:30 | 01:00 | mg/m2 | 100 | 2000 |  | mg/mL | 1.9 | 2 | 20
532 | ONC Infusion Center | cyclophosphamide | Weight Based | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:30 |  | mg/kg | 1 | 50 | 51 | mg/mL | 1.9 | 2 | 20
533 | ONC Infusion Center | dacarbazine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:40 |  | mg/m2 | 175 | 850 |  | mg/mL |  | 0.19 | 20.1
534 | ONC Infusion Center | daratumumab | First Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 02:20 | 02:30 | 20:00 |  | mg/kg | 7 | 16 | 16.1 | mg/mL |  | 0.4 | 4 | 0.22 Filter Needed
535 | ONC Infusion Center | daratumumab | Second Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 02:30 | 04:00 | 10:00 |  | mg/kg | 15 | 16 | 16.1 | mg/mL |  | 0.9 | 5 | 0.22 Filter Needed
536 | ONC Infusion Center | daratumumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 02:30 | 03:50 | 10:00 |  | mg/kg | 15 | 16 | 16.1 | mg/mL |  | 0.9 | 5 | 0.22 Filter Needed
537 | ONC Infusion Center | decitabine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 |  | mg/m2 | 10 | 22 |  | mg/mL | 0.089 | 0.09 | 1.1
538 | ONC Infusion Center | dexamethasone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:30 | 00:15 | mg | 1 | 21 |  | mg/mL | 0.008 | 0.009 | 0.4
539 | ONC Infusion Center | DOCEtaxel |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg/m2 | 16 | 100 |  | mg/mL |  | 0.3 | 0.74 | DEHP
540 | ONC Infusion Center | DOXOrubicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 24:00 |  | mg/m2 | 3 | 90 |  | mg/mL |  | 0.01 | 5
541 | ONC Infusion Center | DOXOrubicin, liposo |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg/m2 | 2.5 | 55 |  | mg/mL |  | 0.1 | 0.36 | NO FILTER
542 | ONC Infusion Center | durvalumab | .Weight >= 30 kg | 1500 | mg | 280 | mL | 5.3571 | mg/mL | Pump | Primary Only |  | 00:55 | 08:00 | 01:00 | mg | 1499 |  | 1501 | n/a |  |  |  | Low Binding Filter
543 | ONC Infusion Center | durvalumab | Weight < 30 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only |  | 00:55 | 08:00 | 01:00 | mg/kg | 9 | 21 |  | mg/mL | 0.9 | 1 | 15 | Low Binding Filter
544 | ONC Infusion Center | enfortumab vedotin | .wt <= 100 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 04:00 | 00:30 | mg/kg | 0.4 | 1.25 | 1.3 | mg/mL | 0.29 | 0.3 | 4
545 | ONC Infusion Center | enfortumab vedotin | .wt > 100 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 04:00 | 00:30 | mg | 49 | 125 | 125.1 | mg/mL | 0.29 | 0.3 | 4
546 | ONC Infusion Center | eriBULin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:02 | 00:20 | 00:10 | mg/m2 | 0.7 | 1.4 |  | mg/mL |  | 0.001 | 0.5 | No Dextrose
547 | ONC Infusion Center | etoposide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 03:00 |  | mg/m2 | 35 | 750 |  | mg/mL |  | 0.2 | 0.4 | Non-PVC
548 | ONC Infusion Center | ferumoxytol | .MG Dosing | 510 | mg | 117 | mL | 4.359 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:30 | mg | 509 | 1021 | 1022 | n/a
549 | ONC Infusion Center | ferumoxytol | .MG Dosing | 1020 | mg | 284 | mL | 3.5915 | mg/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | mg | 509 | 1021 | 1022 | n/a
550 | ONC Infusion Center | ferumoxytol | GM dosing | 0.51 | gram | 117 | mL | 0.0044 | gram/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:30 | gram | 0.509 | 1.021 | 1.022 | n/a
551 | ONC Infusion Center | ferumoxytol | GM dosing | 1.02 | gram | 284 | mL | 0.0036 | gram/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | gram | 0.509 | 1.021 | 1.022 | n/a
552 | ONC Infusion Center | fluorouracil |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 24:00 |  | mg/m2 | 200 | 3000 |  | mg/mL |  | 0.1 | 50
553 | ONC Infusion Center | fosaprepitant |  | 150 | mg | 250 | mL | 0.6 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 00:40 | 00:30 | mg | 149 | 151 |  | n/a
554 | ONC Infusion Center | gemcitabine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg/m2 | 300 | 1250 |  | mg/mL |  | 0.1 | 38
555 | ONC Infusion Center | golimumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:00 | 00:30 | mg/kg | 1.8 |  | 2.2 | mg/mL | 0.49 | 0.5 | 4 | Low Binding Filter
556 | ONC Infusion Center | ifosfamide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 25:00 |  | mg/m2 | 250 | 5000 |  | mg/mL |  | 0.6 | 80
557 | ONC Infusion Center | ipilimumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:25 | 01:35 | 01:30 | mg/kg | 1 | 10 |  | mg/mL |  | 1 | 2
558 | ONC Infusion Center | irinotecan |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:20 |  | 02:30 | 01:30 | mg/m2 | 20 | 600 |  | mg/mL |  | 0.12 | 2.8
559 | ONC Infusion Center | irinotecan liposomal |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:20 | 02:00 | 01:30 | mg/m2 | 34 | 71 |  | mg/mL | 0.05 | 0.1 | 0.6 | NO FILTER
560 | ONC Infusion Center | leucovorin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mg/m2 | 10 | 1000 |  | mg/mL |  | 0.05 | 20
561 | ONC Infusion Center | leucovorin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mg | 1 | 100 |  | mg/mL |  | 0.05 | 20
562 | ONC Infusion Center | mag sulfate |  | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 02:30 | 01:00 | gram | 0.9 | 1.1 |  | n/a
563 | ONC Infusion Center | mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:45 | 08:00 |  | gram | 0.9 | 1.1 |  | No
564 | ONC Infusion Center | mesna |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 25:00 |  | mg/m2 | 300 | 5000 |  | mg/mL |  | 0.01 | 20
565 | ONC Infusion Center | mesna |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 25:00 |  | mg | 1 | 1000 |  | mg/mL |  | 0.01 | 20
566 | ONC Infusion Center | mitoMYcin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg/m2 | 1 | 20 |  | mg/mL |  | 0.02 | 0.04
567 | ONC Infusion Center | mitoMYcin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg | 1 | 20 |  | mg/mL |  | 0.02 | 0.04
568 | ONC Infusion Center | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
569 | ONC Infusion Center | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
570 | ONC Infusion Center | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
571 | ONC Infusion Center | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
572 | ONC Infusion Center | nivolumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | mg/mL |  | 1 | 10 | In-line filter
573 | ONC Infusion Center | obinutuzumab | First Dose | 100 | mg | 104 | mL | 0.9615 | mg/mL | Pump | Primary and Secondary |  | 04:00 | 08:00 | 04:00 | mg | 99 | 100.1 |  | n/a
574 | ONC Infusion Center | obinutuzumab | Subsequent Doses | 900 | mg | 286 | mL | 3.1469 | mg/mL | Pump | Primary and Secondary |  | 02:00 | 60:00 |  | mg | 900 | 1000 |  | n/a
575 | ONC Infusion Center | obinutuzumab | Subsequent Doses | 1000 | mg | 290 | mL | 3.4483 | mg/mL | Pump | Primary and Secondary |  | 02:00 | 60:00 |  | mg | 900 | 1000 |  | n/a
576 | ONC Infusion Center | obinutuzumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 02:00 | 60:00 |  | mg | 900 | 1000 |  | mg/mL |  | 0.9 | 3.5
577 | ONC Infusion Center | ondansetron |  | 8 | mg | 54 | mL | 0.1481 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 4 | 16 |  | n/a
578 | ONC Infusion Center | ondansetron |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 |  | mg | 4 | 16 |  | mg/mL |  | 0.01 | 0.64
579 | ONC Infusion Center | oxaliplatin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:00 | 06:30 | 02:00 | mg/m2 | 20 | 130 |  | mg/mL |  | 0.2 | 1.2
580 | ONC Infusion Center | PACLitaxel |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 24:00 |  | mg/m2 | 40 | 225 |  | mg/mL |  | 0.3 | 1.2 | Filter and DEHP Free
581 | ONC Infusion Center | PACLitaxel protein-b |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/m2 | 49 | 261 |  | mg/mL |  | 4.9 | 5.1 | NO FILTER
582 | ONC Infusion Center | panitumumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg/kg | 3 | 6 |  | mg/mL |  | 2 | 10 | Low Binding Filter
583 | ONC Infusion Center | pembrolizumab |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 00:35 | 00:30 | mg | 100 | 250 |  | n/a |  |  |  | In-line filter
584 | ONC Infusion Center | pembrolizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 00:35 | 00:30 | mg | 100 | 250 |  | mg/mL |  | 1 | 10 | In-line filter
585 | ONC Infusion Center | PEMEtrexed |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:10 | mg/m2 | 240 | 550 | 910 | mg/mL |  | 3 | 13
586 | ONC Infusion Center | pertuzumab | Loading Dose | 840 | mg | 278 | mL | 3.0216 | mg/mL | Pump | Primary and Secondary | 01:00 |  | 02:00 | 01:00 | mg | 839 |  | 841 | n/a
587 | ONC Infusion Center | pertuzumab | Maintenance Dose | 420 | mg | 264 | mL | 1.5909 | mg/mL | Pump | Primary and Secondary | 00:30 |  | 02:00 | 00:30 | mg | 419 |  | 421 | n/a
588 | ONC Infusion Center | polatuzumab vedotin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 03:00 |  | mg/kg | 0.9 | 1.98 | 2 | mg/mL | 0.71 | 0.72 | 2.71 | Low Binding Filter
589 | ONC Infusion Center | ramucirumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 |  | mg/kg | 5 | 10 |  | mg/mL |  | 0.8 | 12 | Low Binding Filter
590 | ONC Infusion Center | raNITIdine |  | 50 | mg | 102 | mL | 0.4902 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:20 | mg | 25 | 50 |  | n/a
591 | ONC Infusion Center | riTUXimab | Based on BSA | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg/m2 | 250 | 500 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
592 | ONC Infusion Center | riTUXimab | Rheumatoid Arthritis | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg | 90 | 1200 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
593 | ONC Infusion Center | temozolomide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:20 | 02:00 | 01:30 | mg/m2 | 74 | 201 |  | mg/mL |  | 2.4 | 2.6 | Flush Before/AfterNS
594 | ONC Infusion Center | topotecan |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg/m2 | 0.2 | 1.5 | 4 | mg/mL |  | 0.01 | 0.5
595 | ONC Infusion Center | trabectedin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 02:45 | 02:50 | 04:00 | 03:00 | mg/m2 | 0.2 | 1.15 | 1.2 | mg/mL |  | 0.001 | 0.01 | 0.22 Filter Needed
596 | ONC Infusion Center | trastuzumab | Loading Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:00 | 03:00 | 01:30 | mg/kg | 4 | 8 |  | mg/mL |  | 0.4 | 8
597 | ONC Infusion Center | trastuzumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 2 | 6 |  | mg/mL |  | 0.4 | 8
598 | ONC Infusion Center | vinBLAStine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 01:00 | 00:15 | mg/m2 | 0.75 | 18.5 | 18.6 | mg/mL |  | 0.001 | 0.6 | Vesicant
599 | ONC Infusion Center | vinCRIStine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:20 | 00:10 | mg | 0.4 | 2 | 2.1 | mg/mL |  | 0.001 | 0.1 | Vesicant
600 | ONC Infusion Center | vinorelbine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:06 | 00:30 | 00:10 | mg/m2 | 5 | 35 |  | mg/mL |  | 0.5 | 2
601 | ONC Infusion Center | ziv-aflibercept |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:55 | 01:05 | 01:00 | mg/kg | 2 | 4 |  | mg/mL |  | 0.6 | 8 | 0.22 Filter Needed
602 | ONC Infusion Center | zoledronic acid |  | 4 | mg | 100 | mL | 0.04 | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 02:00 | 00:30 | mg | 3 | 4 |  | n/a
603 | OutPt InfusionCenter | abatacept |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
604 | OutPt InfusionCenter | abatacept |  | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
605 | OutPt InfusionCenter | abatacept |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
606 | OutPt InfusionCenter | abatacept |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg/kg | 9 |  | 11 | mg/mL | 2 | 3.8 | 10.2 | Low Binding 0.2or1.2
607 | OutPt InfusionCenter | albumin 25% |  | 5 | gram | 20 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:08 | 01:00 | 00:20 | gram | 4.9 | 25.1 |  | n/a
608 | OutPt InfusionCenter | albumin 25% |  | 12.5 | gram | 50 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:16 | 02:00 | 00:50 | gram | 4.9 | 25.1 |  | n/a
609 | OutPt InfusionCenter | albumin 25% |  | 25 | gram | 100 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:32 | 04:00 | 01:40 | gram | 4.9 | 25.1 |  | n/a
610 | OutPt InfusionCenter | amikacin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 |  | mg/kg | 1 | 21 | 39 | mg/mL | 0.2 |  | 5.1
611 | OutPt InfusionCenter | ampho B (liposomal) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only |  | 00:55 | 04:00 | 02:00 | mg/kg | 1.9 | 10.5 |  | mg/mL | 0.8 | 0.9 | 2.1 | AmBisome
612 | OutPt InfusionCenter | belimumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 00:55 |  | 08:00 | 01:00 | mg/kg | 9 |  | 11 | mg/mL | 1.6 |  | 8.5
613 | OutPt InfusionCenter | bezlotoxumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 | 00:51 | 03:00 | 01:00 | mg/kg | 9 | 11 | 11.1 | mg/mL | 0.8 | 0.9 | 10.1 | 0.2 to 5 micron filt
614 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
615 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
616 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
617 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
618 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | mg/mL | 5 | 10 | 40
619 | OutPt InfusionCenter | ceFAZolin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
620 | OutPt InfusionCenter | ceFAZolin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
621 | OutPt InfusionCenter | ceFAZolin | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
622 | OutPt InfusionCenter | ceFAZolin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
623 | OutPt InfusionCenter | ceFAZolin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | gram/mL | 0.005 | 0.01 | 0.04
624 | OutPt InfusionCenter | cefepime | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
625 | OutPt InfusionCenter | cefepime | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
626 | OutPt InfusionCenter | cefepime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
627 | OutPt InfusionCenter | cefepime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
628 | OutPt InfusionCenter | cefepime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
629 | OutPt InfusionCenter | cefepime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
630 | OutPt InfusionCenter | cefepime | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 4 | 4.9 | 40
631 | OutPt InfusionCenter | cefepime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
632 | OutPt InfusionCenter | cefepime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
633 | OutPt InfusionCenter | cefepime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
634 | OutPt InfusionCenter | cefepime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
635 | OutPt InfusionCenter | cefepime | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.004 | 0.009 | 0.04
636 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
637 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
638 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
639 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
640 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
641 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
642 | OutPt InfusionCenter | cefTAZidime | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
643 | OutPt InfusionCenter | cefTAZidime | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
644 | OutPt InfusionCenter | cefTAZidime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
645 | OutPt InfusionCenter | cefTAZidime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
646 | OutPt InfusionCenter | cefTAZidime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
647 | OutPt InfusionCenter | cefTAZidime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
648 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
649 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
650 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
651 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
652 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | mg/mL | 4.9 | 9 | 40
653 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
654 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
655 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
656 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
657 | OutPt InfusionCenter | cefTRIAXone | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | gram/mL | 0.004 | 0.009 | 0.04
658 | OutPt InfusionCenter | ciprofloxacin |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
659 | OutPt InfusionCenter | ciprofloxacin |  | 400 | mg | 200 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
660 | OutPt InfusionCenter | clindamycin |  | 300 | mg | 50 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
661 | OutPt InfusionCenter | clindamycin |  | 300 | mg | 100 | mL | 3 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
662 | OutPt InfusionCenter | clindamycin |  | 600 | mg | 50 | mL | 12 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
663 | OutPt InfusionCenter | clindamycin |  | 600 | mg | 100 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
664 | OutPt InfusionCenter | clindamycin |  | 900 | mg | 50 | mL | 18 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
665 | OutPt InfusionCenter | clindamycin |  | 900 | mg | 100 | mL | 9 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
666 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 375 | mg | 100 | mL | 3.75 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
667 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 375 | mg | 250 | mL | 1.5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
668 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
669 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 500 | mg | 250 | mL | 2 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
670 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
671 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
672 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 1125 | mg | 250 | mL | 4.5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
673 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
674 | OutPt InfusionCenter | dalbavancin | GM dosing | 0.375 | gram | 100 | mL | 0.0038 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
675 | OutPt InfusionCenter | dalbavancin | GM dosing | 0.375 | gram | 250 | mL | 0.0015 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
676 | OutPt InfusionCenter | dalbavancin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
677 | OutPt InfusionCenter | dalbavancin | GM dosing | 0.5 | gram | 250 | mL | 0.002 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
678 | OutPt InfusionCenter | dalbavancin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
679 | OutPt InfusionCenter | dalbavancin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
680 | OutPt InfusionCenter | dalbavancin | GM dosing | 1.125 | gram | 250 | mL | 0.0045 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
681 | OutPt InfusionCenter | dalbavancin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
682 | OutPt InfusionCenter | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg | 100 | 2000 |  | mg/mL | 1 | 2 | 20
683 | OutPt InfusionCenter | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg/kg | 4 | 12.1 |  | mg/mL | 1 | 2 | 20
684 | OutPt InfusionCenter | dexamethasone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:30 | 00:15 | mg | 1 | 21 |  | No
685 | OutPt InfusionCenter | doxycycline |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
686 | OutPt InfusionCenter | doxycycline |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
687 | OutPt InfusionCenter | ertapenem | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
688 | OutPt InfusionCenter | ertapenem | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
689 | OutPt InfusionCenter | ertapenem | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
690 | OutPt InfusionCenter | ertapenem | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
691 | OutPt InfusionCenter | ertapenem | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
692 | OutPt InfusionCenter | ertapenem | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
693 | OutPt InfusionCenter | ertapenem | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
694 | OutPt InfusionCenter | ertapenem | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
695 | OutPt InfusionCenter | erythromycin |  | 250 | mg | 100 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
696 | OutPt InfusionCenter | erythromycin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
697 | OutPt InfusionCenter | erythromycin |  | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:30 | 01:00 | mg | 100 | 1000 |  | n/a
698 | OutPt InfusionCenter | erythromycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | mg/mL | 0.05 | 1 | 5
699 | OutPt InfusionCenter | famotidine |  | 20 | mg | 50 | mL | 0.4 | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg | 8 | 44 | 50 | n/a
700 | OutPt InfusionCenter | famotidine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 02:00 |  | mg | 8 | 44 | 50 | mg/mL | 0.09 | 0.1 | 3.5
701 | OutPt InfusionCenter | ferumoxytol | .MG Dosing | 510 | mg | 117 | mL | 4.359 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 509 | 1021 | 1022 | n/a |  |  |  | Hypersensitivity
702 | OutPt InfusionCenter | ferumoxytol | .MG Dosing | 1020 | mg | 284 | mL | 3.5915 | mg/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | mg | 509 | 1021 | 1022 | n/a |  |  |  | Hypersensitivity
703 | OutPt InfusionCenter | ferumoxytol | GM dosing | 0.51 | gram | 117 | mL | 0.0044 | gram/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | gram | 0.509 | 1.021 | 1.022 | n/a |  |  |  | Hypersensitivity
704 | OutPt InfusionCenter | ferumoxytol | GM dosing | 1.02 | gram | 284 | mL | 0.0036 | gram/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | gram | 0.509 | 1.021 | 1.022 | n/a |  |  |  | Hypersensitivity
705 | OutPt InfusionCenter | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 |  | mg | 10 | 1000 |  | mg/mL | 0.05 | 0.1 | 10
706 | OutPt InfusionCenter | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 |  | mg/kg | 0.5 | 10 |  | mg/mL | 0.05 | 0.1 | 10
707 | OutPt InfusionCenter | golimumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:00 | 00:30 | mg/kg | 1.8 |  | 2.2 | mg/mL | 0.49 | 0.5 | 4 | Low Binding Filter
708 | OutPt InfusionCenter | immune glob (IVIG) |  | 5 | gram | 50 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:00 |  | gram | 1 | 10.1 |  | n/a |  |  |  | Hypersensitivity
709 | OutPt InfusionCenter | immune glob (IVIG) |  | 10 | gram | 100 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:40 | 08:00 |  | gram | 1 | 10.1 |  | n/a |  |  |  | Hypersensitivity
710 | OutPt InfusionCenter | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg | 50 | 2000 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
711 | OutPt InfusionCenter | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg/kg | 2 | 10 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
712 | OutPt InfusionCenter | iron sucrose |  | 100 | mg | 105 | mL | 0.9524 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 08:00 | 00:30 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
713 | OutPt InfusionCenter | iron sucrose |  | 200 | mg | 110 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 08:00 | 00:30 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
714 | OutPt InfusionCenter | iron sucrose |  | 300 | mg | 265 | mL | 1.1321 | mg/mL | Pump | Primary and Secondary | 01:25 |  | 08:00 | 01:30 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
715 | OutPt InfusionCenter | iron sucrose |  | 400 | mg | 270 | mL | 1.4815 | mg/mL | Pump | Primary and Secondary | 02:25 |  | 08:00 | 02:30 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
716 | OutPt InfusionCenter | iron sucrose |  | 500 | mg | 275 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 03:25 |  | 08:00 | 04:00 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
717 | OutPt InfusionCenter | levoFLOXacin |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
718 | OutPt InfusionCenter | levoFLOXacin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
719 | OutPt InfusionCenter | levoFLOXacin |  | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 01:15 | 02:00 | 01:30 | mg | 250 | 750 |  | n/a
720 | OutPt InfusionCenter | mag sulfate |  | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 02:00 | gram | 0.9 | 7 | 8 | n/a
721 | OutPt InfusionCenter | mag sulfate |  | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:30 | 06:00 | 04:00 | gram | 0.9 | 7 | 8 | n/a
722 | OutPt InfusionCenter | mag sulfate |  | 4 | gram | 100 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 01:00 | 09:00 | 08:00 | gram | 0.9 | 7 | 8 | n/a
723 | OutPt InfusionCenter | mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:30 | 12:00 |  | gram | 0.9 | 7 | 8 | gram/mL | 0.005 | 0.009 | 0.12
724 | OutPt InfusionCenter | meropenem | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
725 | OutPt InfusionCenter | meropenem | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
726 | OutPt InfusionCenter | meropenem | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
727 | OutPt InfusionCenter | meropenem | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
728 | OutPt InfusionCenter | meropenem | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
729 | OutPt InfusionCenter | meropenem | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
730 | OutPt InfusionCenter | meropenem | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
731 | OutPt InfusionCenter | meropenem | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
732 | OutPt InfusionCenter | meropenem | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
733 | OutPt InfusionCenter | meropenem | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
734 | OutPt InfusionCenter | methylPREDNISolone |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | n/a
735 | OutPt InfusionCenter | methylPREDNISolone |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | n/a
736 | OutPt InfusionCenter | methylPREDNISolone |  | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 125 | 2000 |  | n/a
737 | OutPt InfusionCenter | methylPREDNISolone |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 125 | 2000 |  | n/a
738 | OutPt InfusionCenter | methylPREDNISolone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 |  | mg | 125 | 2000 |  | mg/mL | 0.5 | 1 | 25
739 | OutPt InfusionCenter | micafungin |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
740 | OutPt InfusionCenter | micafungin |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
741 | OutPt InfusionCenter | micafungin |  | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
742 | OutPt InfusionCenter | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
743 | OutPt InfusionCenter | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
744 | OutPt InfusionCenter | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
745 | OutPt InfusionCenter | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
746 | OutPt InfusionCenter | nivolumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | mg/mL |  | 1 | 10 | In-line filter
747 | OutPt InfusionCenter | ocrelizumab |  | 300 | mg | 260 | mL | 1.1538 | mg/mL | Pump | Primary and Secondary | 02:20 | 02:25 | 24:00 |  | mg | 290 | 610 | 611 | n/a |  |  |  | 0.22 Filter Needed
748 | OutPt InfusionCenter | ocrelizumab |  | 600 | mg | 520 | mL | 1.1538 | mg/mL | Pump | Primary and Secondary | 01:50 | 01:55 | 24:00 |  | mg | 290 | 610 | 611 | n/a |  |  |  | 0.22 Filter Needed
749 | OutPt InfusionCenter | ondansetron |  | 8 | mg | 54 | mL | 0.1481 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 4 | 16 |  | n/a
750 | OutPt InfusionCenter | ondansetron |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 |  | mg | 4 | 16 |  | mg/mL | 0.02 | 0.04 | 0.64
751 | OutPt InfusionCenter | oritavacin | Kimsyra gram | 1.2 | gram | 250 | mL | 0.0048 | gram/mL | Pump | Primary and Secondary | 00:55 |  | 01:15 | 01:00 | gram | 1.199 |  | 1.201 | n/a |  |  |  | Flush Before/After
752 | OutPt InfusionCenter | oritavacin | Kimyrsa mg | 1200 | mg | 250 | mL | 4.8 | mg/mL | Pump | Primary and Secondary | 00:55 |  | 01:15 | 01:00 | mg | 1199 |  | 1201 | n/a |  |  |  | Flush Before/After
753 | OutPt InfusionCenter | oritavacin | Orbactiv gram | 1.2 | gram | 250 | mL | 0.0048 | gram/mL | Pump | Primary and Secondary | 02:45 |  | 03:30 | 03:00 | gram | 1.199 |  | 1.201 | n/a |  |  |  | Flush Before/AfterD5
754 | OutPt InfusionCenter | oritavacin | Orbactiv mg | 1200 | mg | 1000 | mL | 1.2 | mg/mL | Pump | Primary and Secondary | 02:45 |  | 03:30 | 03:00 | mg | 1199 |  | 1201 | n/a |  |  |  | Flush Before/AfterD5
755 | OutPt InfusionCenter | piperacillin/tazo |  | 3.375 | gram | 50 | mL | 0.0675 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
756 | OutPt InfusionCenter | piperacillin/tazo |  | 3.375 | gram | 100 | mL | 0.0338 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
757 | OutPt InfusionCenter | piperacillin/tazo |  | 4.5 | gram | 50 | mL | 0.09 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
758 | OutPt InfusionCenter | piperacillin/tazo |  | 4.5 | gram | 100 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
759 | OutPt InfusionCenter | potassium chloride |  | 10 | mEq | 100 | mL | 0.1 | mEq/mL | Pump | Primary and Secondary | 00:25 | 00:45 | 06:00 | 01:00 | mEq | 5 |  | 41 | n/a
760 | OutPt InfusionCenter | potassium chloride |  | 40 | mEq | 520 | mL | 0.0769 | mEq/mL | Pump | Primary and Secondary | 03:15 |  | 12:00 | 04:00 | mEq | 5 |  | 41 | n/a
761 | OutPt InfusionCenter | potassium chloride |  | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:50 | 01:45 | 12:00 |  | mEq | 5 |  | 41 | mEq/mL | 0.04 | 0.07 | 0.11
762 | OutPt InfusionCenter | promethazine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | mg | 6.25 | 25 |  | mg/mL | 1 | 1.25 | 25
763 | OutPt InfusionCenter | raNITIdine |  | 50 | mg | 102 | mL | 0.4902 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:20 | mg | 25 | 50 |  | n/a
764 | OutPt InfusionCenter | rifAMPin |  | 300 | mg | 250 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
765 | OutPt InfusionCenter | rifAMPin |  | 600 | mg | 500 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
766 | OutPt InfusionCenter | rifAMPin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | mg/mL | 0.1 | 0.5 | 6
767 | OutPt InfusionCenter | riTUXimab | Based on BSA | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg/m2 | 250 | 500 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
768 | OutPt InfusionCenter | riTUXimab | Rheumatoid Arthritis | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg | 90 | 1200 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
769 | OutPt InfusionCenter | spesolimab |  | 900 | mg | 100 | mL | 9 | mg/mL | Pump | Primary Only |  | 01:25 | 03:00 | 01:30 | mg | 899 |  | 901 | n/a |  |  |  | Low Binding Filter
770 | OutPt InfusionCenter | teprotumumab | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 | 00:55 | 08:00 |  | mg/kg | 7 | 21 | 22 | mg/mL | 2.9 | 3 | 18.1
771 | OutPt InfusionCenter | teprotumumab | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 | 00:55 | 08:00 |  | gram/kg | 0.007 | 0.021 | 0.022 | gram/mL | 0.002 | 0.003 | 0.019
772 | OutPt InfusionCenter | tocilizumab |  | 800 | mg | 100 | mL | 8 | mg/mL | Pump | Primary and Secondary | 00:55 |  | 04:00 | 01:00 | mg | 161 |  | 801 | n/a |  |  |  | Dedicated IV Line
773 | OutPt InfusionCenter | tocilizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:55 |  | 04:00 | 01:00 | mg/kg | 3.9 |  | 8.1 | mg/mL | 1 | 1.1 | 8.1 | Dedicated IV Line
774 | OutPt InfusionCenter | ustekinumab |  | 260 | mg | 250 | mL | 1.04 | mg/mL | Pump | Primary Only | 00:55 |  | 03:00 | 01:00 | mg | 259 |  | 521 | n/a |  |  |  | Low Binding Filter
775 | OutPt InfusionCenter | ustekinumab |  | 390 | mg | 250 | mL | 1.56 | mg/mL | Pump | Primary Only | 00:55 |  | 03:00 | 01:00 | mg | 259 |  | 521 | n/a |  |  |  | Low Binding Filter
776 | OutPt InfusionCenter | ustekinumab |  | 520 | mg | 250 | mL | 2.08 | mg/mL | Pump | Primary Only | 00:55 |  | 03:00 | 01:00 | mg | 259 |  | 521 | n/a |  |  |  | Low Binding Filter
777 | OutPt InfusionCenter | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 500 | 2000 |  | n/a
778 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
779 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
780 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
781 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
782 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | mg | 500 | 2000 |  | n/a
783 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg | 500 | 2000 |  | n/a
784 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg | 500 | 2000 |  | n/a
785 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
786 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
787 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
788 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
789 | OutPt InfusionCenter | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | mg | 500 | 2000 |  | mg/mL | 1 | 2.9 | 5.1
790 | OutPt InfusionCenter | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg/kg | 5 |  | 30 | n/a
791 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
792 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
793 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
794 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
795 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
796 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
797 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
798 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
799 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
800 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
801 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
802 | OutPt InfusionCenter | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 |  | mg/kg | 5 |  | 30 | mg/mL | 1 | 3 | 5
803 | OutPt InfusionCenter | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.5 | 2 |  | n/a
804 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
805 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
806 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
807 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
808 | OutPt InfusionCenter | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | gram | 0.5 | 2 |  | n/a
809 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram | 0.5 | 2 |  | n/a
810 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram | 0.5 | 2 |  | n/a
811 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
812 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
813 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
814 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
815 | OutPt InfusionCenter | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.002 | 0.006
816 | OutPt InfusionCenter | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram/kg | 0.005 |  | 0.03 | n/a
817 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
818 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
819 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
820 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
821 | OutPt InfusionCenter | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
822 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
823 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
824 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
825 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
826 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
827 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
828 | OutPt InfusionCenter | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | gram/kg | 0.005 |  | 0.03 | gram/mL | 0.001 | 0.002 | 0.006
829 | OutPt InfusionCenter | vedolizumab |  | 300 | mg | 250 | mL | 1.2 | mg/mL | Pump | Primary and Secondary | 00:28 |  | 02:00 | 00:30 | mg | 299 |  | 301 | n/a |  |  |  | Flush After NS-LR
830 | OutPt InfusionCenter | zoledronic acid |  | 4 | mg | 100 | mL | 0.04 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | n/a
831 | OutPt InfusionCenter | zoledronic acid |  | 5 | mg | 100 | mL | 0.05 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | n/a
832 | OutPt InfusionCenter | zoledronic acid |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | mg/mL | 0.028 | 0.029 | 0.051
833 | Pediatrics | acetaminophen |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 999 | 1001 |  | n/a
834 | Pediatrics | acetaminophen |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg/kg | 7 |  | 16 | mg/mL | 9 |  | 11
835 | Pediatrics | acetylcysteine | 1st Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg/kg | 142.5 |  | 157.5 | mg/mL |  | 1 | 75
836 | Pediatrics | acetylcysteine | 1st Dose >100kg | 15000 | mg | 275 | mL | 54.5455 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 14999 |  | 15001 | n/a
837 | Pediatrics | acetylcysteine | 2nd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg/kg | 47.5 |  | 52.5 | mg/mL |  | 1 | 75
838 | Pediatrics | acetylcysteine | 2nd Dose >100kg | 5000 | mg | 525 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg | 4999 |  | 5001 | n/a
839 | Pediatrics | acetylcysteine | 3rd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg/kg | 95 |  | 105 | mg/mL |  | 1 | 75
840 | Pediatrics | acetylcysteine | 3rd Dose >100kg | 10000 | mg | 1050 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg | 9999 |  | 10001 | n/a
841 | Pediatrics | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 4 | 25 |  | mg/mL | 0.1 | 1 | 7
842 | Pediatrics | albumin 25% |  | 5 | gram | 20 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:08 | 01:00 | 00:20 | gram | 4.9 | 25.1 |  | n/a
843 | Pediatrics | albumin 25% |  | 12.5 | gram | 50 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:16 | 02:00 | 00:50 | gram | 4.9 | 25.1 |  | n/a
844 | Pediatrics | albumin 25% |  | 25 | gram | 100 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:32 | 04:00 | 01:40 | gram | 4.9 | 25.1 |  | n/a
845 | Pediatrics | ampicillin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:30 |  | mg/kg | 4 | 100 |  | mg/mL | 1 | 5 | 35
846 | Pediatrics | ampicillin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:30 | 00:30 | mg | 500 | 3000 |  | mg/mL | 1 | 5 | 35
847 | Pediatrics | ampicillin/sulbact | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 |  | mg/kg | 4 | 80 |  | mg/mL | 5 | 15 | 35
848 | Pediatrics | ampicillin/sulbact | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 |  | mg | 1000 | 3000 |  | mg/mL | 5 | 15 | 35
849 | Pediatrics | antivenin crotalidae |  | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 01:00 | 11:30 | 10:00 | unit | 1 | 6 |  | unit/mL | 0.004 |  | 0.024 | Antivenin
850 | Pediatrics | azithromycin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 3 | 14 |  | mg/mL | 0.1 | 1 | 2.2
851 | Pediatrics | azithromycin | pt weight >= 40 kg | 250 | mg | 250 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:40 | 02:00 | 01:00 | mg | 100 | 600 |  | n/a
852 | Pediatrics | azithromycin | pt weight >= 40 kg | 500 | mg | 250 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:40 | 02:00 | 01:00 | mg | 100 | 600 |  | n/a
853 | Pediatrics | aztreonam | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:15 | 02:00 |  | mg/kg | 25.5 | 50 |  | mg/mL | 4 | 5 | 40
854 | Pediatrics | aztreonam | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:15 | 02:00 |  | mg | 500 | 2000 |  | mg/mL | 4 | 5 | 40
855 | Pediatrics | caffeine citrate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 06:00 |  | mg | 1 | 1200 |  | mg/mL |  | 0.01 | 10
856 | Pediatrics | caffeine citrate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 06:00 |  | mg/kg | 1 | 22 |  | mg/mL |  | 0.01 | 10
857 | Pediatrics | caffeine/sod benzoa |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 04:00 | 01:00 | mg | 50 | 501 | 1001 | mg/mL |  | 0.24 | 6 | 0.22 Filter Needed
858 | Pediatrics | caffeine/sod benzoa |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 04:00 | 01:00 | mg/kg | 1 | 22 |  | mg/mL |  | 0.24 | 6 | 0.22 Filter Needed
859 | Pediatrics | ceFAZolin | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg/kg | 9 | 55 |  | mg/mL | 1 | 5 | 40
860 | Pediatrics | ceFAZolin | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
861 | Pediatrics | ceFAZolin | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram/kg | 0.009 | 0.055 |  | gram/mL | 0.001 | 0.005 | 0.04
862 | Pediatrics | ceFAZolin | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
863 | Pediatrics | cefepime | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 15 | 55 |  | mg/mL | 1 | 5 | 40
864 | Pediatrics | cefepime | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
865 | Pediatrics | cefepime | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | gram/kg | 0.015 | 0.055 |  | gram/mL | 0.001 | 0.005 | 0.04
866 | Pediatrics | cefepime | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
867 | Pediatrics | cefotaxime | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 22 | 80 |  | mg/mL | 5 | 10 | 55
868 | Pediatrics | cefotaxime | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 55
869 | Pediatrics | cefotaxime | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram/kg | 0.022 | 0.08 |  | gram/mL | 0.005 | 0.01 | 0.055
870 | Pediatrics | cefotaxime | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.055
871 | Pediatrics | cefOXitin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg/kg | 20 | 55 |  | mg/mL | 5 | 10 | 40
872 | Pediatrics | cefOXitin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 40
873 | Pediatrics | cefTAZidime | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | mg/kg | 25 | 67 | 135 | mg/mL | 1 | 5 | 50
874 | Pediatrics | cefTAZidime | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | mg | 250 | 2000 |  | mg/mL | 1 | 5 | 50
875 | Pediatrics | cefTAZidime | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | gram/kg | 0.025 | 0.067 | 0.135 | gram/mL | 0.001 | 0.005 | 0.05
876 | Pediatrics | cefTAZidime | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | gram | 0.25 | 2 |  | gram/mL | 0.001 | 0.005 | 0.05
877 | Pediatrics | cefTRIAXone | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg/kg | 21 | 105 |  | mg/mL | 5 | 10 | 50
878 | Pediatrics | cefTRIAXone | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 50
879 | Pediatrics | cefTRIAXone | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | gram/kg | 0.021 | 0.105 |  | gram/mL | 0.005 | 0.01 | 0.05
880 | Pediatrics | cefTRIAXone | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.05
881 | Pediatrics | cefuroxime | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg/kg | 18 | 55 |  | mg/mL | 5 | 7.5 | 30
882 | Pediatrics | cefuroxime | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg | 750 | 1500 |  | mg/mL | 5 | 7.5 | 30
883 | Pediatrics | chlorothiazide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 250 | 1000 |  | mg/mL | 0.5 | 1 | 30
884 | Pediatrics | chlorothiazide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg/kg | 0.4 | 11 |  | mg/mL | 0.5 | 1 | 30
885 | Pediatrics | clindamycin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 4 | 23 |  | mg/mL | 1 | 2 | 20
886 | Pediatrics | clindamycin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 300 | 2400 |  | mg/mL | 1 | 2 | 20
887 | Pediatrics | erythromycin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg/kg | 2.5 | 15 |  | mg/mL | 1 | 2 | 5.5
888 | Pediatrics | erythromycin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg | 100 | 1000 |  | mg/mL | 1 | 2 | 5.5
889 | Pediatrics | fluconazole | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 02:00 | 01:00 | mg/kg | 3 | 12 | 14 | mg/mL | 1.5 |  | 2.5
890 | Pediatrics | fluconazole | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg | 100 | 800 |  | mg/mL | 1.5 |  | 2.5
891 | Pediatrics | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:08 | 01:00 | 00:30 | mg | 100 | 1500 |  | mg/mL | 0.9 | 1 | 30
892 | Pediatrics | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:08 | 01:00 | 00:30 | mg/kg | 2 |  | 29 | mg/mL | 0.9 | 1 | 30
893 | Pediatrics | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:15 | 00:30 | mg/kg | 1 | 10 |  | mg/mL | 0.1 | 0.5 | 11
894 | Pediatrics | iron sucrose | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 01:35 |  | mg/kg | 0.4 | 7 | 7.3 | mg/mL | 0.9 | 1 | 2
895 | Pediatrics | iron sucrose | .mg, pt weight>=40kg | 100 | mg | 55 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 00:30 | mg | 90 |  | 315 | n/a
896 | Pediatrics | iron sucrose | .mg, pt weight>=40kg | 100 | mg | 105 | mL | 0.9524 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 00:30 | mg | 90 |  | 315 | n/a
897 | Pediatrics | iron sucrose | .mg, pt weight>=40kg | 200 | mg | 110 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 00:55 | 00:57 | 03:00 | 01:00 | mg | 90 |  | 315 | n/a
898 | Pediatrics | iron sucrose | .mg, pt weight>=40kg | 300 | mg | 265 | mL | 1.1321 | mg/mL | Pump | Primary and Secondary | 01:20 | 01:25 | 04:00 | 01:30 | mg | 90 |  | 315 | n/a
899 | Pediatrics | levETIRAcetam | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:08 | 00:10 | 01:00 |  | mg/kg | 4.3 | 57 | 115 | mg/mL | 0.5 | 1 | 30
900 | Pediatrics | levETIRAcetam | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:08 | 00:10 | 01:00 |  | mg | 200 | 3000 | 5000 | mg/mL | 0.5 | 1 | 30
901 | Pediatrics | mannitol 20% |  | 100 | gram | 500 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 00:31 | 02:00 |  | gram | 1 | 101 |  | n/a
902 | Pediatrics | mannitol 20% |  | 100 | gram | 500 | mL | 0.2 | gram/mL | Pump | Primary and Secondary | 00:25 | 00:31 | 02:00 |  | gram/kg | 0.2 | 2 |  | n/a |  |  |  | In-line filter
903 | Pediatrics | mannitol 20% |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:10 | 02:00 |  | gram/kg | 0.2 | 2 |  | gram/mL |  | 0.1 | 0.5 | In-line filter
904 | Pediatrics | meropenem | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | mg/kg | 10 | 40 |  | mg/mL | 1 | 5 | 40
905 | Pediatrics | meropenem | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
906 | Pediatrics | meropenem | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | gram/kg | 0.01 | 0.04 |  | gram/mL | 0.001 | 0.005 | 0.04
907 | Pediatrics | meropenem | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
908 | Pediatrics | metroNIDAZOLE |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 03:00 | 01:00 | mg | 250 | 500 |  | n/a
909 | Pediatrics | metroNIDAZOLE |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 03:00 | 01:00 | mg | 250 | 500 |  | n/a
910 | Pediatrics | metroNIDAZOLE |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 03:00 | 01:00 | mg/kg | 7 | 16 |  | mg/mL | 4 |  | 6
911 | Pediatrics | micafungin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 0.85 | 7 |  | mg/mL | 0.4 |  | 4
912 | Pediatrics | micafungin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg | 40 | 150 |  | mg/mL | 0.4 |  | 4
913 | Pediatrics | nafcillin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 10.6 | 57.5 |  | mg/mL | 1 | 5 | 40
914 | Pediatrics | nafcillin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
915 | Pediatrics | oxacillin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:08 | 01:30 |  | mg/kg | 5 | 50 |  | mg/mL | 5 | 10 | 40
916 | Pediatrics | oxacillin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:08 | 01:30 |  | mg | 500 | 2000 |  | mg/mL | 5 | 10 | 40
917 | Pediatrics | penG x Million Units |  | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | unit | 1 | 6 |  | unit/mL | 0.01 | 0.02 | 0.12 | 1unit=1million
918 | Pediatrics | pentamidine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 04:30 | 02:00 | mg/kg | 2.9 |  | 4.1 | mg/mL |  | 0.6 | 30
919 | Pediatrics | PHENobarbital |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:08 | 02:00 |  | mg/kg | 0.5 | 7.5 |  | mg/mL | 0.1 | 0.5 | 135 | High Risk
920 | Pediatrics | phenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:15 | 02:00 |  | mg/kg | 1 |  | 29 | mg/mL | 1 | 4 | 55 | In-line filter
921 | Pediatrics | piperacillin/tazo | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | mg/kg | 20 | 140 |  | mg/mL | 0.5 | 10 | 100
922 | Pediatrics | piperacillin/tazo | GM dosing | 2.25 | gram | 50 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
923 | Pediatrics | piperacillin/tazo | GM dosing | 2.25 | gram | 100 | mL | 0.0225 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
924 | Pediatrics | piperacillin/tazo | GM dosing | 3.375 | gram | 50 | mL | 0.0675 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
925 | Pediatrics | piperacillin/tazo | GM dosing | 3.375 | gram | 100 | mL | 0.0338 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
926 | Pediatrics | piperacillin/tazo | GM dosing | 4.5 | gram | 50 | mL | 0.09 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
927 | Pediatrics | piperacillin/tazo | GM dosing | 4.5 | gram | 100 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
928 | Pediatrics | piperacillin/tazo | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 04:30 |  | gram/kg | 0.025 | 0.14 |  | gram/mL | 0.005 | 0.01 | 0.1
929 | Pediatrics | potassium chloride |  | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:45 | 00:50 | 06:00 |  | mEq | 1 |  | 41 | mEq/mL | 0.04 | 0.05 | 0.2
930 | Pediatrics | potassium chloride |  | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:45 | 00:50 | 06:00 |  | mEq/kg | 0.2 | 1.5 | 2.5 | mEq/mL | 0.04 | 0.05 | 0.2 | High Risk
931 | Pediatrics | potassium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 01:30 | 03:00 | 12:30 |  | mmol | 5 | 20 | 40 | mmol/mL | 0.009 | 0.01 | 0.3
932 | Pediatrics | raNITIdine | pt weight < 50 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:10 | 01:00 | 00:30 | mg/kg | 0.4 | 2.5 |  | mg/mL | 0.1 | 0.5 | 3
933 | Pediatrics | raNITIdine | pt weight >= 50 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:10 | 01:00 |  | mg | 25 | 50 | 150 | mg/mL | 0.1 | 0.5 | 3
934 | Pediatrics | rifAMPin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 03:00 |  | mg/kg | 4.3 | 23 |  | mg/mL | 0.5 | 1 | 7
935 | Pediatrics | rifAMPin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 03:00 |  | mg | 300 | 600 |  | mg/mL | 0.5 | 1 | 7
936 | Pediatrics | SMX/TMP |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:30 | mg/kg | 1.7 | 8 | 10 | mg/mL | 0.01 | 0.04 | 1 | 48 Hour Tubing
937 | Pediatrics | sodium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 00:45 | 01:00 | 10:30 |  | mmol | 1 |  | 31 | mmol/mL | 0.009 | 0.01 | 0.3 | Calcium Precipitate
938 | Pediatrics | sodium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 00:45 | 01:00 | 10:30 |  | mmol/kg | 0.04 | 0.4 | 0.6 | mmol/mL | 0.009 | 0.01 | 0.3 | Calcium Precipitate
939 | Pediatrics | tobramycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:15 |  | mg/kg | 0.8 | 7 | 12 | mg/mL | 0.01 | 0.1 | 11
940 | Pediatrics | tocilizumab |  | 800 | mg | 100 | mL | 8 | mg/mL | Pump | Primary and Secondary | 00:55 |  | 04:00 | 01:00 | mg | 161 |  | 801 | n/a
941 | Pediatrics | tocilizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 00:55 |  | 04:00 | 01:00 | mg/kg | 3.9 |  | 12.1 | mg/mL | 1 | 1.1 | 8.1
942 | Pediatrics | valproate sodium |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 |  | mg/kg | 5 | 60 |  | mg/mL | 1 |  | 25
943 | Pediatrics | vancomycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:30 | 00:50 | 04:30 |  | mg/kg | 5 | 30 |  | mg/mL | 0.5 | 1 | 5
944 | Pediatrics | voriconazole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:10 |  | mg/kg | 2.6 | 9 |  | mg/mL | 0.1 | 0.4 | 5.5
945 | Pediatrics | zidovudine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 0.8 | 9 |  | mg/mL | 0.5 | 1 | 4
Note on Y26: Existing IV lines must be flushed with D5w prior to infusion. If unable, must start a new line. **An in-fine filter >= 1 micron may be used with infusion.
Note on Y46: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y47: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y48: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y49: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y50: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y51: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y71: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y73: Use a 0.22 micron in-line filter when administering this product.
Note on Y74: Use a 0.22 micron in-line filter when administering this product.
Note on Y93: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y94: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y95: Do not infuse any other drugs through this line.
Note on Y96: Do not infuse any other drugs through this line.
Note on Y97: Do not infuse any other drugs through this line.
Note on Y200: Advise patient to lie down during and 30 minutes after infusion to prevent hypotension.
Note on Y212: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y213: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y214: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y215: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y216: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y217: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y218: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y219: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y220: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y221: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y222: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y223: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y236: Protect this drug from light.
Note on Y237: Protect this drug from light.
Note on Y254: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y255: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y256: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y257: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y258: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y259: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y260: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y261: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y270: Infusion requires the use of an in-line filter.
Note on Y284: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y285: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y286: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y287: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y288: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y289: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y290: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y291: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y292: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y293: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y294: Infusion requires the use of an in-line filter.
Note on Y295: Infusion requires the use of an in-line filter.
Note on Y296: INR 2 - 4 ** Max dose of Kcentra should not exceed 2500 units **
Note on Y297: INR 4 - 6 ** Max dose of Kcentra should not exceed 3500 units **
Note on Y298: INR > 6 ** Max dose of Kcentra should not exceed 5000 units **
Note on Y311: Infusion requires the use of an in-line filter.
Note on Y312: Infusion requires the use of an in-line filter.
Note on Y313: Infusion requires the use of an in-line filter.
Note on Y359: Do not infuse any other drugs through this line.
Note on Y361: 1 unit = 1 million units
Note on Y362: 1 unit = 1 million units
Note on Y363: 1 unit = 1 million units
Note on Y364: 1 unit = 1 million units
Note on Y365: 1 unit = 1 million units
Note on Y366: 1 unit = 1 million units
Note on Y372: Infusion requires the use of an in-line filter.
Note on Y408: Premedicate before each infusion.
Note on Y409: Change tubing every 48 hours.
Note on Y411: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y412: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y423: Do not infuse any other drugs through this line.
Note on Y424: Do not infuse any other drugs through this line.
Note on Y490: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y491: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y492: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y493: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y494: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y495: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y500: Adminster over 60 minutes to decrease risk of bilirubin encephalopathy.
Note on Y504: 1 unit in pump = 1,000 units on label. Eg. 200 units in pump = 200,000 units on drug label.
Note on Y507: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y508: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y509: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y510: Use a 0.22 micron in-line filter when administering this product.
Note on Y511: This medication must be infused using DEHP FREE tubing.
Note on Y512: This medication must be infused using DEHP FREE tubing.
Note on Y523: MAXIMUM dose of brentuximab is 180 mg.
Note on Y524: Use a 0.22 micron in-line filter when administering this product.
Note on Y528: Infusion requires the use of an in-line filter.
Note on Y529: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y530: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y532: Use a 0.22 micron in-line filter when administering this product.
Note on Y536: Use a 0.22 micron in-line filter when administering this product.
Note on Y537: Use a 0.22 micron in-line filter when administering this product.
Note on Y538: Use a 0.22 micron in-line filter when administering this product.
Note on Y541: This medication must be infused using DEHP FREE tubing.
Note on Y543: Do NOT use an in-line filter for administration.
Note on Y544: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y545: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y548: Do NOT administer through lines containing dextrose.
Note on Y549: This medication must be infused using Non-PVC tubing.
Note on Y557: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y561: Do NOT use an in-line filter for administration.
Note on Y570: Infusion requires the use of an in-line filter.
Note on Y571: Infusion requires the use of an in-line filter.
Note on Y572: Infusion requires the use of an in-line filter.
Note on Y573: Infusion requires the use of an in-line filter.
Note on Y574: Infusion requires the use of an in-line filter.
Note on Y582: Requires Low protein binding 0.2 micron filter and DEHP free bag
Note on Y583: Do NOT use an in-line filter for administration.
Note on Y584: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y585: Infusion requires the use of an in-line filter.
Note on Y586: Infusion requires the use of an in-line filter.
Note on Y590: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y591: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y593: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y594: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y595: Flush line before and after administration. Can be given on same line as sodium chloride 0.9%.
Note on Y597: Use a 0.22 micron in-line filter when administering this product.
Note on Y600: Drug is a vesicant. Check blood return.
Note on Y601: Drug is a vesicant. Check blood return.
Note on Y603: Use a 0.22 micron in-line filter when administering this product.
Note on Y605: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y606: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y607: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y608: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y613: Existing IV lines must be flushed with D5w prior to infusion. If unable, must start a new line. **An in-fine filter >= 1 micron may be used with infusion.
Note on Y615: Use a low-protein binding in-line or add-on filter between 0.2 and 5 microns.
Note on Y668: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y669: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y670: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y671: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y672: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y673: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y674: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y675: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y687: Protect this drug from light.
Note on Y688: Protect this drug from light.
Note on Y703: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y704: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y705: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y706: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y709: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y710: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y711: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y712: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y713: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y714: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y715: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y716: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y717: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y718: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y744: Infusion requires the use of an in-line filter.
Note on Y745: Infusion requires the use of an in-line filter.
Note on Y746: Infusion requires the use of an in-line filter.
Note on Y747: Infusion requires the use of an in-line filter.
Note on Y748: Infusion requires the use of an in-line filter.
Note on Y749: Use a 0.22 micron in-line filter when administering this product.
Note on Y750: Use a 0.22 micron in-line filter when administering this product.
Note on Y753: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w or NS.
Note on Y754: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w or NS.
Note on Y755: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y756: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y769: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y770: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y771: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y774: Do not infuse any other drugs through this line.
Note on Y775: Do not infuse any other drugs through this line.
Note on Y776: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y777: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y778: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y831: Following infusion, flush with 30 mL of sterile normal saline or lactated ringers.
Note on Y851: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y859: Use a 0.22 micron in-line filter when administering this product.
Note on Y860: Use a 0.22 micron in-line filter when administering this product.
Note on Y904: Infusion requires the use of an in-line filter.
Note on Y905: Infusion requires the use of an in-line filter.
Note on Y919: 1 unit = 1 million units
Note on Y921: This is a HIGH RISK medication. 2nd RN to verify prior to initiating therapy.
Note on Y922: Infusion requires the use of an in-line filter.
Note on Y932: This is a HIGH RISK medication. 2nd RN to verify prior to initiating therapy.
Note on Y938: Change tubing every 48 hours.
Note on Y939: Do not infuse Calcium and Phosphate though the same line.
Note on Y940: Do not infuse Calcium and Phosphate though the same line.

## Fluids
Fluids
Line Item | Profile | Fluid Name | Therapy Name | Module Availability | Supports Secondary (Y/N) | Rate Units | Rate Limits (mL/h) Soft Min | Rate Limits (mL/h) Soft Max | Rate Limits (mL/h) Hard Max | Clinical Advisory
1 | Adult | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 999
2 | Adult | .IVF + KCL | 10 mEq/L | Pump | Y | mL/h | 1 | 999
3 | Adult | .IVF + KCL | 15 mEq/L | Pump | Y | mL/h | 1 | 666 | 700
4 | Adult | .IVF + KCL | 20 mEq/L | Pump | Y | mL/h | 1 | 500 | 550
5 | Adult | .IVF + KCL | 30 mEq/L | Pump | Y | mL/h | 1 | 333 | 366
6 | Adult | .IVF + KCL | 40 mEq/L | Pump | Y | mL/h | 1 | 250 | 275
7 | Adult | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 1 | 999
8 | Adult | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 500 | 550
9 | Adult | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 1 | 250 | 275
10 | Adult | .Maint IVF+Additives |  | Pump | Y | mL/h | 1 |  | 999
11 | Adult | 3% Sodium Chloride |  | Pump | Y | mL/h | 1 | 50 | 100
12 | Adult | Bicarb Drip |  | Pump | Y | mL/h | 1 |  | 999
13 | Adult | Blood (whole) |  | Pump | N | mL/h | 1 | 301 |  | Dedicated IV Line
14 | Adult | Dextrose 10% |  | Pump | Y | mL/h | 1 | 150
15 | Adult | Fresh Frozen Plasma |  | Pump | N | mL/h | 1 | 201 |  | Dedicated IV Line
16 | Adult | Insulin Maternity |  | Pump | Y | mL/h | 70 | 160 | 200
17 | Adult | Parenteral Nutrition |  | Pump | N | mL/h | 19 | 95 |  | Midline Warning
18 | Adult | PRBCs |  | Pump | N | mL/h | 1 | 301 |  | Dedicated IV Line
19 | Nursery | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 15 | 25
20 | Nursery | Dextrose 10% |  | Pump | Y | mL/h | 1 | 15 | 25
21 | ONC Infusion Center | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 999
22 | ONC Infusion Center | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 1 | 999
23 | ONC Infusion Center | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 667
24 | ONC Infusion Center | .IVF + KCL 30 mEq/L |  | Pump | Y | mL/h | 1 | 334 | 667
25 | ONC Infusion Center | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 1 | 250 | 500
26 | ONC Infusion Center | .Maint IVF+Additives |  | Pump | Y | mL/h | 1 | 250
27 | ONC Infusion Center | Dextrose 10% |  | Pump | Y | mL/h | 15 | 200
28 | ONC Infusion Center | Rituximab (Rituxan) |  | Pump | Y | mL/h | 0.1 | 250
29 | OutPt InfusionCenter | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 999
30 | OutPt InfusionCenter | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 1 | 999
31 | OutPt InfusionCenter | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 667
32 | OutPt InfusionCenter | .IVF + KCL 30 mEq/L |  | Pump | Y | mL/h | 1 | 334 | 667
33 | OutPt InfusionCenter | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 1 | 250 | 500
34 | OutPt InfusionCenter | .Maint IVF+Additives |  | Pump | Y | mL/h | 1 | 250
35 | OutPt InfusionCenter | Dextrose 10% |  | Pump | Y | mL/h | 15 | 200
36 | OutPt InfusionCenter | PRBCs |  | Pump and Syringe | N | mL/h | 1 | 150 |  | Dedicated IV Line
37 | OutPt InfusionCenter | Rituximab (Rituxan) |  | Pump | Y | mL/h | 0.1 | 250
38 | Pediatrics | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 300
39 | Pediatrics | .IVF + KCL | 10 mEq/L | Pump | Y | mL/h | 1 | 999
40 | Pediatrics | .IVF + KCL | 15 mEq/L | Pump | Y | mL/h | 1 | 666 | 700
41 | Pediatrics | .IVF + KCL | 20 mEq/L | Pump | Y | mL/h | 1 | 500 | 550
42 | Pediatrics | .IVF + KCL | 30 mEq/L | Pump | Y | mL/h | 1 | 333 | 366
43 | Pediatrics | .IVF + KCL | 40 mEq/L | Pump | Y | mL/h | 1 | 250 | 275
44 | Pediatrics | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 0.1 | 200 | 999
45 | Pediatrics | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 200 | 500
46 | Pediatrics | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 0.1 | 200 | 275
47 | Pediatrics | Cryoprecipitate |  | Pump | N | mL/h | 0.1 | 300
48 | Pediatrics | Dextrose 10% |  | Pump | Y | mL/h | 1 | 50 | 150
49 | Pediatrics | Fresh Frozen Plasma |  | Pump | N | mL/h | 0.1 | 150 |  | Dedicated IV Line
50 | Pediatrics | Parenteral Nutrition |  | Pump | N | mL/h | 35 | 95 |  | Midline Warning
51 | Pediatrics | PRBCs |  | Pump | N | mL/h | 0.1 | 150 |  | Dedicated IV Line
Note on K15: Do not infuse any other drugs through this line.
Note on K17: Do not infuse any other drugs through this line.
Note on K19: Do NOT infuse through a midline catheter.
Note on K20: Do not infuse any other drugs through this line.
Note on K38: Do not infuse any other drugs through this line.
Note on K51: Do not infuse any other drugs through this line.
Note on K52: Do NOT infuse through a midline catheter.
Note on K53: Do not infuse any other drugs through this line.

## PCA Drugs
PCA Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount | Drug Amount Units | Diluent Volume | Diluent Volume Units | Final Concentration | Final Concentration Units | Max Accumulated Lockout | Dosing Units | Max Accumulated Dose Range Soft Min | Max Accumulated Dose Range Soft Max | Max Accumulated Dose Range Hard Max | Max Accumulated Dose Range Initial Value | PCA Dose Limits Soft Min | PCA Dose Limits Soft Max | PCA Dose Limits Hard Max | PCA Dose Limits Initial Value | Lock Out Interval (minutes) Hard Min | Lock Out Interval (minutes) Soft Min | Lock Out Interval (minutes) Soft Max | Lock Out Interval (minutes) Initial Value | Continuous Dose Limits Soft Min | Continuous Dose Limits Soft Max | Continuous Dose Limits Hard Max | Continuous Dose Limits Initial Value | Bolus Dose Limits Soft Min | Bolus Dose Limits Soft Max | Bolus Dose Limits Hard Max | Bolus Dose included in Accum Max Dose (Y/N) | Loading Dose Limits Soft Min | Loading Dose Limits Soft Max | Loading Dose Limits Hard Max | Conc. Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max | Supports PCA Pause Protocol (Y/N) | Clinical Advisory
1 | Adult | No drugs exist in this section
2 | Nursery | No drugs exist in this section
3 | ONC Infusion Center | No drugs exist in this section
4 | OutPt InfusionCenter | No drugs exist in this section
5 | Pediatrics | No drugs exist in this section

## PC Unit Config
| PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit
Line Item | Profile | Alarm Audio Profile | Minimum Audio Volume | Default Audio Volume | Anesthesia Mode | Battery Meter Display | Color Display | Key Click Audio | Limit Checking | Max Patient BSA (m2) | Max Patient Weight (kg) | Patient ID Entry | Preventive Maintenance Reminder | Tamper Resistant | Authorized User Mode | No Guardrails Basic Infusion Clinical Advisory | Clinical Advisory Message
1 | Adult | Profile 1 | 3 | 4 | Enabled | Enabled | Enabled | Enabled | Always | 3 | 500 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
2 | Nursery | Profile 1 | 2 | 3 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 7 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
3 | ONC Infusion Center | Profile 1 | 3 | 4 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 500 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
4 | OutPt InfusionCenter | Profile 1 | 3 | 4 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 500 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
5 | Pediatrics | Profile 1 | 3 | 4 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 250 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection

## Pump Config
| Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module
Line Item | Profile | Accumulated Air-In-Line | Air In Line (microliter) | Auto-Restart Attempts | KVO Rate Adjust (mL/h) | Max Occlusion Pressure (mmHg) | Default Pressure (mmHg) | Maximum Rate (mL/h) | Maximum VTBI (mL) | Pressure Mode Lock Status | Pressure Mode Selection | Secondary Infusion | SEC-PRI Audio Alert
1 | Adult | Enabled | 250 | 9 | 20 | 525 | 525 | 999 | 3000 | Unlocked | Pump | Enabled | Enabled
2 | Nursery | Enabled | 50 | 1 | 2 | 525 | 525 | 25 | 1200 | Unlocked | Pump | Enabled | Enabled
3 | ONC Infusion Center | Enabled | 250 | 9 | 20 | 525 | 525 | 999 | 3000 | Unlocked | Pump | Enabled | Enabled
4 | OutPt InfusionCenter | Enabled | 250 | 9 | 20 | 525 | 525 | 999 | 3000 | Unlocked | Pump | Enabled | Enabled
5 | Pediatrics | Enabled | 75 | 9 | 5 | 525 | 525 | 999 | 2200 | Unlocked | Pump | Enabled | Enabled

## Syringe Config
| Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module
Line Item | Profile | All Mode | Auto Occlusion Pressure | Back-Off After Occlusion | Fast-Start | KVO Option | KVO Rate Adjust (mL/h) | KVO Volume Adjust (%) | Near End of Infusion (Continuous, Fluids, Basic) | Alert Time (Continuous, Fluids, Basic) (minutes) | Near End of Infusion (Intermittent) | Alert Time (Intermittent) (minutes) | Occlusion Pressure Limit with Disc (mmHg) | Occlusion Pressure Limit No Disc | Priming
1 | Adult | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled
2 | Nursery | Enabled | Enabled | Enabled | Enabled | Enabled | 2 | 5 | Enabled | 2 | Enabled | 2 | 1000 | High | Enabled
3 | ONC Infusion Center | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled
4 | OutPt InfusionCenter | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled
5 | Pediatrics | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled

## Pump & Syringe Shared Config
| Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe)
Line Item | Profile | Delay Start Options | Delay Start Callback | Drug Calculation | Drug Calculation Bolus Dose | Pressure Dynamic Display | Volume/Duration
1 | Adult | Disabled | n/a | Enabled | Enabled | Enabled | Enabled
2 | Nursery | Disabled | n/a | Enabled | Enabled | Enabled | Enabled
3 | ONC Infusion Center | Disabled | n/a | Disabled | n/a | Enabled | Enabled
4 | OutPt InfusionCenter | Disabled | n/a | Disabled | n/a | Enabled | Enabled
5 | Pediatrics | Disabled | n/a | Enabled | Enabled | Enabled | Enabled

## PCA Config
| PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module
Line Item | Profile | Bolus Delivery Rate (mL/h) | Dose Request Cord (Profile) | Max Dose Limit | Time Window (h) | Module Location Enforcement | Near End of Infusion | Near End of Infusion Alert Time (%) | Pressure Limit | Priming | Security Access Level | Security Code
1 | Adult | 150 | 2 | Enabled | 1 | Enabled | Enabled | 5 | High | Enabled | 1 | n/a
2 | Nursery | 150 | 2 | Enabled | 4 | Enabled | Disabled | n/a | High | Enabled | 1 | n/a
3 | ONC Infusion Center | 150 | 2 | Enabled | 4 | Enabled | Disabled | n/a | High | Enabled | 1 | n/a
4 | OutPt InfusionCenter | 150 | 2 | Enabled | 4 | Enabled | Disabled | n/a | High | Enabled | 1 | n/a
5 | Pediatrics | 150 | 2 | Enabled | 1 | Enabled | Enabled | 5 | High | Enabled | 1 | n/a

## PCA Pause Protocol
| PCA Pause Protocol | PCA Pause Protocol | PCA Pause Protocol | PCA Pause Protocol | PCA Pause Protocol
Line Item | Profile | PCA Pause Protocol Enabled | Available For | Monitoring Module Attach Enforcement | PCA Pause Protocol
1 | Adult | Enabled | Etco2 | Disabled | PCA Infusion has paused due to a decline in respiratory status. Check Patient!
2 | Nursery | Disabled | n/a | n/a | n/a
3 | ONC Infusion Center | Disabled | n/a | n/a | n/a
4 | OutPt InfusionCenter | Disabled | n/a | n/a | n/a
5 | Pediatrics | Disabled | n/a | n/a | n/a

## PCA & Syringe Shared Config
| Shared Infusion Settings (PCA & Syringe) | Shared Infusion Settings (PCA & Syringe)
Line Item | Profile | Maximum Rate (mL/h) | NEOI Snooze | NEOI Snooze Time (minutes)
1 | Adult | 999 | Disabled | n/a
2 | Nursery | 25 | Enabled | 5
3 | ONC Infusion Center | 999 | Disabled | n/a
4 | OutPt InfusionCenter | 999 | Disabled | n/a
5 | Pediatrics | 999 | Disabled | n/a

## EtCO2 Config
| EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module
Line Item | Profile | Alarm Limit Type | EtCO2 High (mmHg) | EtCO2 Low (mmHg) | Respiratory Rate High (bpm) | Respiratory Rate Low (bpm) | No Breath Alarm (Seconds) | FiCO2 High (mmHg) | Waveform Time Scale (Seconds) | PCA Pause Resp. Rate Lower Limit (bpm) | PCA Pause Resp. Rate Initial Value (bpm)
1 | Adult | Adult | 60 | 10 | 35 | 7 | 30 | 8 | 10 | 5 | 6
2 | Nursery | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a
3 | ONC Infusion Center | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a
4 | OutPt InfusionCenter | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a
5 | Pediatrics | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a

## Master Drugs
Master Drugs | Master Drugs | Master Drugs | Master Drugs | Master Drugs
Line Item | Drug Name | Concentration | Used As | Aliases
1 | abatacept | 500 mg / 100 mL (5 mg / mL) | Intermittent | 4144595I
2 | abatacept | 750 mg / 100 mL (7.5 mg / mL) | Intermittent | 4144595I
3 | abatacept | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 4144595I
4 | abatacept | --- mg / --- mL | Intermittent | 4144595I
5 | abciximab | 7.2 mg / 250 mL (0.0288 mg / mL) |  | 2748001I
6 | abciximab | 7.4 mg / 250 mL (0.0296 mg / mL) |  | 2748001I
7 | abciximab | 9 mg / 250 mL (0.036 mg / mL) |  | 2748001I
8 | abciximab | --- mg / --- mL |  | 2748001I
9 | acetaminophen | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2748022I
10 | acetaminophen | --- mg / --- mL | Intermittent | 2748022I
11 | acetaminophen | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2748022I
12 | acetaZOLAMIDE | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2748400I
13 | acetaZOLAMIDE | --- mg / --- mL | Intermittent | 2748400I
14 | acetylcysteine | 5,000 mg / 525 mL (9.5238 mg / mL) | Intermittent | 2748462I
15 | acetylcysteine | 10,000 mg / 1,050 mL (9.5238 mg / mL) | Intermittent | 2748462I
16 | acetylcysteine | 15,000 mg / 275 mL (54.5455 mg / mL) | Intermittent | 2748462I
17 | acetylcysteine | --- mg / --- mL | Intermittent | 2748462I
18 | acyclovir | --- mg / --- mL | Intermittent | 2748487I
19 | adenosine (CATH LAB) | 60 mg / 50 mL (1.2 mg / mL) | Cont/Bolus or PCA | 2748529C
20 | adenosine (CATH LAB) | 90 mg / 90 mL (1 mg / mL) | Cont/Bolus or PCA | 2748529C
21 | adenosine (CATH LAB) | 180 mg / 90 mL (2 mg / mL) | Cont/Bolus or PCA | 2748529C
22 | ado-trastuzumab | --- mg / --- mL | Intermittent | 24501603I
23 | aflibercept | --- mg / --- mL |  | 19949112I
24 | agalsidase beta | --- mg / --- mL |  | 2748541I
25 | albumin 25% | 5 gram / 20 mL (0.25 gram / mL) | Intermittent | 2748593I
26 | albumin 25% | 12.5 gram / 50 mL (0.25 gram / mL) | Intermittent | 2748593I
27 | albumin 25% | 25 gram / 100 mL (0.25 gram / mL) | Intermittent | 2748593I
28 | albumin 25% | --- gram / --- mL | Intermittent | 2748593I
29 | albumin 5% | 12.5 gram / 250 mL (0.05 gram / mL) |  | 2748593
30 | albumin 5% | 25 gram / 500 mL (0.05 gram / mL) |  | 2748593
31 | aldesleukin | --- unit / --- mL |  | 2748678I
32 | alemtuzumab | --- mg / --- mL |  | 2748690I
33 | alfentanil | --- mcg / --- mL |  | 2748710I
34 | allopurinol | --- mg / --- mL |  | 2748758I
35 | alprostadil | 500 mcg / 250 mL (2 mcg / mL) | Cont/Bolus or PCA | 2748821C
36 | alprostadil | --- mcg / --- mL | Cont/Bolus or PCA | 2748821C
37 | alteplase | 100 mg / 100 mL (1 mg / mL) |  | 2748851
38 | alteplase | --- mg / --- mL |  | 2748851
39 | alteplase -stroke | 81 mg / 81 mL (1 mg / mL) | Cont/Bolus or PCA | 2748851C
40 | alteplase -stroke | 90 mg / 90 mL (1 mg / mL) | Cont/Bolus or PCA | 2748851C
41 | alteplase -stroke | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2748851C
42 | alteplase -stroke | --- mg / --- mL | Cont/Bolus or PCA | 2748851C
43 | alteplase-non-stroke | 2 mg / 50 mL (0.04 mg / mL) | Intermittent | 2748851I
44 | alteplase-non-stroke | 2 mg / 100 mL (0.02 mg / mL) | Intermittent | 2748851I
45 | alteplase-non-stroke | 100 mg / 100 mL (1 mg / mL) | Intermittent | 2748851I
46 | alteplase-non-stroke | --- mg / --- mL | Intermittent | 2748851I
47 | amifostine | --- mg / --- mL |  | 2748941I
48 | amikacin | 250 mg / 250 mL (1 mg / mL) | Intermittent | 2748946I
49 | amikacin | --- mg / --- mL | Intermittent | 2748946I
50 | aminocaproic acid | --- mg / --- mL | Intermittent | 2748973I
51 | aminocaproic acid | 5 gram / 250 mL (0.02 gram / mL) | Intermittent | 2748973I
52 | aminocaproic acid | --- gram / --- mL | Intermittent | 2748973I
53 | aminophylline | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2748995C
54 | aminophylline | 500 mg / 500 mL (1 mg / mL) | Cont/Bolus or PCA | 2748995C
55 | aminophylline | 1,000 mg / 500 mL (2 mg / mL) | Cont/Bolus or PCA | 2748995C
56 | aminophylline | --- mg / --- mL | Cont/Bolus or PCA | 2748995C
57 | amiodarone | 150 mg / 100 mL (1.5 mg / mL) |  | 2749009
58 | amiodarone | 450 mg / 250 mL (1.8 mg / mL) |  | 2749009
59 | amiodarone | 900 mg / 500 mL (1.8 mg / mL) |  | 2749009
60 | amiodarone *Loading* | 150 mg / 100 mL (1.5 mg / mL) | Intermittent | 2749009I
61 | amiodarone *Loading* | 150 mg / 103 mL (1.4563 mg / mL) | Intermittent | 2749009I
62 | amiodarone Drip | 360 mg / 200 mL (1.8 mg / mL) | Cont/Bolus or PCA | 2749009C
63 | amiodarone Drip | 450 mg / 250 mL (1.8 mg / mL) | Cont/Bolus or PCA | 2749009C
64 | amiodarone Drip | 900 mg / 500 mL (1.8 mg / mL) | Cont/Bolus or PCA | 2749009C
65 | ammonium chloride | --- mEq / --- mL |  | 2749089I
66 | ampho B (LIPID) | --- mg / --- mL |  | 2749273I
67 | ampho B (liposomal) | --- mg / --- mL | Intermittent | 2749278I
68 | ampho B cholesteryl | --- mg / --- mL |  | 2749266I
69 | amphotericin B | --- mg / --- mL | Intermittent | 2749258I
70 | ampicillin | 250 mg / 50 mL (5 mg / mL) | Intermittent | 2749283I
71 | ampicillin | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2749283I
72 | ampicillin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2749283I
73 | ampicillin | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2749283I
74 | ampicillin | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2749283I
75 | ampicillin | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2749283I
76 | ampicillin | --- mg / --- mL | Intermittent | 2749283I
77 | ampicillin | 0.5 gram / 50 mL (0.01 gram / mL) | Intermittent | 2749283I
78 | ampicillin | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2749283I
79 | ampicillin | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2749283I
80 | ampicillin | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2749283I
81 | ampicillin | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2749283I
82 | ampicillin | --- gram / --- mL | Intermittent | 2749283I
83 | ampicillin/sulbact | 1,500 mg / 50 mL (30 mg / mL) | Intermittent | 2749313I
84 | ampicillin/sulbact | 1,500 mg / 100 mL (15 mg / mL) | Intermittent | 2749313I
85 | ampicillin/sulbact | 3,000 mg / 100 mL (30 mg / mL) | Intermittent | 2749313I
86 | ampicillin/sulbact | --- mg / --- mL | Intermittent | 2749313I
87 | ampicillin/sulbact | 1.5 gram / 50 mL (0.03 gram / mL) | Intermittent | 2749313I
88 | ampicillin/sulbact | 1.5 gram / 100 mL (0.015 gram / mL) | Intermittent | 2749313I
89 | ampicillin/sulbact | 3 gram / 100 mL (0.03 gram / mL) | Intermittent | 2749313I
90 | andulafungin | --- mg / --- mL | Intermittent | 4144679I
91 | antithymocyte EQUIN | --- mg / --- mL |  | 19949425I
92 | antithymocyte RABBI | --- mg / --- mL |  | 19949425
93 | antivenin crotalidae | 4 unit / 250 mL (0.016 unit / mL) | Intermittent | 2749465I
94 | antivenin crotalidae | 6 unit / 250 mL (0.024 unit / mL) | Intermittent | 2749465I
95 | antivenin crotalidae | 10 unit / 250 mL (0.04 unit / mL) | Intermittent | 2749465I
96 | antivenin crotalidae | --- unit / --- mL | Intermittent | 2749465I
97 | aprepitant | 100 mg / 100 mL (1 mg / mL) | Intermittent | 2749846I
98 | aprepitant | 130 mg / 100 mL (1.3 mg / mL) | Intermittent | 2749846I
99 | argatroban | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2749866C
100 | argatroban | 500 mg / 500 mL (1 mg / mL) | Cont/Bolus or PCA | 2749866C
101 | arsenic trioxide | --- mg / --- mL |  | 2749895I
102 | asparaginase | --- unit / --- mL |  | 2750061I
103 | atezolizumab | 800 mg / 250 mL (3.2 mg / mL) | Intermittent | 2554527703I
104 | atezolizumab | 840 mg / 250 mL (3.36 mg / mL) | Intermittent | 2554527703I
105 | atezolizumab | 1,200 mg / 250 mL (4.8 mg / mL) | Intermittent | 2554527703I
106 | atezolizumab | 1,680 mg / 250 mL (6.72 mg / mL) | Intermittent | 2554527703I
107 | atracurium | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2750296C
108 | atracurium | --- mg / --- mL | Cont/Bolus or PCA | 2750296C
109 | avibactam/ceftazidim | --- gram / --- mL | Intermittent | 271758255I
110 | azaCITIDine | --- mg / --- mL |  | 3316777I
111 | azaTHIOprine | --- mg / --- mL |  | 2750417I
112 | azithromycin | 250 mg / 250 mL (1 mg / mL) | Intermittent | 2750452I
113 | azithromycin | 500 mg / 250 mL (2 mg / mL) | Intermittent | 2750452I
114 | azithromycin | --- mg / --- mL | Intermittent | 2750452I
115 | aztreonam | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2750475I
116 | aztreonam | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2750475I
117 | aztreonam | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2750475I
118 | aztreonam | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2750475I
119 | aztreonam | --- mg / --- mL | Intermittent | 2750475I
120 | aztreonam | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2750475I
121 | aztreonam | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2750475I
122 | aztreonam | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2750475I
123 | aztreonam | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2750475I
124 | bamlaniv/estesivimab | 700 mg / 110 mL (6.3636 mg / mL) | Intermittent | 1000000200I
125 | bamlaniv/estesivimab | 700 mg / 160 mL (4.375 mg / mL) | Intermittent | 1000000200I
126 | bamlaniv/estesivimab | 700 mg / 210 mL (3.3333 mg / mL) | Intermittent | 1000000200I
127 | bamlaniv/estesivimab | 700 mg / 310 mL (2.2581 mg / mL) | Intermittent | 1000000200I
128 | bamlaniv/estesivimab | --- mg / --- mL | Intermittent | 1000000200I
129 | bamlanivimab | 700 mg / 200 mL (3.5 mg / mL) | Intermittent | 3749717983I
130 | bamlanivimab | 700 mg / 270 mL (2.5926 mg / mL) | Intermittent | 3749717983I
131 | basiliximab | 20 mg / 50 mL (0.4 mg / mL) |  | 2750634I
132 | basiliximab | --- mg / --- mL |  | 2750634I
133 | belimumab | --- mg / --- mL | Intermittent | 19949906I
134 | belinostat | --- mg / --- mL |  | 45868857I
135 | Bendamusti (Bendeka) | --- mg / --- mL | Intermittent | 17213979bend
136 | Bendamusti (Treanda) | --- mg / --- mL | Intermittent | 17213979trea
137 | bendamustine HCl | --- mg / --- mL |  | 17213979I
138 | bevacizumab | --- mg / --- mL | Intermittent | 3316812I
139 | bezlotoxumab | --- mg / --- mL | Intermittent | 2557014467I
140 | bivalirudin | 250 mg / 50 mL (5 mg / mL) | Cont/Bolus or PCA | 2751435C
141 | bivalirudin | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2751435C
142 | bivalirudin | 250 mg / 500 mL (0.5 mg / mL) | Cont/Bolus or PCA | 2751435C
143 | bivalirudin | 500 mg / 100 mL (5 mg / mL) | Cont/Bolus or PCA | 2751435C
144 | bivalirudin | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA | 2751435C
145 | bivalirudin | 1,250 mg / 250 mL (5 mg / mL) | Cont/Bolus or PCA | 2751435C
146 | bleomycin | --- unit / --- mL | Intermittent | 2751447I
147 | bortezomib | --- mg / --- mL |  | 2751476I
148 | brentuximab | --- mg / --- mL | Intermittent | 19950376I
149 | bretylium | 1,000 mg / 250 mL (4 mg / mL) |  | 2751503I
150 | bretylium | 2,000 mg / 250 mL (8 mg / mL) |  | 2751503I
151 | bretylium | 2,000 mg / 500 mL (4 mg / mL) |  | 2751503I
152 | bumetanide | 10 mg / 100 mL (0.1 mg / mL) |  | 2751668I
153 | bumetanide | 12.5 mg / 50 mL (0.25 mg / mL) |  | 2751668I
154 | bumetanide | 12.5 mg / 250 mL (0.05 mg / mL) |  | 2751668I
155 | bumetanide | 25 mg / 100 mL (0.25 mg / mL) |  | 2751668I
156 | bumetanide | --- mg / --- mL |  | 2751668I
157 | bumetanide continuos | 24 mg / 96 mL (0.25 mg / mL) | Cont/Bolus or PCA | 2751668C
158 | bumetanide continuos | --- mg / --- mL | Cont/Bolus or PCA | 2751668C
159 | bupivacaine | --- mg / --- mL |  | 2751679I
160 | busulfan | --- mg / --- mL |  | 2751784I
161 | cabazitaxel | --- mg / --- mL | Intermittent | 19950393I
162 | caffeine citrate | --- mg / --- mL | Intermittent | 2751845I
163 | caffeine/sod benzoa | 500 mg / 1,000 mL (0.5 mg / mL) | Intermittent | 2751881I
164 | caffeine/sod benzoa | --- mg / --- mL | Intermittent | 2751881I
165 | calcium chloride | 1,000 mg / 60 mL (16.6667 mg / mL) | Intermittent | 2752076I
166 | calcium chloride | 1,000 mg / 110 mL (9.0909 mg / mL) | Intermittent | 2752076I
167 | calcium chloride | 2,000 mg / 120 mL (16.6667 mg / mL) | Intermittent | 2752076I
168 | calcium chloride | --- mg / --- mL | Intermittent | 2752076I
169 | calcium chloride | 1 gram / 60 mL (0.0167 gram / mL) | Intermittent | 2752076I
170 | calcium chloride | 1 gram / 110 mL (0.0091 gram / mL) | Intermittent | 2752076I
171 | calcium chloride | 2 gram / 120 mL (0.0167 gram / mL) | Intermittent | 2752076I
172 | calcium chloride | --- gram / --- mL | Intermittent | 2752076I
173 | calcium chloride | --- mEq / --- mL | Intermittent | 2752076I
174 | calcium gluconate | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752096I
175 | calcium gluconate | 1,000 mg / 60 mL (16.6667 mg / mL) | Intermittent | 2752096I
176 | calcium gluconate | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752096I
177 | calcium gluconate | 1,000 mg / 110 mL (9.0909 mg / mL) | Intermittent | 2752096I
178 | calcium gluconate | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752096I
179 | calcium gluconate | 2,000 mg / 70 mL (28.5714 mg / mL) | Intermittent | 2752096I
180 | calcium gluconate | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752096I
181 | calcium gluconate | 2,000 mg / 120 mL (16.6667 mg / mL) | Intermittent | 2752096I
182 | calcium gluconate | --- mg / --- mL | Intermittent | 2752096I
183 | calcium gluconate | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752096I
184 | calcium gluconate | 1 gram / 60 mL (0.0167 gram / mL) | Intermittent | 2752096I
185 | calcium gluconate | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752096I
186 | calcium gluconate | 1 gram / 110 mL (0.0091 gram / mL) | Intermittent | 2752096I
187 | calcium gluconate | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752096I
188 | calcium gluconate | 2 gram / 70 mL (0.0286 gram / mL) | Intermittent | 2752096I
189 | calcium gluconate | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752096I
190 | calcium gluconate | 2 gram / 120 mL (0.0167 gram / mL) | Intermittent | 2752096I
191 | calcium gluconate | --- gram / --- mL | Intermittent | 2752096I
192 | calcium gluconate | --- mEq / --- mL | Intermittent | 2752096I
193 | cangrelor | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 273246945C
194 | carbenicillin | --- mg / --- mL |  | 2752348I
195 | CARBOplatin | --- mg / --- mL | Intermittent | 2752528I
196 | carfilzomib | --- mg / --- mL | Intermittent | 24495877I
197 | carmustine | --- mg / --- mL |  | 2752551I
198 | casirivimab/imdevima | 1,200 mg / 260 mL (4.6154 mg / mL) | Intermittent | 3958645811I
199 | casirivimab/imdevima | 2,400 mg / 250 mL (9.6 mg / mL) | Intermittent | 3958645811I
200 | casirivimab/imdevima | --- mg / --- mL | Intermittent | 3958645811I
201 | caspofungin | 35 mg / 100 mL (0.35 mg / mL) | Intermittent | 2752634I
202 | caspofungin | 50 mg / 100 mL (0.5 mg / mL) | Intermittent | 2752634I
203 | caspofungin | 50 mg / 250 mL (0.2 mg / mL) | Intermittent | 2752634I
204 | caspofungin | 70 mg / 250 mL (0.28 mg / mL) | Intermittent | 2752634I
205 | caspofungin | --- mg / --- mL | Intermittent | 2752634I
206 | cefamandole | --- mg / --- mL |  | ???????????
207 | ceFAZolin | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2752689I
208 | ceFAZolin | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752689I
209 | ceFAZolin | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752689I
210 | ceFAZolin | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752689I
211 | ceFAZolin | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752689I
212 | ceFAZolin | 3,000 mg / 100 mL (30 mg / mL) | Intermittent | 2752689I
213 | ceFAZolin | --- mg / --- mL | Intermittent | 2752689I
214 | ceFAZolin | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752689I
215 | ceFAZolin | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752689I
216 | ceFAZolin | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752689I
217 | ceFAZolin | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752689I
218 | ceFAZolin | 3 gram / 100 mL (0.03 gram / mL) | Intermittent | 2752689I
219 | ceFAZolin | --- gram / --- mL | Intermittent | 2752689I
220 | cefepime | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2752723I
221 | cefepime | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2752723I
222 | cefepime | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752723I
223 | cefepime | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752723I
224 | cefepime | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752723I
225 | cefepime | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752723I
226 | cefepime | --- mg / --- mL | Intermittent | 2752723I
227 | cefepime | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752723I
228 | cefepime | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752723I
229 | cefepime | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752723I
230 | cefepime | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752723I
231 | cefepime | --- gram / --- mL | Intermittent | 2752723I
232 | cefiderocol | 750 mg / 100 mL (7.5 mg / mL) | Intermittent | 3602322021I
233 | cefiderocol | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 3602322021I
234 | cefiderocol | 1,500 mg / 100 mL (15 mg / mL) | Intermittent | 3602322021I
235 | cefiderocol | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 3602322021I
236 | cefiderocol | 0.75 gram / 100 mL (0.0075 gram / mL) | Intermittent | 3602322021I
237 | cefiderocol | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 3602322021I
238 | cefiderocol | 1.5 gram / 100 mL (0.015 gram / mL) | Intermittent | 3602322021I
239 | cefiderocol | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 3602322021I
240 | cefoperazone | --- mg / --- mL |  | ????????????
241 | cefotaxime | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752732I
242 | cefotaxime | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752732I
243 | cefotaxime | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752732I
244 | cefotaxime | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752732I
245 | cefotaxime | --- mg / --- mL | Intermittent | 2752732I
246 | cefotaxime | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752732I
247 | cefotaxime | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752732I
248 | cefotaxime | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752732I
249 | cefotaxime | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752732I
250 | cefotaxime | --- gram / --- mL | Intermittent | 2752732I
251 | cefoTEtan | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752747I
252 | cefoTEtan | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752747I
253 | cefoTEtan | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752747I
254 | cefoTEtan | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752747I
255 | cefoTEtan | --- mg / --- mL | Intermittent | 2752747I
256 | cefoTEtan | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752747I
257 | cefoTEtan | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752747I
258 | cefoTEtan | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752747I
259 | cefoTEtan | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752747I
260 | cefoTEtan | --- gram / --- mL | Intermittent | 2752747I
261 | cefOXitin | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752760I
262 | cefOXitin | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752760I
263 | cefOXitin | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752760I
264 | cefOXitin | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752760I
265 | cefOXitin | --- mg / --- mL | Intermittent | 2752760I
266 | cefOXitin | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752760I
267 | cefOXitin | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752760I
268 | cefOXitin | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752760I
269 | cefOXitin | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752760I
270 | cefOXitin | --- gram / --- mL | Intermittent | 2752760I
271 | ceftaroline | 200 mg / 50 mL (4 mg / mL) | Intermittent | 19950738I
272 | ceftaroline | 200 mg / 100 mL (2 mg / mL) | Intermittent | 19950738I
273 | ceftaroline | 200 mg / 250 mL (0.8 mg / mL) | Intermittent | 19950738I
274 | ceftaroline | 300 mg / 50 mL (6 mg / mL) | Intermittent | 19950738I
275 | ceftaroline | 300 mg / 100 mL (3 mg / mL) | Intermittent | 19950738I
276 | ceftaroline | 300 mg / 250 mL (1.2 mg / mL) | Intermittent | 19950738I
277 | ceftaroline | 400 mg / 50 mL (8 mg / mL) | Intermittent | 19950738I
278 | ceftaroline | 400 mg / 100 mL (4 mg / mL) | Intermittent | 19950738I
279 | ceftaroline | 400 mg / 250 mL (1.6 mg / mL) | Intermittent | 19950738I
280 | ceftaroline | 600 mg / 50 mL (12 mg / mL) | Intermittent | 19950738I
281 | ceftaroline | 600 mg / 100 mL (6 mg / mL) | Intermittent | 19950738I
282 | ceftaroline | 600 mg / 250 mL (2.4 mg / mL) | Intermittent | 19950738I
283 | ceftaroline | --- mg / --- mL | Intermittent | 19950738I
284 | cefTAZidime | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2752795I
285 | cefTAZidime | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2752795I
286 | cefTAZidime | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752795I
287 | cefTAZidime | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752795I
288 | cefTAZidime | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752795I
289 | cefTAZidime | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752795I
290 | cefTAZidime | --- mg / --- mL | Intermittent | 2752795I
291 | cefTAZidime | 0.5 gram / 50 mL (0.01 gram / mL) | Intermittent | 2752795I
292 | cefTAZidime | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2752795I
293 | cefTAZidime | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752795I
294 | cefTAZidime | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752795I
295 | cefTAZidime | 1.5 gram / 50 mL (0.03 gram / mL) | Intermittent | 2752795I
296 | cefTAZidime | 1.5 gram / 100 mL (0.015 gram / mL) | Intermittent | 2752795I
297 | cefTAZidime | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752795I
298 | cefTAZidime | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752795I
299 | cefTAZidime | --- gram / --- mL | Intermittent | 2752795I
300 | ceftizoxime | 1,000 mg / 50 mL (20 mg / mL) |  | 2752834I
301 | ceftizoxime | 2,000 mg / 50 mL (40 mg / mL) |  | 2752834I
302 | ceftizoxime | --- mg / --- mL |  | 2752834I
303 | ceftizoxime | 1 gram / 50 mL (0.02 gram / mL) |  | 2752834I
304 | ceftizoxime | 2 gram / 50 mL (0.04 gram / mL) |  | 2752834I
305 | ceftizoxime | --- gram / --- mL |  | 2752834I
306 | ceftolozane/tazo | 150 mg / 100 mL (1.5 mg / mL) | Intermittent | 202016343I
307 | ceftolozane/tazo | 375 mg / 100 mL (3.75 mg / mL) | Intermittent | 202016343I
308 | ceftolozane/tazo | 750 mg / 100 mL (7.5 mg / mL) | Intermittent | 202016343I
309 | ceftolozane/tazo | 1,500 mg / 100 mL (15 mg / mL) | Intermittent | 202016343I
310 | ceftolozane/tazo | 3,000 mg / 100 mL (30 mg / mL) | Intermittent | 202016343I
311 | ceftolozane/tazo | --- mg / --- mL | Intermittent | 202016343I
312 | ceftolozane/tazo | 1.5 gram / 100 mL (0.015 gram / mL) | Intermittent | 202016343I
313 | ceftolozane/tazo | 3 gram / 100 mL (0.03 gram / mL) | Intermittent | 202016343I
314 | ceftolozane/tazo | --- gram / --- mL | Intermittent | 202016343I
315 | cefTRIAXone | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752847I
316 | cefTRIAXone | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752847I
317 | cefTRIAXone | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752847I
318 | cefTRIAXone | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752847I
319 | cefTRIAXone | --- mg / --- mL | Intermittent | 2752847I
320 | cefTRIAXone | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752847I
321 | cefTRIAXone | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752847I
322 | cefTRIAXone | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752847I
323 | cefTRIAXone | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752847I
324 | cefTRIAXone | --- gram / --- mL | Intermittent | 2752847I
325 | cefuroxime | 750 mg / 50 mL (15 mg / mL) | Intermittent | 2752869I
326 | cefuroxime | 750 mg / 100 mL (7.5 mg / mL) | Intermittent | 2752869I
327 | cefuroxime | 1,500 mg / 50 mL (30 mg / mL) | Intermittent | 2752869I
328 | cefuroxime | 1,500 mg / 100 mL (15 mg / mL) | Intermittent | 2752869I
329 | cefuroxime | --- mg / --- mL | Intermittent | 2752869I
330 | cefuroxime | 0.75 gram / 50 mL (0.015 gram / mL) | Intermittent | 2752869I
331 | cefuroxime | 0.75 gram / 100 mL (0.0075 gram / mL) | Intermittent | 2752869I
332 | cefuroxime | 1.5 gram / 50 mL (0.03 gram / mL) | Intermittent | 2752869I
333 | cefuroxime | 1.5 gram / 100 mL (0.015 gram / mL) | Intermittent | 2752869I
334 | cefuroxime | --- gram / --- mL | Intermittent | 2752869I
335 | cemiplimab | --- mg / --- mL | Intermittent | 2970621255I
336 | cetuximab | --- mg / --- mL | Intermittent | 3316890I
337 | chloramphenicol | --- mg / --- mL | Intermittent | 2753038I
338 | chlorothiazide | --- mg / --- mL | Intermittent | 2753128I
339 | chlorproMAZINE | --- mg / --- mL | Intermittent | 2753638I
340 | cidofovir | --- mg / --- mL |  | 2753804I
341 | cimetidine | 300 mg / 50 mL (6 mg / mL) |  | 2753818I
342 | cimetidine | --- mg / --- mL |  | 2753818I
343 | ciprofloxacin | 200 mg / 100 mL (2 mg / mL) | Intermittent | 2753846I
344 | ciprofloxacin | 400 mg / 200 mL (2 mg / mL) | Intermittent | 2753846I
345 | ciprofloxacin | --- mg / --- mL | Intermittent | 2753846I
346 | cisatracurium | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2753896C
347 | cisatracurium | 100 mg / 250 mL (0.4 mg / mL) | Cont/Bolus or PCA | 2753896C
348 | cisatracurium | 200 mg / 200 mL (1 mg / mL) | Cont/Bolus or PCA | 2753896C
349 | cisatracurium | 200 mg / 250 mL (0.8 mg / mL) | Cont/Bolus or PCA | 2753896C
350 | cisatracurium | --- mg / --- mL | Cont/Bolus or PCA | 2753896C
351 | CISplatin | --- mg / --- mL | Intermittent | 2753903I
352 | cladribine | --- mg / --- mL |  | 2753956I
353 | clindamycin | 300 mg / 50 mL (6 mg / mL) | Intermittent | 2753995I
354 | clindamycin | 300 mg / 100 mL (3 mg / mL) | Intermittent | 2753995I
355 | clindamycin | 600 mg / 50 mL (12 mg / mL) | Intermittent | 2753995I
356 | clindamycin | 600 mg / 100 mL (6 mg / mL) | Intermittent | 2753995I
357 | clindamycin | 900 mg / 50 mL (18 mg / mL) | Intermittent | 2753995I
358 | clindamycin | 900 mg / 100 mL (9 mg / mL) | Intermittent | 2753995I
359 | clindamycin | --- mg / --- mL | Intermittent | 2753995I
360 | clodronate | --- mg / --- mL |  | 4333632I
361 | clofarabine | --- mg / --- mL |  | 3617188I
362 | cloxacillin | --- mg / --- mL |  | 2754214I
363 | CMV immunoglobulin | --- mg / --- mL |  | 2754840I
364 | colistimethate | --- mg / --- mL | Intermittent | 2754536I
365 | conivaptan | 20 mg / 100 mL (0.2 mg / mL) | Cont/Bolus or PCA | 4145032C
366 | conivaptan | 20 mg / 250 mL (0.08 mg / mL) | Cont/Bolus or PCA | 4145032C
367 | conivaptan | --- mg / --- mL | Cont/Bolus or PCA | 4145032C
368 | conivaptan Loading | 20 mg / 100 mL (0.2 mg / mL) | Intermittent | 4145032I
369 | cotrimoxazole | --- mg / --- mL |  | ????????
370 | crizanlizumab | --- mg / --- mL | Intermittent | 3488204201I
371 | cyclophosphamide | --- mg / --- mL | Intermittent | 2754749I
372 | cycloSPORINE | --- mg / --- mL |  | 2754780I
373 | cycloSPORINE cont | --- mg / --- mL |  | 2754780C
374 | cytarabine | --- mg / --- mL |  | 2754823I
375 | dacarbazine | --- mg / --- mL | Intermittent | 2754845I
376 | daclizumab | --- mg / --- mL |  | 2754853I
377 | DACTINomycin | --- mcg / --- mL |  | 2754858I
378 | dalbavancin | 375 mg / 100 mL (3.75 mg / mL) | Intermittent | 45868475I
379 | dalbavancin | 375 mg / 250 mL (1.5 mg / mL) | Intermittent | 45868475I
380 | dalbavancin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 45868475I
381 | dalbavancin | 500 mg / 250 mL (2 mg / mL) | Intermittent | 45868475I
382 | dalbavancin | 750 mg / 250 mL (3 mg / mL) | Intermittent | 45868475I
383 | dalbavancin | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 45868475I
384 | dalbavancin | 1,125 mg / 250 mL (4.5 mg / mL) | Intermittent | 45868475I
385 | dalbavancin | 1,500 mg / 500 mL (3 mg / mL) | Intermittent | 45868475I
386 | dalbavancin | --- mg / --- mL | Intermittent | 45868475I
387 | dalbavancin | 0.375 gram / 100 mL (0.0038 gram / mL) | Intermittent | 45868475I
388 | dalbavancin | 0.375 gram / 250 mL (0.0015 gram / mL) | Intermittent | 45868475I
389 | dalbavancin | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 45868475I
390 | dalbavancin | 0.5 gram / 250 mL (0.002 gram / mL) | Intermittent | 45868475I
391 | dalbavancin | 0.75 gram / 250 mL (0.003 gram / mL) | Intermittent | 45868475I
392 | dalbavancin | 1 gram / 250 mL (0.004 gram / mL) | Intermittent | 45868475I
393 | dalbavancin | 1.125 gram / 250 mL (0.0045 gram / mL) | Intermittent | 45868475I
394 | dalbavancin | 1.5 gram / 500 mL (0.003 gram / mL) | Intermittent | 45868475I
395 | dalbavancin | --- gram / --- mL | Intermittent | 45868475I
396 | danaparoid | --- unit / --- mL |  | 4333829I
397 | dantrolene | --- mg / --- mL | Intermittent | 2754894I
398 | DAPTOmycin | --- mg / --- mL | Intermittent | 2754915I
399 | daratumumab | --- mg / --- mL | Intermittent | 275110833I
400 | DAUNOrubicin | --- mg / --- mL |  | 2754939I
401 | DAUNOrubicin liposo | --- mg / --- mL |  | 2754945I
402 | decitabine | --- mg / --- mL | Intermittent | 4145048I
403 | deferoxamine | --- mg / --- mL | Intermittent | 2754950I
404 | denileukin | --- mcg / --- mL |  | 2754986I
405 | desferoxamine | --- mg / --- mL |  | 2754950
406 | desmopressin | --- mcg / --- mL | Intermittent | 2755032I
407 | dexamethasone | --- mg / --- mL | Intermittent | 2755093I
408 | dexmedetomidine | 200 mcg / 50 mL (4 mcg / mL) | Cont/Bolus or PCA | 2755261C
409 | dexmedetomidine | 200 mcg / 100 mL (2 mcg / mL) | Cont/Bolus or PCA | 2755261C
410 | dexmedetomidine | 400 mcg / 100 mL (4 mcg / mL) | Cont/Bolus or PCA | 2755261C
411 | dexmedetomidine | --- mcg / --- mL | Cont/Bolus or PCA | 2755261C
412 | dexrazoxane | --- mg / --- mL |  | 2755283I
413 | dextran 40 | 50 gram / 500 mL (0.1 gram / mL) |  | 2755295I
414 | dextran 40 | --- gram / --- mL |  | 2755295I
415 | diazePAM | 50 mg / 50 mL (1 mg / mL) |  | 2755701I
416 | diazePAM | --- mg / --- mL |  | 2755701I
417 | digoxin | --- mcg / --- mL |  | 2755862I
418 | digoxin immune fab | --- mg / --- mL | Intermittent | 2755888I
419 | dihydroergotamine | --- mg / --- mL |  | 2755900I
420 | dilTIAZem | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2755935C
421 | dilTIAZem | 125 mg / 125 mL (1 mg / mL) | Cont/Bolus or PCA | 2755935C
422 | dilTIAZem | --- mg / --- mL | Cont/Bolus or PCA | 2755935C
423 | dimenhyDRINATE | --- mg / --- mL |  | 2756003I
424 | diphenhydrAMINE | --- mg / --- mL |  | 2756043I
425 | dipyridamole | --- mg / --- mL |  | 2756254I
426 | DOBUTamine | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2756310C
427 | DOBUTamine | 500 mg / 250 mL (2 mg / mL) | Cont/Bolus or PCA | 2756310C
428 | DOBUTamine | 500 mg / 500 mL (1 mg / mL) | Cont/Bolus or PCA | 2756310C
429 | DOBUTamine | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA | 2756310C
430 | DOBUTamine | 1,000 mg / 500 mL (2 mg / mL) | Cont/Bolus or PCA | 2756310C
431 | DOBUTamine | --- mg / --- mL | Cont/Bolus or PCA | 2756310C,2756310I
432 | Dobutamine Stress | --- mg / --- mL |  | 2756310
433 | DOCEtaxel | --- mg / --- mL | Intermittent | 2756319I
434 | dolasetron | --- mg / --- mL |  | 2756447I
435 | DOPamine | 400 mg / 250 mL (1.6 mg / mL) | Cont/Bolus or PCA | 2756465C
436 | DOPamine | 800 mg / 250 mL (3.2 mg / mL) | Cont/Bolus or PCA | 2756465C
437 | DOPamine | --- mg / --- mL | Cont/Bolus or PCA | 2756465C
438 | doripenem | 250 mg / 50 mL (5 mg / mL) | Intermittent | 4316756I
439 | doripenem | 250 mg / 100 mL (2.5 mg / mL) | Intermittent | 4316756I
440 | doripenem | 500 mg / 100 mL (5 mg / mL) | Intermittent | 4316756I
441 | doxaCURIUM | --- mg / --- mL |  | 2756488I
442 | doxapram | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2756493C
443 | doxapram | --- mg / --- mL | Cont/Bolus or PCA | 2756493C
444 | DOXOrubicin | --- mg / --- mL | Intermittent | 2756540I
445 | DOXOrubicin, liposo | --- mg / --- mL | Intermittent | 2756552I
446 | doxycycline | 100 mg / 100 mL (1 mg / mL) | Intermittent | 2756557I
447 | doxycycline | 200 mg / 250 mL (0.8 mg / mL) | Intermittent | 2756557I
448 | doxycycline | --- mg / --- mL | Intermittent | 2756557I
449 | droperidol | --- mg / --- mL |  | 2756618I
450 | drotrecogin | --- mg / --- mL |  | 2756633I
451 | durvalumab | 1,500 mg / 280 mL (5.3571 mg / mL) | Intermittent | 2581288283I
452 | durvalumab | --- mg / --- mL | Intermittent | 2581288283I
453 | eculizumab | 300 mg / 60 mL (5 mg / mL) |  | 4245702I
454 | eculizumab | 600 mg / 120 mL (5 mg / mL) |  | 4245702I
455 | eculizumab | 900 mg / 180 mL (5 mg / mL) |  | 4245702I
456 | eculizumab | 1,200 mg / 240 mL (5 mg / mL) |  | 4245702I
457 | eculizumab | --- mg / --- mL |  | 4245702I
458 | EDTA | --- mg / --- mL |  | edta
459 | enalaprilat | --- mcg / --- mL |  | 2757060I
460 | enfortumab vedotin | --- mg / --- mL | Intermittent | 3488204431I
461 | epidural w/ fentanyl | 100 unit / 100 mL (1 unit / mL) | Cont/Bolus or PCA | epiduralfent
462 | epidural w/o fentany | 100 unit / 100 mL (1 unit / mL) | Cont/Bolus or PCA | epidnofent
463 | EPINEPHrine | 4 mg / 250 mL (0.016 mg / mL) | Cont/Bolus or PCA | 2757140C
464 | EPINEPHrine | 5 mg / 250 mL (0.02 mg / mL) | Cont/Bolus or PCA | 2757140C
465 | EPINEPHrine | 8 mg / 250 mL (0.032 mg / mL) | Cont/Bolus or PCA | 2757140C
466 | EPINEPHrine | 10 mg / 500 mL (0.02 mg / mL) | Cont/Bolus or PCA | 2757140C
467 | EPINEPHrine | 16 mg / 250 mL (0.064 mg / mL) | Cont/Bolus or PCA | 2757140C
468 | EPINEPHrine | 30 mg / 250 mL (0.12 mg / mL) | Cont/Bolus or PCA | 2757140C
469 | EPINEPHrine | --- mg / --- mL | Cont/Bolus or PCA | 2757140C
470 | epiRUBicin | --- mg / --- mL |  | 2757216I
471 | epoprostenol | 500 mcg / 100 mL (5 mcg / mL) | Cont/Bolus or PCA | 2757252C
472 | epoprostenol | 1,000 mcg / 100 mL (10 mcg / mL) | Cont/Bolus or PCA | 2757252C
473 | epoprostenol | 3,000 mcg / 100 mL (30 mcg / mL) | Cont/Bolus or PCA | 2757252C
474 | epoprostenol | --- mcg / --- mL | Cont/Bolus or PCA | 2757252C
475 | epoprostenol | 1.5 mg / 50 mL (0.03 mg / mL) | Cont/Bolus or PCA | 2757252C
476 | eptifibatide | 75 mg / 100 mL (0.75 mg / mL) | Cont/Bolus or PCA | 2757273C
477 | eptifibatide | 200 mg / 100 mL (2 mg / mL) | Cont/Bolus or PCA | 2757273C
478 | eriBULin | --- mg / --- mL | Intermittent | 19952644I
479 | ertapenem | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2757309I
480 | ertapenem | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2757309I
481 | ertapenem | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2757309I
482 | ertapenem | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2757309I
483 | ertapenem | 0.5 gram / 50 mL (0.01 gram / mL) | Intermittent | 2757309I
484 | ertapenem | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2757309I
485 | ertapenem | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2757309I
486 | ertapenem | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2757309I
487 | erythromycin | 250 mg / 100 mL (2.5 mg / mL) | Intermittent | 2757314I
488 | erythromycin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2757314I
489 | erythromycin | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 2757314I
490 | erythromycin | --- mg / --- mL | Intermittent | 2757314I
491 | erythromycin | --- gram / --- mL | Intermittent | 2757314I
492 | esmolol | 2,500 mg / 250 mL (10 mg / mL) | Cont/Bolus or PCA | 2757437C
493 | esmolol | 5,000 mg / 250 mL (20 mg / mL) | Cont/Bolus or PCA | 2757437C
494 | esmolol | --- mg / --- mL | Cont/Bolus or PCA | 2757437C
495 | esomeprazole | --- mg / --- mL |  | 2757446C,2757446I
496 | ethanol 10% | --- gram / --- mL |  | 2757676I
497 | ethanol 5% | --- gram / --- mL |  | 2757676
498 | etidronate | --- mg / --- mL |  | 2757876I
499 | etomidate | --- mg / --- mL |  | 2757905I
500 | etoposide | --- mg / --- mL | Intermittent | 2757912I
501 | factor IX | --- unit / --- mL |  | 2757938I
502 | factor VIII | --- unit / --- mL |  | 24606459I
503 | factor VIIIa | --- unit / --- mL |  | 24606459
504 | famotidine | 20 mg / 50 mL (0.4 mg / mL) | Intermittent | 2757966I
505 | famotidine | --- mg / --- mL | Intermittent | 2757966I
506 | fat emulsion 10% | 10 gram / 100 mL (0.1 gram / mL) |  | 2757989
507 | fat emulsion 20% | 20 gram / 100 mL (0.2 gram / mL) | Intermittent | 2757989I
508 | fat emulsion 20% | 25 gram / 125 mL (0.2 gram / mL) | Intermittent | 2757989I
509 | fat emulsion 20% | 50 gram / 250 mL (0.2 gram / mL) | Intermittent | 2757989I
510 | fat emulsion 30% | 30 gram / 100 mL (0.3 gram / mL) |  | 2757989-30%
511 | fenoldopam | 10 mg / 250 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2758046C
512 | fentaNYL | 1,000 mcg / 100 mL (10 mcg / mL) | Cont/Bolus or PCA | 2758060C
513 | fentaNYL | 1,250 mcg / 25 mL (50 mcg / mL) | Cont/Bolus or PCA | 2758060C
514 | fentaNYL | 2,500 mcg / 100 mL (25 mcg / mL) | Cont/Bolus or PCA | 2758060C
515 | fentaNYL | 2,500 mcg / 250 mL (10 mcg / mL) | Cont/Bolus or PCA | 2758060C
516 | fentaNYL | --- mcg / --- mL | Cont/Bolus or PCA | 2758060C
517 | ferric carboxymaltos | 750 mg / 265 mL (2.8302 mg / mL) | Intermittent | 24606832I
518 | ferric carboxymaltos | --- mg / --- mL | Intermittent | 24606832I
519 | ferric gluconate Na | 125 mg / 100 mL (1.25 mg / mL) |  | 2758129I
520 | ferumoxytol | 510 mg / 117 mL (4.359 mg / mL) | Intermittent | 19952174I
521 | ferumoxytol | 1,020 mg / 284 mL (3.5915 mg / mL) | Intermittent | 19952174I
522 | ferumoxytol | 0.51 gram / 117 mL (0.0044 gram / mL) | Intermittent | 19952174I
523 | ferumoxytol | 1.02 gram / 284 mL (0.0036 gram / mL) | Intermittent | 19952174I
524 | filgrastim | --- mcg / --- mL |  | 2758220I
525 | floxuridine | --- mg / --- mL |  | 2758251I
526 | fluconazole | 100 mg / 50 mL (2 mg / mL) | Intermittent | 2758256I
527 | fluconazole | 200 mg / 100 mL (2 mg / mL) | Intermittent | 2758256I
528 | fluconazole | 400 mg / 200 mL (2 mg / mL) | Intermittent | 2758256I
529 | fluconazole | 800 mg / 400 mL (2 mg / mL) | Intermittent | 2758256I
530 | fluconazole | --- mg / --- mL | Intermittent | 2758256I
531 | fludarabine | --- mg / --- mL |  | 2758282I
532 | fluorouracil | --- mg / --- mL | Intermittent | 2758517C,2758517I
533 | folic acid | --- mg / --- mL | Intermittent | 2758687I
534 | fomepizole | --- mg / --- mL | Intermittent | 2758716I
535 | fosaprepitant | 150 mg / 250 mL (0.6 mg / mL) | Intermittent | 17213916I
536 | foscarnet | 100 mg / 50 mL (2 mg / mL) |  | 2758754I
537 | foscarnet | --- mg / --- mL |  | 2758754I
538 | fosphenytoin | --- mg / --- mL | Intermittent | 2758779I
539 | furosemide | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2758798C
540 | furosemide | 200 mg / 100 mL (2 mg / mL) | Cont/Bolus or PCA | 2758798C
541 | furosemide | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2758798C
542 | furosemide | 300 mg / 30 mL (10 mg / mL) | Cont/Bolus or PCA | 2758798C
543 | furosemide | 500 mg / 100 mL (5 mg / mL) | Cont/Bolus or PCA | 2758798C
544 | furosemide | --- mg / --- mL | Cont/Bolus or PCA | 2758798C
545 | ganciclovir | --- mg / --- mL | Intermittent | 2758863I
546 | gancyclovir | --- mg / --- mL |  | 2758863
547 | ganisetron | --- mcg / --- mL |  | 2759194
548 | gatifloxacin | 200 mg / 100 mL (2 mg / mL) |  | 2758883I
549 | gatifloxacin | 400 mg / 200 mL (2 mg / mL) |  | 2758883I
550 | gatifloxacin | --- mg / --- mL |  | 2758883I
551 | gemcitabine | --- mg / --- mL | Intermittent | 2758906I
552 | gemtuzumab | --- mg / --- mL |  | 2758923I
553 | gentamicin | 80 mg / 100 mL (0.8 mg / mL) | Intermittent | 2758928I
554 | gentamicin | 120 mg / 100 mL (1.2 mg / mL) | Intermittent | 2758928I
555 | gentamicin | --- mg / --- mL | Intermittent | 2758928I
556 | glucagon | 10 mg / 50 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2759022C
557 | glucagon | 10 mg / 100 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2759022C
558 | glucagon | --- mg / --- mL | Cont/Bolus or PCA | 2759022C
559 | golimumab | --- mg / --- mL | Intermittent | 19952385I
560 | granisetron | --- mg / --- mL |  | 2759194I
561 | haloperidol | --- mg / --- mL |  | 2759713I
562 | heparin | 100 unit / 1 mL (100 unit / mL) | Cont/Bolus or PCA | 2759747C
563 | heparin | 1,250 unit / 50 mL (25 unit / mL) | Cont/Bolus or PCA | 2759747C
564 | heparin | 2,500 unit / 25 mL (100 unit / mL) | Cont/Bolus or PCA | 2759747C
565 | heparin | 2,500 unit / 50 mL (50 unit / mL) | Cont/Bolus or PCA | 2759747C
566 | heparin | 5,000 unit / 50 mL (100 unit / mL) | Cont/Bolus or PCA | 2759747C
567 | heparin | 5,000 unit / 100 mL (50 unit / mL) | Cont/Bolus or PCA | 2759747C
568 | heparin | 25,000 unit / 250 mL (100 unit / mL) | Cont/Bolus or PCA | 2759747C
569 | heparin | 25,000 unit / 500 mL (50 unit / mL) | Cont/Bolus or PCA | 2759747C
570 | heparin | --- unit / --- mL | Cont/Bolus or PCA | 2759747C
571 | hetastarch 6% | 30 gram / 500 mL (0.06 gram / mL) | Intermittent | 2759815C,2759815I
572 | hydrALAZINE | --- mg / --- mL |  | 2759877I
573 | hydrochloric acid | 0.1 mEq / 1 mL (0.1 mEq / mL) |  | HCL
574 | hydrochloric acid | 0.2 mEq / 1 mL (0.2 mEq / mL) |  | HCL
575 | hydrochloric acid | --- mEq / --- mL |  | HCL
576 | hydrocortisone | --- mg / --- mL | Intermittent | 2760102I
577 | HYDROmorphone | 20 mg / 100 mL (0.2 mg / mL) |  | 2760421C
578 | HYDROmorphone | 100 mg / 100 mL (1 mg / mL) |  | 2760421C
579 | HYDROmorphone | --- mg / --- mL |  | 2760421C
580 | ibuprofen lysine | --- mg / --- mL |  | 2760660I
581 | ibutilide | 0.5 mg / 50 mL (0.01 mg / mL) | Intermittent | 2760741I
582 | ibutilide | 0.5 mg / 100 mL (0.005 mg / mL) | Intermittent | 2760741I
583 | ibutilide | 1 mg / 50 mL (0.02 mg / mL) | Intermittent | 2760741I
584 | ibutilide | 1 mg / 60 mL (0.0167 mg / mL) | Intermittent | 2760741I
585 | ibutilide | 1 mg / 100 mL (0.01 mg / mL) | Intermittent | 2760741I
586 | ibutilide | --- mg / --- mL | Intermittent | 2760741I
587 | IDArubicin | --- mg / --- mL |  | 2760752I
588 | Ifex / Mesna | --- mg / --- mL |  | 2760768I
589 | Ifex / Mesna | --- gram / --- mL |  | 2760768I
590 | ifosfamide | --- mg / --- mL | Intermittent | 2760761I
591 | ifosfamide | --- gram / --- mL | Intermittent | 2760761I
592 | iloprost | --- mcg / --- mL |  | 3617215I
593 | imiglucerase | --- unit / --- mL |  | 2760783I
594 | imipenem/cilastatin | 250 mg / 100 mL (2.5 mg / mL) | Intermittent | 2760792I
595 | imipenem/cilastatin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2760792I
596 | imipenem/cilastatin | 750 mg / 250 mL (3 mg / mL) | Intermittent | 2760792I
597 | imipenem/cilastatin | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 2760792I
598 | imipenem/cilastatin | --- mg / --- mL | Intermittent | 2760792I
599 | imipenem/cilastatin | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2760792I
600 | imipenem/cilastatin | 1 gram / 250 mL (0.004 gram / mL) | Intermittent | 2760792I
601 | imipenem/cilastatin | --- gram / --- mL | Intermittent | 2760792I
602 | immune glob (IVIG) | 5 gram / 50 mL (0.1 gram / mL) | Intermittent | 2760832I
603 | immune glob (IVIG) | 5 gram / 100 mL (0.05 gram / mL) | Intermittent | 2760832I
604 | immune glob (IVIG) | 10 gram / 100 mL (0.1 gram / mL) | Intermittent | 2760832I
605 | immune glob (IVIG) | 10 gram / 200 mL (0.05 gram / mL) | Intermittent | 2760832I
606 | immune glob (IVIG) | 20 gram / 200 mL (0.1 gram / mL) | Intermittent | 2760832I
607 | immune glob (IVIG) | --- gram / --- mL | Intermittent | 2760832I
608 | immune globulin Rho | --- gram / --- mL |  | 2768135I
609 | inamrinone | 100 mg / 100 mL (1 mg / mL) |  | 2760881I
610 | inamrinone | 300 mg / 60 mL (5 mg / mL) |  | 2760881I
611 | indomethacin | --- mg / --- mL |  | 2760908I
612 | inFLIXimab | --- mg / --- mL | Intermittent | 2760925I
613 | insulin | 10 unit / 1,000 mL (0.01 unit / mL) | Cont/Bolus or PCA | 2761029C
614 | insulin | 100 unit / 100 mL (1 unit / mL) | Cont/Bolus or PCA | 2761029C
615 | insulin | 250 unit / 250 mL (1 unit / mL) | Cont/Bolus or PCA | 2761029C
616 | insulin | --- unit / --- mL | Cont/Bolus or PCA | 2761029C
617 | interferon alpha-2b | --- unit / --- mL |  | 2761065I
618 | ipilimumab | --- mg / --- mL | Intermittent | 19952974I
619 | irinotecan | --- mg / --- mL | Intermittent | 2761340I
620 | irinotecan liposomal | --- mg / --- mL | Intermittent | 275110257I
621 | iron dextran | 25 mg / 50 mL (0.5 mg / mL) |  | 2761345I
622 | iron dextran | --- mg / --- mL |  | 2761345I
623 | iron sucrose | 100 mg / 55 mL (1.8182 mg / mL) | Intermittent | 2761376I
624 | iron sucrose | 100 mg / 105 mL (0.9524 mg / mL) | Intermittent | 2761376I
625 | iron sucrose | 200 mg / 110 mL (1.8182 mg / mL) | Intermittent | 2761376I
626 | iron sucrose | 300 mg / 265 mL (1.1321 mg / mL) | Intermittent | 2761376I
627 | iron sucrose | 400 mg / 270 mL (1.4815 mg / mL) | Intermittent | 2761376I
628 | iron sucrose | 500 mg / 275 mL (1.8182 mg / mL) | Intermittent | 2761376I
629 | iron sucrose | --- mg / --- mL | Intermittent | 2761376I
630 | isoproterenol | 1 mg / 250 mL (0.004 mg / mL) | Cont/Bolus or PCA | 2761419C
631 | isoproterenol | 2 mg / 250 mL (0.008 mg / mL) | Cont/Bolus or PCA | 2761419C
632 | isoproterenol | 2 mg / 500 mL (0.004 mg / mL) | Cont/Bolus or PCA | 2761419C
633 | isoproterenol | --- mg / --- mL | Cont/Bolus or PCA | 2761419C
634 | itraconazole | 200 mg / 60 mL (3.3333 mg / mL) | Intermittent | 2761517I
635 | itraconazole | --- mg / --- mL | Intermittent | 2761517I
636 | ixabepilone | --- mg / --- mL |  | 4316936I
637 | Kcentra (Factor IX) | --- unit / --- mL | Intermittent | 24608305I
638 | ketamine | --- mg / --- mL |  | 2761553I
639 | ketamine ANES | 50 mg / 50 mL (1 mg / mL) | Cont/Bolus or PCA | 2761553C
640 | ketamine ANES | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2761553C
641 | ketamine ANES | --- mg / --- mL | Cont/Bolus or PCA | 2761553C
642 | Kinevac (sincalide) | --- mcg / --- mL | Intermittent | 2768854I
643 | Kinevac (sincalide) | --- mg / --- mL | Intermittent | 2768854I
644 | labetalol | 200 mg / 200 mL (1 mg / mL) | Cont/Bolus or PCA | 2761620C
645 | labetalol | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2761620C
646 | labetalol | 500 mg / 250 mL (2 mg / mL) | Cont/Bolus or PCA | 2761620C
647 | labetalol | --- mg / --- mL | Cont/Bolus or PCA | 2761620C
648 | lacosamide | --- mg / --- mL | Intermittent | 19952919I
649 | lepirudin | 100 mg / 250 mL (0.4 mg / mL) |  | 2761819I
650 | lepirudin | --- mg / --- mL |  | 2761819I
651 | leucovorin | --- mg / --- mL | Intermittent | 2761829I
652 | levETIRAcetam | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2761878I
653 | levETIRAcetam | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2761878I
654 | levETIRAcetam | 1,500 mg / 100 mL (15 mg / mL) | Intermittent | 2761878I
655 | levETIRAcetam | --- mg / --- mL | Intermittent | 2761878I
656 | levETIRAcetam | --- gram / --- mL | Intermittent | 2761878I
657 | levobupivicaine | --- mg / --- mL |  | 2761905I
658 | levOCARNitine | --- mg / --- mL |  | 2761919I
659 | levoFLOXacin | 250 mg / 50 mL (5 mg / mL) | Intermittent | 2761938I
660 | levoFLOXacin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2761938I
661 | levoFLOXacin | 750 mg / 150 mL (5 mg / mL) | Intermittent | 2761938I
662 | levoFLOXacin | --- mg / --- mL | Intermittent | 2761938I
663 | LEVOleucovorin | --- mg / --- mL |  | 17213770I
664 | levothyroxine | --- mcg / --- mL |  | 2761978I
665 | lidocaine | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA | 2762076C
666 | lidocaine | 2,000 mg / 500 mL (4 mg / mL) | Cont/Bolus or PCA | 2762076C
667 | lidocaine | --- mg / --- mL | Cont/Bolus or PCA | 2762076C
668 | linezolid | 200 mg / 100 mL (2 mg / mL) | Intermittent | 2762196I
669 | linezolid | 600 mg / 300 mL (2 mg / mL) | Intermittent | 2762196I
670 | linezolid | --- mg / --- mL | Intermittent | 2762196I
671 | liothyronine | --- mcg / --- mL |  | 2762207I
672 | LORazepam | 20 mg / 100 mL (0.2 mg / mL) |  | 2762369C
673 | LORazepam | 40 mg / 100 mL (0.4 mg / mL) |  | 2762369C
674 | LORazepam | 120 mg / 60 mL (2 mg / mL) |  | 2762369C
675 | LORazepam | --- mg / --- mL |  | 2762369C
676 | mag OB maintenance | 20 gram / 500 mL (0.04 gram / mL) | Cont/Bolus or PCA | 2762580C
677 | mag sulfate | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2762580I
678 | mag sulfate | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2762580I
679 | mag sulfate | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2762580I
680 | mag sulfate | 4 gram / 100 mL (0.04 gram / mL) | Intermittent | 2762580I
681 | mag sulfate | --- gram / --- mL | Intermittent | 2762580I
682 | magnesium sulfate | 1 gram / 50 mL (0.02 gram / mL) |  | 2762580
683 | magnesium sulfate | 4 gram / 100 mL (0.04 gram / mL) |  | 2762580
684 | magnesium sulfate | 40 gram / 500 mL (0.08 gram / mL) |  | 2762580
685 | magnesium sulfate | --- gram / --- mL |  | 2762580
686 | mannitol 20% | 50 gram / 250 mL (0.2 gram / mL) | Intermittent | 2762620I
687 | mannitol 20% | 100 gram / 500 mL (0.2 gram / mL) | Intermittent | 2762620I
688 | mannitol 20% | --- gram / --- mL | Intermittent | 2762620I
689 | mannitol 10% | 100 gram / 1,000 mL (0.1 gram / mL) |  | 2762620
690 | mannitol 10% | --- gram / --- mL |  | 2762620
691 | mannitol 15% | 75 gram / 500 mL (0.15 gram / mL) |  | 2762620-15
692 | mannitol 15% | --- gram / --- mL |  | 2762620-15
693 | mannitol 20% | 50 gram / 250 mL (0.2 gram / mL) |  | 2762620-20
694 | mannitol 20% | --- gram / --- mL |  | 2762620-20
695 | mannitol 25% | 12.5 gram / 50 mL (0.25 gram / mL) |  | 2762620-12.5
696 | mannitol 25% | --- gram / --- mL |  | 2762620-12.5
697 | mannitol 5% | 50 gram / 1,000 mL (0.05 gram / mL) |  | 2762620-5
698 | mannitol 5% | --- gram / --- mL |  | 2762620-5
699 | mechlorethamine | --- mg / --- mL |  | 2762671I
700 | melphalan | --- mg / --- mL |  | 2762749I
701 | meperidine | 10 mg / 1 mL (10 mg / mL) |  | 2762823I
702 | meperidine | 500 mg / 50 mL (10 mg / mL) |  | 2762823C
703 | meperidine | --- mg / --- mL |  | 2762823C
704 | meropenem | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2762894I
705 | meropenem | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2762894I
706 | meropenem | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2762894I
707 | meropenem | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2762894I
708 | meropenem | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2762894I
709 | meropenem | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2762894I
710 | meropenem | --- mg / --- mL | Intermittent | 2762894I
711 | meropenem | 0.5 gram / 50 mL (0.01 gram / mL) | Intermittent | 2762894I
712 | meropenem | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2762894I
713 | meropenem | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2762894I
714 | meropenem | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2762894I
715 | meropenem | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2762894I
716 | meropenem | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2762894I
717 | meropenem | --- gram / --- mL | Intermittent | 2762894I
718 | meropenem/vaborbacta | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2670044615I
719 | meropenem/vaborbacta | 1 gram / 250 mL (0.004 gram / mL) | Intermittent | 2670044615I
720 | meropenem/vaborbacta | 2 gram / 250 mL (0.008 gram / mL) | Intermittent | 2670044615I
721 | meropenem/vaborbacta | 2 gram / 500 mL (0.004 gram / mL) | Intermittent | 2670044615I
722 | meropenem/vaborbacta | 4 gram / 250 mL (0.016 gram / mL) | Intermittent | 2670044615I
723 | meropenem/vaborbacta | 4 gram / 500 mL (0.008 gram / mL) | Intermittent | 2670044615I
724 | meropenem/vaborbacta | 4 gram / 1,000 mL (0.004 gram / mL) | Intermittent | 2670044615I
725 | meropenem/vaborbacta | --- gram / --- mL | Intermittent | 2670044615I
726 | mesna | --- mg / --- mL | Intermittent | 2762918I
727 | methadone | --- mg / --- mL |  | 2763004C
728 | methocarbamol | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 2763080I
729 | methocarbamol | --- mg / --- mL | Intermittent | 2763080I
730 | methohexitol | --- mg / --- mL |  | 2763092I
731 | methotrexate | --- mg / --- mL |  | 2763101I
732 | methotrexate | --- gram / --- mL |  | 2763101I
733 | methyldopa | --- mg / --- mL |  | 2763256I
734 | methylPREDNISolone | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2763323I
735 | methylPREDNISolone | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2763323I
736 | methylPREDNISolone | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2763323I
737 | methylPREDNISolone | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2763323I
738 | methylPREDNISolone | --- mg / --- mL | Intermittent | 2763323I
739 | metoclopramide | --- mg / --- mL |  | 2763400I
740 | metoprolol | --- mg / --- mL |  | 2763424C,2763424I
741 | metroNIDAZOLE | 250 mg / 50 mL (5 mg / mL) | Intermittent | 2763443I
742 | metroNIDAZOLE | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2763443I
743 | metroNIDAZOLE | --- mg / --- mL | Intermittent | 2763443I
744 | mezlocillin | --- mg / --- mL |  | unknown
745 | micafungin | 50 mg / 100 mL (0.5 mg / mL) | Intermittent | 3830619I
746 | micafungin | 100 mg / 100 mL (1 mg / mL) | Intermittent | 3830619I
747 | micafungin | 150 mg / 100 mL (1.5 mg / mL) | Intermittent | 3830619I
748 | micafungin | --- mg / --- mL | Intermittent | 3830619I
749 | miconazole | --- mg / --- mL |  | 19953615I
750 | midazolam | 50 mg / 50 mL (1 mg / mL) | Cont/Bolus or PCA | 2763568C
751 | midazolam | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2763568C
752 | midazolam | 125 mg / 125 mL (1 mg / mL) | Cont/Bolus or PCA | 2763568C
753 | midazolam | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2763568C
754 | midazolam | --- mg / --- mL | Cont/Bolus or PCA | 2763568C
755 | milrinone | 20 mg / 100 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2763605C
756 | milrinone | 40 mg / 200 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2763605C
757 | milrinone | --- mg / --- mL | Cont/Bolus or PCA | 2763605C
758 | minocycline | --- mg / --- mL |  | 2763628I
759 | mitoMYcin | --- mg / --- mL | Intermittent | 2763694I
760 | mitoXANTRONE | --- mg / --- mL |  | 2763708I
761 | morphine | 1 mg / 1 mL (1 mg / mL) |  | 2763791I
762 | morphine | 30 mg / 30 mL (1 mg / mL) |  | 2763791C,2763791I
763 | morphine | 50 mg / 50 mL (1 mg / mL) |  | 2763791C
764 | morphine | 100 mg / 100 mL (1 mg / mL) |  | 2763791C
765 | morphine | --- mg / --- mL |  | 2763791C
766 | moxifloxacin | 400 mg / 250 mL (1.6 mg / mL) |  | 2763901I
767 | moxifloxacin | --- mg / --- mL |  | 2763901I
768 | muromonab | --- mg / --- mL |  | 2764672I
769 | MVI + Thiame + Folic | 10 unit / 1,011 mL (0.0099 unit / mL) | Intermittent | 1000000002I
770 | MVI Only | 10 unit / 510 mL (0.0196 unit / mL) | Intermittent | 1000000003I
771 | MVI Only | 10 unit / 1,010 mL (0.0099 unit / mL) | Intermittent | 1000000003I
772 | mycophenolate | --- mg / --- mL |  | 3320344I
773 | nafcillin | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2764713I
774 | nafcillin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2764713I
775 | nafcillin | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2764713I
776 | nafcillin | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2764713I
777 | nafcillin | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2764713I
778 | nafcillin | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2764713I
779 | nafcillin | --- mg / --- mL | Intermittent | 2764713I
780 | nafcillin | 0.5 gram / 50 mL (0.01 gram / mL) | Intermittent | 2764713I
781 | nafcillin | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2764713I
782 | nafcillin | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2764713I
783 | nafcillin | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2764713I
784 | nafcillin | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2764713I
785 | nafcillin | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2764713I
786 | nafcillin | --- gram / --- mL | Intermittent | 2764713I
787 | nalbuphine | --- mg / --- mL |  | 2764738I
788 | naloxone | 2 mg / 500 mL (0.004 mg / mL) | Cont/Bolus or PCA | 2764760C
789 | naloxone | 4 mg / 100 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2764760C
790 | naloxone | --- mg / --- mL | Cont/Bolus or PCA | 2764760C
791 | nelarabine | --- mg / --- mL |  | 3976209I
792 | nesiritide | 1.5 mg / 250 mL (0.006 mg / mL) | Cont/Bolus or PCA | 2764982C
793 | niCARdipine | 20 mg / 200 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2765030C
794 | niCARdipine | 25 mg / 250 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2765030C
795 | niCARdipine | 40 mg / 200 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2765030C
796 | niCARdipine | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2765030C
797 | niCARdipine | 50 mg / 500 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2765030C
798 | niCARdipine | 500 mg / 250 mL (2 mg / mL) | Cont/Bolus or PCA | 2765030C
799 | niCARdipine | --- mg / --- mL | Cont/Bolus or PCA | 2765030C
800 | nitroglycerin | 25 mg / 250 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2765191C
801 | nitroglycerin | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2765191C
802 | nitroglycerin | --- mg / --- mL | Cont/Bolus or PCA | 2765191C
803 | nitroprusside | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2765267C
804 | nitroprusside | 50 mg / 500 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2765267C
805 | nitroprusside | 100 mg / 250 mL (0.4 mg / mL) | Cont/Bolus or PCA | 2765267C
806 | nitroprusside | --- mg / --- mL | Cont/Bolus or PCA | 2765267C
807 | nivolumab | 240 mg / 100 mL (2.4 mg / mL) | Intermittent | 202016817I
808 | nivolumab | 480 mg / 100 mL (4.8 mg / mL) | Intermittent | 202016817I
809 | nivolumab | --- mg / --- mL | Intermittent | 202016817I
810 | norepinephrine | 4 mg / 250 mL (0.016 mg / mL) | Cont/Bolus or PCA | 2765336C
811 | norepinephrine | 8 mg / 250 mL (0.032 mg / mL) | Cont/Bolus or PCA | 2765336C
812 | norepinephrine | 8 mg / 500 mL (0.016 mg / mL) | Cont/Bolus or PCA | 2765336C
813 | norepinephrine | 16 mg / 250 mL (0.064 mg / mL) | Cont/Bolus or PCA | 2765336C
814 | norepinephrine | 16 mg / 500 mL (0.032 mg / mL) | Cont/Bolus or PCA | 2765336C
815 | norepinephrine | --- mg / --- mL | Cont/Bolus or PCA | 2765336C
816 | obinutuzumab | 100 mg / 100 mL (1 mg / mL) | Intermittent | 25920903I
817 | obinutuzumab | 100 mg / 104 mL (0.9615 mg / mL) | Intermittent | 25920903I
818 | obinutuzumab | 900 mg / 250 mL (3.6 mg / mL) | Intermittent | 25920903I
819 | obinutuzumab | 900 mg / 286 mL (3.1469 mg / mL) | Intermittent | 25920903I
820 | obinutuzumab | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 25920903I
821 | obinutuzumab | 1,000 mg / 290 mL (3.4483 mg / mL) | Intermittent | 25920903I
822 | obinutuzumab | --- mg / --- mL | Intermittent | 25920903I
823 | ocrelizumab | 300 mg / 260 mL (1.1538 mg / mL) | Intermittent | 2581289359I
824 | ocrelizumab | 600 mg / 520 mL (1.1538 mg / mL) | Intermittent | 2581289359I
825 | octreotide | 500 mcg / 100 mL (5 mcg / mL) | Cont/Bolus or PCA | 2765439C
826 | octreotide | 600 mcg / 240 mL (2.5 mcg / mL) | Cont/Bolus or PCA | 2765439C
827 | octreotide | 1,000 mcg / 250 mL (4 mcg / mL) | Cont/Bolus or PCA | 2765439C
828 | octreotide | 1,250 mcg / 250 mL (5 mcg / mL) | Cont/Bolus or PCA | 2765439C
829 | octreotide | --- mcg / --- mL | Cont/Bolus or PCA | 2765439C
830 | ofatumumab | --- mg / --- mL |  | 19954113I
831 | ondansetron | 4 mg / 50 mL (0.08 mg / mL) | Intermittent | 2765642I
832 | ondansetron | 8 mg / 50 mL (0.16 mg / mL) | Intermittent | 2765642I
833 | ondansetron | 8 mg / 54 mL (0.1481 mg / mL) | Intermittent | 2765642I
834 | ondansetron | 32 mg / 50 mL (0.64 mg / mL) | Intermittent | 2765642I
835 | ondansetron | --- mg / --- mL | Intermittent | 2765642I
836 | oritavacin | 1,200 mg / 250 mL (4.8 mg / mL) | Intermittent | 155136453I
837 | oritavacin | 1,200 mg / 1,000 mL (1.2 mg / mL) | Intermittent | 155136453I
838 | oritavacin | 1.2 gram / 250 mL (0.0048 gram / mL) | Intermittent | 155136453I
839 | oritavacin | 1.2 gram / 1,000 mL (0.0012 gram / mL) | Intermittent | 155136453I
840 | oxacillin | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2765734I
841 | oxacillin | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2765734I
842 | oxacillin | --- mg / --- mL | Intermittent | 2765734I
843 | oxacillin | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2765734I
844 | oxacillin | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2765734I
845 | oxaliplatin | --- mg / --- mL | Intermittent | 2765754I
846 | oxytocin | 10 unit / 1,000 mL (0.01 unit / mL) | Cont/Bolus or PCA | 2765930C
847 | oxytocin | 20 unit / 1,000 mL (0.02 unit / mL) | Cont/Bolus or PCA | 2765930C
848 | oxytocin | 30 unit / 500 mL (0.06 unit / mL) | Cont/Bolus or PCA | 2765930C
849 | oxytocin | 30 unit / 1,000 mL (0.03 unit / mL) | Cont/Bolus or PCA | 2765930C
850 | oxytocin | 40 unit / 1,000 mL (0.04 unit / mL) | Cont/Bolus or PCA | 2765930C
851 | oxytocin | --- unit / --- mL | Cont/Bolus or PCA | 2765930C
852 | PACLitaxel | --- mg / --- mL | Intermittent | 2765935I
853 | PACLitaxel protein-b | --- mg / --- mL | Intermittent | 3830361I
854 | pamidronate | 90 mg / 250 mL (0.36 mg / mL) | Intermittent | 2765957I
855 | pamidronate | --- mg / --- mL | Intermittent | 2765957I
856 | pancuronium | --- mg / --- mL | Cont/Bolus or PCA | 2766061C
857 | panitumumab | --- mg / --- mL | Intermittent | 4246362I
858 | pantoprazole | 40 mg / 100 mL (0.4 mg / mL) |  | 2766068
859 | pantoprazole | 80 mg / 100 mL (0.8 mg / mL) |  | 2766068
860 | pantoprazole | 96 mg / 240 mL (0.4 mg / mL) |  | 2766068
861 | pantoprazole | --- mg / --- mL |  | 2766068
862 | pantoprazole bolus | 40 mg / 100 mL (0.4 mg / mL) | Intermittent | 2766068I
863 | pantoprazole bolus | 80 mg / 100 mL (0.8 mg / mL) | Intermittent | 2766068I
864 | pantoprazole drip | 40 mg / 100 mL (0.4 mg / mL) | Cont/Bolus or PCA | 2766068C
865 | pantoprazole drip | 96 mg / 240 mL (0.4 mg / mL) | Cont/Bolus or PCA | 2766068C
866 | pantoprazole drip | --- mg / --- mL | Cont/Bolus or PCA | 2766068C
867 | papaverine | 300 mg / 300 mL (1 mg / mL) |  | 2766098I
868 | papaverine | --- mg / --- mL |  | 2766098I
869 | pegaspargase | --- unit / --- mL |  | 2766160I
870 | pembrolizumab | 200 mg / 100 mL (2 mg / mL) | Intermittent | 114589861I
871 | pembrolizumab | --- mg / --- mL | Intermittent | 114589861I
872 | PEMEtrexed | --- mg / --- mL | Intermittent | 3320387I
873 | penG 2.5 Mill Units | 2.5 unit / 50 mL (0.05 unit / mL) | Intermittent | 19955177I2.5
874 | penG 2.5 Mill Units | 2.5 unit / 100 mL (0.025 unit / mL) | Intermittent | 19955177I2.5
875 | penG 3 Million Units | 3 unit / 50 mL (0.06 unit / mL) | Intermittent | 19955177I-3
876 | penG 5 Million Units | 5 unit / 50 mL (0.1 unit / mL) | Intermittent | 19955177I-5
877 | penG 5 Million Units | 5 unit / 100 mL (0.05 unit / mL) | Intermittent | 19955177
878 | PenG Potass Syringe | --- unit / --- mL | Intermittent | 19955177I
879 | penG x Million Units | --- unit / --- mL | Intermittent | 19955177-x
880 | pentamidine | --- mg / --- mL | Intermittent | 2766284I
881 | PENTobarbital | 2,500 mg / 500 mL (5 mg / mL) |  | 2766302I
882 | PENTobarbital | --- mg / --- mL |  | 2766302I
883 | pentostatin | --- mg / --- mL |  | 2766314I
884 | pertuzumab | 240 mg / 250 mL (0.96 mg / mL) | Intermittent | 24607638I
885 | pertuzumab | 420 mg / 250 mL (1.68 mg / mL) | Intermittent | 24607638I
886 | pertuzumab | 420 mg / 264 mL (1.5909 mg / mL) | Intermittent | 24607638I
887 | pertuzumab | 840 mg / 250 mL (3.36 mg / mL) | Intermittent | 24607638I
888 | pertuzumab | 840 mg / 278 mL (3.0216 mg / mL) | Intermittent | 24607638I
889 | PHENobarbital | 130 mg / 52 mL (2.5 mg / mL) | Intermittent | 2766460I
890 | PHENobarbital | 130 mg / 102 mL (1.2745 mg / mL) | Intermittent | 2766460I
891 | PHENobarbital | 260 mg / 54 mL (4.8148 mg / mL) | Intermittent | 2766460I
892 | PHENobarbital | 260 mg / 104 mL (2.5 mg / mL) | Intermittent | 2766460I
893 | PHENobarbital | --- mg / --- mL | Intermittent | 2766460I
894 | phentolamine | --- mg / --- mL |  | 2766536I
895 | phenylephrine | 10 mg / 250 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2766541C
896 | phenylephrine | 10 mg / 251 mL (0.0398 mg / mL) | Cont/Bolus or PCA | 2766541C
897 | phenylephrine | 20 mg / 250 mL (0.08 mg / mL) | Cont/Bolus or PCA | 2766541C
898 | phenylephrine | 20 mg / 500 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2766541C
899 | phenylephrine | 20 mg / 502 mL (0.0398 mg / mL) | Cont/Bolus or PCA | 2766541C
900 | phenylephrine | 40 mg / 250 mL (0.16 mg / mL) | Cont/Bolus or PCA | 2766541C
901 | phenylephrine | 40 mg / 1,000 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2766541C
902 | phenylephrine | --- mg / --- mL | Cont/Bolus or PCA | 2766541C
903 | phenylephrine ANES | 10 mg / 100 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2766541
904 | phenylephrine ANES | 10 mg / 101 mL (0.099 mg / mL) | Cont/Bolus or PCA | 2766541
905 | phenylephrine ANES | 10 mg / 250 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2766541
906 | phenylephrine ANES | 10 mg / 251 mL (0.0398 mg / mL) | Cont/Bolus or PCA | 2766541
907 | phenylephrine ANES | --- mg / --- mL | Cont/Bolus or PCA | 2766541
908 | phenytoin | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 2766638I
909 | phenytoin | --- mg / --- mL | Intermittent | 2766638I
910 | phytonadione | --- mg / --- mL | Intermittent | 2766694I
911 | piperacillin | 2.25 gram / 50 mL (0.045 gram / mL) |  | 2766817I
912 | piperacillin | 3.37 gram / 50 mL (0.0674 gram / mL) |  | 2766817I
913 | piperacillin | 4.5 gram / 50 mL (0.09 gram / mL) |  | 2766817I
914 | piperacillin/tazo | --- mg / --- mL | Intermittent | 2766819I
915 | piperacillin/tazo | 2.25 gram / 50 mL (0.045 gram / mL) | Intermittent | 2766819I
916 | piperacillin/tazo | 2.25 gram / 100 mL (0.0225 gram / mL) | Intermittent | 2766819I
917 | piperacillin/tazo | 3.375 gram / 50 mL (0.0675 gram / mL) | Intermittent | 2766819I
918 | piperacillin/tazo | 3.375 gram / 100 mL (0.0338 gram / mL) | Intermittent | 2766819I
919 | piperacillin/tazo | 4.5 gram / 50 mL (0.09 gram / mL) | Intermittent | 2766819I
920 | piperacillin/tazo | 4.5 gram / 100 mL (0.045 gram / mL) | Intermittent | 2766819I
921 | piperacillin/tazo | --- gram / --- mL | Intermittent | 2766819I
922 | plicamycin | --- mcg / --- mL |  | 2766901I
923 | polatuzumab vedotin | --- mg / --- mL | Intermittent | 3173912591I
924 | porfirmer | --- mg / --- mL |  | 2767007I
925 | potassium acetate | 10 mEq / 100 mL (0.1 mEq / mL) | Intermittent | 2767012I
926 | potassium acetate | 20 mEq / 100 mL (0.2 mEq / mL) | Intermittent | 2767012I
927 | potassium acetate | 30 mEq / 100 mL (0.3 mEq / mL) | Intermittent | 2767012I
928 | potassium acetate | 40 mEq / 100 mL (0.4 mEq / mL) | Intermittent | 2767012I
929 | potassium acetate | 40 mEq / 520 mL (0.0769 mEq / mL) | Intermittent | 2767012I
930 | potassium acetate | --- mEq / --- mL | Intermittent | 2767012I
931 | potassium chloride | 10 mEq / 50 mL (0.2 mEq / mL) | Intermittent | 1000000100I
932 | potassium chloride | 10 mEq / 100 mL (0.1 mEq / mL) | Intermittent | 1000000100I
933 | potassium chloride | 20 mEq / 100 mL (0.2 mEq / mL) | Intermittent | 1000000100I
934 | potassium chloride | 40 mEq / 100 mL (0.4 mEq / mL) | Intermittent | 1000000100I
935 | potassium chloride | 40 mEq / 520 mL (0.0769 mEq / mL) | Intermittent | 1000000100I
936 | potassium chloride | --- mEq / --- mL | Intermittent | 1000000100I
937 | potassium phosphate | --- mEq / --- mL | Intermittent | 2767185I
938 | potassium phosphate | --- mmol / --- mL | Intermittent | 2767185I
939 | pralidoxime | --- mg / --- mL | Intermittent | 2767273I
940 | pralidoxime | --- gram / --- mL | Intermittent | 2767273I
941 | procainamide | 1,000 mg / 50 mL (20 mg / mL) | Cont/Bolus or PCA | 2767500C
942 | procainamide | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA | 2767500C
943 | procainamide | 2,000 mg / 250 mL (8 mg / mL) | Cont/Bolus or PCA | 2767500C
944 | procainamide | 2,000 mg / 500 mL (4 mg / mL) | Cont/Bolus or PCA | 2767500C
945 | prochlorperazine | --- mg / --- mL |  | 2767546I
946 | promethazine | 25 mg / 51 mL (0.4902 mg / mL) | Intermittent | 2767590I
947 | promethazine | --- mg / --- mL | Intermittent | 2767590I
948 | propofol | 500 mg / 50 mL (10 mg / mL) | Cont/Bolus or PCA | 2767660C
949 | propofol | 1,000 mg / 100 mL (10 mg / mL) | Cont/Bolus or PCA | 2767660C
950 | propofol | --- mg / --- mL | Cont/Bolus or PCA | 2767660C
951 | propofol ANES | 10 mg / 1 mL (10 mg / mL) | Cont/Bolus or PCA | 2767660I
952 | propofol ANES | 500 mg / 50 mL (10 mg / mL) | Cont/Bolus or PCA | 2767660I
953 | propofol ANES | 1,000 mg / 100 mL (10 mg / mL) | Cont/Bolus or PCA | 2767660I
954 | propofol ANES | --- mg / --- mL | Cont/Bolus or PCA | 2767660I
955 | protamine sulfate | --- mg / --- mL | Intermittent | 2767711I
956 | quiNIDine | --- mg / --- mL |  | 2767995I
957 | ramucirumab | --- mg / --- mL | Intermittent | 38821169I
958 | raNITIdine | 50 mg / 50 mL (1 mg / mL) | Intermittent | 2768055I
959 | raNITIdine | 50 mg / 100 mL (0.5 mg / mL) | Intermittent | 2768055I
960 | raNITIdine | 50 mg / 102 mL (0.4902 mg / mL) | Intermittent | 2768055I
961 | raNITIdine | 150 mg / 250 mL (0.6 mg / mL) | Intermittent | 2768055I
962 | raNITIdine | 250 mg / 250 mL (1 mg / mL) | Intermittent | 2768055I
963 | raNITIdine | --- mg / --- mL | Intermittent | 2768055I
964 | rasburicase | --- mg / --- mL |  | 2768080I
965 | remdesivir | 100 mg / 230 mL (0.4348 mg / mL) | Intermittent | 3520886709I
966 | remdesivir | 100 mg / 250 mL (0.4 mg / mL) | Intermittent | 3520886709I
967 | remdesivir | 200 mg / 210 mL (0.9524 mg / mL) | Intermittent | 3520886709I
968 | remdesivir | 200 mg / 250 mL (0.8 mg / mL) | Intermittent | 3520886709I
969 | remifentanil | 50 mcg / 1 mL (50 mcg / mL) |  | 2768094
970 | remifentanil | --- mcg / --- mL |  | 2768094
971 | remifentanil ANES | 50 mcg / 1 mL (50 mcg / mL) | Cont/Bolus or PCA | 2768094C
972 | remifentanil ANES | 1,000 mcg / 40 mL (25 mcg / mL) | Cont/Bolus or PCA | 2768094C
973 | remifentanil ANES | 2,000 mcg / 80 mL (25 mcg / mL) | Cont/Bolus or PCA | 2768094C
974 | remifentanil ANES | --- mcg / --- mL | Cont/Bolus or PCA | 2768094C
975 | reteplase | --- unit / --- mL |  | 2768131I
976 | rifAMPin | 300 mg / 250 mL (1.2 mg / mL) | Intermittent | 2768179I
977 | rifAMPin | 600 mg / 500 mL (1.2 mg / mL) | Intermittent | 2768179I
978 | rifAMPin | --- mg / --- mL | Intermittent | 2768179I
979 | ritodrine | --- mg / --- mL |  | 2768244I
980 | riTUXimab | 500 mg / 250 mL (2 mg / mL) | Intermittent | 2768259I
981 | riTUXimab | 1,000 mg / 500 mL (2 mg / mL) | Intermittent | 2768259I
982 | riTUXimab | --- mg / --- mL | Intermittent | 2768259I
983 | rocuronium | 500 mg / 500 mL (1 mg / mL) | Cont/Bolus or PCA | 2768289C
984 | rocuronium | --- mg / --- mL | Cont/Bolus or PCA | 2768289C
985 | rocuronium ANES | 500 mg / 500 mL (1 mg / mL) | Cont/Bolus or PCA | 2768289I
986 | rocuronium ANES | --- mg / --- mL | Cont/Bolus or PCA | 2768289I
987 | romiDEPsin | --- mg / --- mL |  | 19955002I
988 | ropivacaine | 200 mg / 100 mL (2 mg / mL) |  | 2768324C
989 | ropivacaine | 400 mg / 200 mL (2 mg / mL) |  | 2768324C
990 | ropivacaine | --- mg / --- mL |  | 2768324C
991 | salbutamol | --- mg / --- mL |  | salbutamol
992 | sargramostim | --- mcg / --- mL |  | 2768593I
993 | siltuximab | --- mg / --- mL |  | 45868529I
994 | SMX/TMP | --- mg / --- mL | Intermittent | 2769391I
995 | sodium bicarbonate | --- mEq / --- mL |  | 2768888
996 | sodium chloride 3% | 256.6 mEq / 500 mL (0.5132 mEq / mL) |  | 2778727
997 | sodium chloride 5% | 427.5 mEq / 500 mL (0.855 mEq / mL) |  | 2778729
998 | sodium phosphate | --- mmol / --- mL | Intermittent | 2769071I
999 | sotrovimab | 500 mg / 50 mL (10 mg / mL) | Intermittent | 4112134745I
1000 | sotrovimab | 500 mg / 100 mL (5 mg / mL) | Intermittent | 4112134745I
1001 | spesolimab | 900 mg / 100 mL (9 mg / mL) | Intermittent | 4788045753I
1002 | streptozocin | --- mg / --- mL |  | 2769343I
1003 | succinylcholine | --- mg / --- mL |  | 2769352I
1004 | SUFentanil | --- mcg / --- mL |  | 2769370I
1005 | SUFentanil ANES | 100 mcg / 100 mL (1 mcg / mL) | Cont/Bolus or PCA | 2769370C
1006 | SUFentanil ANES | --- mcg / --- mL | Cont/Bolus or PCA | 2769370C
1007 | tacrolimus | --- mg / --- mL |  | 2769525I
1008 | tedizolid phosphate | 200 mg / 250 mL (0.8 mg / mL) | Intermittent | 45868623I
1009 | temozolomide | --- mg / --- mL | Intermittent | 2769599I
1010 | temsirolimus | --- mg / --- mL |  | 4246596I
1011 | teniposide | --- mg / --- mL |  | 2769615I
1012 | teprotumumab | --- mg / --- mL | Intermittent | 3602321791I
1013 | teprotumumab | --- gram / --- mL | Intermittent | 3602321791I
1014 | terbutaline | --- mg / --- mL |  | 2769660I
1015 | theophylline | 400 mg / 500 mL (0.8 mg / mL) |  | 2769843I
1016 | theophylline | 800 mg / 500 mL (1.6 mg / mL) |  | 2769843I
1017 | theophylline | --- mg / --- mL |  | 2769843I
1018 | thiamine | 100 mg / 50 mL (2 mg / mL) | Intermittent | 2769929I
1019 | thiamine | 100 mg / 100 mL (1 mg / mL) | Intermittent | 2769929I
1020 | thiamine | --- mg / --- mL | Intermittent | 2769929I
1021 | thiopental | --- mg / --- mL |  | 2769954I
1022 | thiotepa | --- mg / --- mL |  | 2769992I
1023 | ticarcillin | 3.1 gram / 100 mL (0.031 gram / mL) |  | 2770096I
1024 | ticarcillin/clav | --- mg / --- mL | Intermittent | 2770103I
1025 | ticarcillin/clav | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2770103I
1026 | ticarcillin/clav | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2770103I
1027 | ticarcillin/clav | 3.1 gram / 50 mL (0.062 gram / mL) | Intermittent | 2770103I
1028 | ticarcillin/clav | 3.1 gram / 100 mL (0.031 gram / mL) | Intermittent | 2770103I
1029 | tigecycline | 50 mg / 50 mL (1 mg / mL) | Intermittent | 3831121I
1030 | tigecycline | 50 mg / 100 mL (0.5 mg / mL) | Intermittent | 3831121I
1031 | tigecycline | 100 mg / 100 mL (1 mg / mL) | Intermittent | 3831121I
1032 | tirofiban | 12.5 mg / 250 mL (0.05 mg / mL) | Cont/Bolus or PCA | 2770172C
1033 | tobramycin | 80 mg / 50 mL (1.6 mg / mL) | Intermittent | 2770186I
1034 | tobramycin | 80 mg / 100 mL (0.8 mg / mL) | Intermittent | 2770186I
1035 | tobramycin | 120 mg / 100 mL (1.2 mg / mL) | Intermittent | 2770186I
1036 | tobramycin | --- mg / --- mL | Intermittent | 2770186I
1037 | tocilizumab | 800 mg / 100 mL (8 mg / mL) | Intermittent | 19955605I
1038 | tocilizumab | --- mg / --- mL | Intermittent | 19955605I
1039 | tolazoline | --- mg / --- mL |  | tolazoline
1040 | topotecan | --- mg / --- mL | Intermittent | 2770314I
1041 | torsemide | --- mg / --- mL |  | 2770324I
1042 | trabectedin | --- mg / --- mL | Intermittent | 275110021I
1043 | tranexamic BOLUS | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2770412I
1044 | tranexamic BOLUS | --- mg / --- mL | Intermittent | 2770412I
1045 | tranexamic BOLUS | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2770412I
1046 | tranexamic BOLUS | --- gram / --- mL | Intermittent | 2770412I
1047 | trastuzumab | --- mg / --- mL | Intermittent | 2770424I
1048 | trimetrexate | --- mg / --- mL |  | 2770672I
1049 | tromethamine | --- mEq / --- mL |  | 2770725I
1050 | urokinase | --- unit / --- mL |  | 2770879I
1051 | ustekinumab | 260 mg / 250 mL (1.04 mg / mL) | Intermittent | 19955913I
1052 | ustekinumab | 390 mg / 250 mL (1.56 mg / mL) | Intermittent | 19955913I
1053 | ustekinumab | 520 mg / 250 mL (2.08 mg / mL) | Intermittent | 19955913I
1054 | valproate sodium | --- mg / --- mL | Intermittent | 2770918I
1055 | valproic acid | --- mg / --- mL |  | 2770918
1056 | vancomycin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2770944I
1057 | vancomycin | 750 mg / 150 mL (5 mg / mL) | Intermittent | 2770944I
1058 | vancomycin | 750 mg / 250 mL (3 mg / mL) | Intermittent | 2770944I
1059 | vancomycin | 1,000 mg / 200 mL (5 mg / mL) | Intermittent | 2770944I
1060 | vancomycin | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 2770944I
1061 | vancomycin | 1,250 mg / 250 mL (5 mg / mL) | Intermittent | 2770944I
1062 | vancomycin | 1,500 mg / 300 mL (5 mg / mL) | Intermittent | 2770944I
1063 | vancomycin | 1,500 mg / 500 mL (3 mg / mL) | Intermittent | 2770944I
1064 | vancomycin | 1,750 mg / 350 mL (5 mg / mL) | Intermittent | 2770944I
1065 | vancomycin | 1,750 mg / 500 mL (3.5 mg / mL) | Intermittent | 2770944I
1066 | vancomycin | 2,000 mg / 400 mL (5 mg / mL) | Intermittent | 2770944I
1067 | vancomycin | 2,000 mg / 500 mL (4 mg / mL) | Intermittent | 2770944I
1068 | vancomycin | --- mg / --- mL | Intermittent | 2770944I
1069 | vancomycin | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2770944I
1070 | vancomycin | 0.75 gram / 150 mL (0.005 gram / mL) | Intermittent | 2770944I
1071 | vancomycin | 0.75 gram / 250 mL (0.003 gram / mL) | Intermittent | 2770944I
1072 | vancomycin | 1 gram / 200 mL (0.005 gram / mL) | Intermittent | 2770944I
1073 | vancomycin | 1 gram / 250 mL (0.004 gram / mL) | Intermittent | 2770944I
1074 | vancomycin | 1.25 gram / 250 mL (0.005 gram / mL) | Intermittent | 2770944I
1075 | vancomycin | 1.5 gram / 300 mL (0.005 gram / mL) | Intermittent | 2770944I
1076 | vancomycin | 1.5 gram / 500 mL (0.003 gram / mL) | Intermittent | 2770944I
1077 | vancomycin | 1.75 gram / 350 mL (0.005 gram / mL) | Intermittent | 2770944I
1078 | vancomycin | 1.75 gram / 500 mL (0.0035 gram / mL) | Intermittent | 2770944I
1079 | vancomycin | 2 gram / 400 mL (0.005 gram / mL) | Intermittent | 2770944I
1080 | vancomycin | 2 gram / 500 mL (0.004 gram / mL) | Intermittent | 2770944I
1081 | vancomycin | --- gram / --- mL | Intermittent | 2770944I
1082 | vasopressin | 0.02 unit / 1 mL (0.02 unit / mL) | Cont/Bolus or PCA | 2770987C
1083 | vasopressin | 0.04 unit / 1 mL (0.04 unit / mL) | Cont/Bolus or PCA | 2770987C
1084 | vasopressin | 1 unit / 1 mL (1 unit / mL) | Cont/Bolus or PCA | 2770987C
1085 | vasopressin | 4 unit / 100 mL (0.04 unit / mL) | Cont/Bolus or PCA | 2770987C
1086 | vasopressin | 20 unit / 100 mL (0.2 unit / mL) | Cont/Bolus or PCA | 2770987C
1087 | vasopressin | 40 unit / 100 mL (0.4 unit / mL) | Cont/Bolus or PCA | 2770987C
1088 | vasopressin | 80 unit / 100 mL (0.8 unit / mL) | Cont/Bolus or PCA | 2770987C
1089 | vasopressin | 100 unit / 100 mL (1 unit / mL) | Cont/Bolus or PCA | 2770987C
1090 | vasopressin | --- unit / --- mL | Cont/Bolus or PCA | 2770987C
1091 | vecuronium | 1 mg / 1 mL (1 mg / mL) | Cont/Bolus or PCA | 2770992C
1092 | vecuronium | 2 mg / 1 mL (2 mg / mL) | Cont/Bolus or PCA | 2770992C
1093 | vecuronium | 20 mg / 50 mL (0.4 mg / mL) | Cont/Bolus or PCA | 2770992C
1094 | vecuronium | 25 mg / 50 mL (0.5 mg / mL) | Cont/Bolus or PCA | 2770992C
1095 | vecuronium | 30 mg / 50 mL (0.6 mg / mL) | Cont/Bolus or PCA | 2770992C
1096 | vecuronium | 50 mg / 50 mL (1 mg / mL) | Cont/Bolus or PCA | 2770992C
1097 | vecuronium | 50 mg / 100 mL (0.5 mg / mL) | Cont/Bolus or PCA | 2770992C
1098 | vecuronium | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2770992C
1099 | vecuronium | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2770992C
1100 | vecuronium | --- mg / --- mL | Cont/Bolus or PCA | 2770992C
1101 | vedolizumab | 300 mg / 250 mL (1.2 mg / mL) | Intermittent | 45868365I
1102 | verapamil | --- mg / --- mL |  | 2771023I
1103 | vidarabine | --- mg / --- mL |  | vidarabine
1104 | vinBLAStine | --- mg / --- mL | Intermittent | 2771078I
1105 | vinCRIStine | --- mg / --- mL | Intermittent | 2771083I
1106 | vinCRIStine Liposom | --- mg / --- mL |  | 24608273I
1107 | vinorelbine | --- mg / --- mL | Intermittent | 2771090I
1108 | voriconazole | 200 mg / 50 mL (4 mg / mL) | Intermittent | 2771180I
1109 | voriconazole | 400 mg / 100 mL (4 mg / mL) | Intermittent | 2771180I
1110 | voriconazole | --- mg / --- mL | Intermittent | 2771180I
1111 | zidovudine | --- mg / --- mL | Intermittent | 2771278I
1112 | ziv-aflibercept | --- mg / --- mL | Intermittent | 24608076I
1113 | zoledronic acid | 4 mg / 100 mL (0.04 mg / mL) | Intermittent | 2771383I
1114 | zoledronic acid | 5 mg / 100 mL (0.05 mg / mL) | Intermittent | 2771383I
1115 | zoledronic acid | --- mg / --- mL | Intermittent | 2771383I
1116 | zz amiodarone BOLUS | 150 mg / 100 mL (1.5 mg / mL) | Intermittent
1117 | zz amiodarone DRIP | 360 mg / 200 mL (1.8 mg / mL) | Cont/Bolus or PCA
1118 | zz amiodarone DRIP | 900 mg / 500 mL (1.8 mg / mL) | Cont/Bolus or PCA
1119 | zz DOBUTamine | 2,000 mcg / 1 mL (2,000 mcg / mL) | Cont/Bolus or PCA
1120 | zz DOBUTamine | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA
1121 | zz DOBUTamine | 500 mg / 250 mL (2 mg / mL) | Cont/Bolus or PCA
1122 | zz DOBUTamine | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA
1123 | zz DOBUTamine | --- mg / --- mL | Cont/Bolus or PCA
1124 | zz DOPamine | 1,600 mcg / 1 mL (1,600 mcg / mL) | Cont/Bolus or PCA
1125 | zz DOPamine | 400 mg / 250 mL (1.6 mg / mL) | Cont/Bolus or PCA
1126 | zz DOPamine | 800 mg / 250 mL (3.2 mg / mL) | Cont/Bolus or PCA
1127 | zz DOPamine | --- mg / --- mL | Cont/Bolus or PCA
1128 | zz EPINEPHrine | 10 mcg / 1 mL (10 mcg / mL) | Cont/Bolus or PCA
1129 | zz EPINEPHrine | 4 mg / 250 mL (0.016 mg / mL) | Cont/Bolus or PCA
1130 | zz EPINEPHrine | 5 mg / 250 mL (0.02 mg / mL) | Cont/Bolus or PCA
1131 | zz EPINEPHrine | 8 mg / 250 mL (0.032 mg / mL) | Cont/Bolus or PCA
1132 | zz EPINEPHrine | 10 mg / 100 mL (0.1 mg / mL) | Cont/Bolus or PCA
1133 | zz EPINEPHrine | 10 mg / 500 mL (0.02 mg / mL) | Cont/Bolus or PCA
1134 | zz EPINEPHrine | 16 mg / 250 mL (0.064 mg / mL) | Cont/Bolus or PCA
1135 | zz lidocaine | 8 mg / 1 mL (8 mg / mL) | Cont/Bolus or PCA
1136 | zz lidocaine | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA
1137 | zz lidocaine | 2,000 mg / 500 mL (4 mg / mL) | Cont/Bolus or PCA
1138 | zz lidocaine | --- mg / --- mL | Cont/Bolus or PCA
1139 | zz lidocaine | 2 gram / 500 mL (0.004 gram / mL) | Cont/Bolus or PCA
1140 | zz mag sulfate | --- mg / --- mL | Intermittent
1141 | zz mag sulfate | 1 gram / 50 mL (0.02 gram / mL) | Intermittent
1142 | zz mag sulfate | 1 gram / 100 mL (0.01 gram / mL) | Intermittent
1143 | zz mag sulfate | 2 gram / 50 mL (0.04 gram / mL) | Intermittent
1144 | zz mag sulfate | 4 gram / 100 mL (0.04 gram / mL) | Intermittent
1145 | zz mag sulfate | 40 gram / 500 mL (0.08 gram / mL) | Intermittent
1146 | zz mag sulfate | --- gram / --- mL | Intermittent
1147 | zz nitroglycerin | 25 mg / 250 mL (0.1 mg / mL) | Cont/Bolus or PCA
1148 | zz nitroglycerin | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA
1149 | zz nitroglycerin | --- mg / --- mL | Cont/Bolus or PCA
1150 | zz nitroprusside | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA
1151 | zz nitroprusside | 100 mg / 250 mL (0.4 mg / mL) | Cont/Bolus or PCA
1152 | zz nitroprusside | --- mg / --- mL | Cont/Bolus or PCA
1153 | zz norepinephrine | 16 mcg / 1 mL (16 mcg / mL) | Cont/Bolus or PCA
1154 | zz norepinephrine | 4 mg / 250 mL (0.016 mg / mL) | Cont/Bolus or PCA
1155 | zz norepinephrine | 8 mg / 250 mL (0.032 mg / mL) | Cont/Bolus or PCA
1156 | zz norepinephrine | 8 mg / 500 mL (0.016 mg / mL) | Cont/Bolus or PCA
1157 | zz norepinephrine | 16 mg / 250 mL (0.064 mg / mL) | Cont/Bolus or PCA
1158 | zz norepinephrine | 16 mg / 500 mL (0.032 mg / mL) | Cont/Bolus or PCA
1159 | zz norepinephrine | --- mg / --- mL | Cont/Bolus or PCA
1160 | zz phenylephrine | 10 mg / 250 mL (0.04 mg / mL) | Cont/Bolus or PCA
1161 | zz phenylephrine | 20 mg / 500 mL (0.04 mg / mL) | Cont/Bolus or PCA
1162 | zz phenylephrine | 40 mg / 1,000 mL (0.04 mg / mL) | Cont/Bolus or PCA

## Master Fluids
Master Fluids | Master Fluids | Master Fluids | Master Fluids
Line Item | Fluid Name | Supports Secondary | Aliases
1 | ..IVF Maintenance | Yes | 1000000000F,1000000000I
2 | .Bolus IVF | Yes
3 | .IVF + KCL | Yes | 1000000100F
4 | .IVF + KCL 10 mEq/L | Yes | 1000000010F
5 | .IVF + KCL 20 mEq/L | Yes | 1000000020F
6 | .IVF + KCL 30 mEq/L | Yes
7 | .IVF + KCL 40 mEq/L | Yes | 1000000040F
8 | .IVF + Oxytocin | Yes
9 | .IVF + Vitamins | Yes
10 | .Maint IVF + KCL | Yes
11 | .Maint IVF+Additives | Yes
12 | 0.45%NS+KCL10 mEq/L | Yes
13 | 0.45%NS+KCL20 mEq/L | Yes
14 | 0.45%NS+KCL30 mEq/L | Yes
15 | 0.45%NS+KCL40 mEq/L | Yes
16 | 0.9% Normal Saline | Yes
17 | 0.9%NS+KCL 10mEq/L | Yes
18 | 0.9%NS+KCL 20mEq/L | Yes
19 | 0.9%NS+KCL 30mEq/L | Yes
20 | 0.9%NS+KCL 40mEq/L | Yes
21 | 1/2NS Hep 0.5unit/mL | Yes
22 | 1/4NS Hep 0.5unit/mL | Yes
23 | 3% Sodium Chloride | Yes | 2778727F,2778727I
24 | Bicarb Drip | Yes | 2768888F,2768888I
25 | Blood (RBCs) | No
26 | Blood (whole) | No
27 | Blood Reinfusion | No
28 | Cryoprecipitate | No
29 | D10W 1/4NS | Yes
30 | D10W 1/4NS Hep 0.5un | Yes
31 | D10W Hep 0.5units/mL | Yes
32 | D5 + Normal Saline | Yes
33 | D5 1/2Normal Saline | Yes
34 | D5 1/2NS+KCL10mEq/L | Yes
35 | D5 1/2NS+KCL20mEq/L | Yes
36 | D5 1/2NS+KCL30mEq/L | Yes
37 | D5 1/2NS+KCL40mEq/L | Yes
38 | D5 1/3Normal Saline | Yes
39 | D5 1/3NS+KCL10mEq/L | Yes
40 | D5 1/3NS+KCL20mEq/L | Yes
41 | D5 1/3NS+KCL30mEq/L | Yes
42 | D5 1/3NS+KCL40mEq/L | Yes
43 | D5 1/4Normal Saline | Yes
44 | D5 1/4NS Hep 0.5unit | Yes
45 | D5 1/4NS+KCL10mEq/L | Yes
46 | D5 1/4NS+KCL20mEq/L | Yes
47 | D5 1/4NS+KCL30mEq/L | Yes
48 | D5 1/4NS+KCL40mEq/L | Yes
49 | D5 Lactated Ringers | Yes
50 | D5NS+KCL 10 mEq/L | Yes
51 | D5NS+KCL 20 mEq/L | Yes
52 | D5NS+KCL 30 mEq/L | Yes
53 | D5NS+KCL 40 mEq/L | Yes
54 | D5W | Yes
55 | D5W Hep 0.5units/mL | Yes
56 | D5W+KCL 10 mEq/L | Yes
57 | D5W+KCL 20 mEq/L | Yes
58 | D5W+KCL 30 mEq/L | Yes
59 | D5W+KCL 40 mEq/L | Yes
60 | Dextrose 10% | Yes | 2778755F,2778755I
61 | Dextrose in Ringers | Yes
62 | FLUSH | Yes
63 | Fresh Frozen Plasma | No
64 | Hetastarch 6% NS | Yes
65 | Infliximab(Remicade) | Yes
66 | Insulin Maternity | Yes | 2761029F
67 | IVF +Hep 0.5units/mL | Yes
68 | Lactated Ringers | Yes
69 | Lipids 20% | No
70 | Mannitol | No
71 | NS w Hep 0.5units/mL | Yes
72 | Parenteral Nutrition | No | 1000000001F,1000000001I
73 | Platelets | No
74 | PPN | No
75 | PRBCs | No
76 | Replacement Fluid | Yes
77 | Rituximab (Rituxan) | Yes
78 | TPN | No
79 | UAC | No
80 | UVC | No

## Master Clinical Advisories
Master Clinical Advisories | Master Clinical Advisories | Master Clinical Advisories
Line Item | Advisory Name | Advisory Text
1 | 0.2 to 5 micron filt | Use a low-protein binding in-line or add-on filter between 0.2 and 5 microns.
2 | 0.22 Filter Needed | Use a 0.22 micron in-line filter when administering this product.
3 | 1 unit = 1 mL | 1 unit = 1 mL
4 | 12 Hour Tubing | Tubing must be changed every 12 hours
5 | 1unit=1million | 1 unit = 1 million units
6 | 2nd RN to verify | 2 RNs to verify and document concentration and programming prior to infusion.
7 | 48 Hour Tubing | Change tubing every 48 hours.
8 | Alaris Unit | 1 Unit = 1 million ex. 0.6units = 600,000 units
9 | Allergic Reaction | Infuse slowly for the first 10 minutes (25-50 mL/hr) to watch for allergic reaction, if none, rate may be increased to 250 mL/hr.
10 | Alteplase wt based | Dose and volume to be infused need to be set according to weight based protocol to avoid potential over infusion of entire contents of bag.
11 | AmBisome | Existing IV lines must be flushed with D5w prior to infusion. If unable, must start a new line. **An in-fine filter >= 1 micron may be used with infusion.
12 | Anaphylaxis Warning | Observe patient for 1 hour after starting infusion for signs and symptoms of anaphylaxis.
13 | Antivenin | 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
14 | Bleeding Precautions | Monitor for signs of bleeding.
15 | Calcium Precipitate | Do not infuse Calcium and Phosphate though the same line.
16 | CeftriSyringe | Adminster over 60 minutes to decrease risk of bilirubin encephalopathy.
17 | Central Line Only | Run through a CENTRAL line only.
18 | Cerebyx | Not to be infused faster than 150 mgPE / min, due to risk of hypotension.
19 | Check Concentration | *CHECK CONCENTRATION* 2nd RN to verify and document concentration and programming prior to starting infusion.
20 | Code only drugs | This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
21 | Dedicated IV Line | Do not infuse any other drugs through this line.
22 | DEHP | This medication must be infused using DEHP FREE tubing.
23 | Double Concentrated | ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
24 | EKG Required | EKG Montoring required for infusions greater than 10mEq per hour.
25 | Filter and DEHP Free | Requires Low protein binding 0.2 micron filter and DEHP free bag
26 | Fluid Bolus | Administer a fluid bolus prior to starting the cidofovir infusion.
27 | Flush After NS-LR | Following infusion, flush with 30 mL of sterile normal saline or lactated ringers.
28 | Flush Before/After | If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w or NS.
29 | Flush Before/AfterD5 | If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
30 | Flush Before/AfterNS | Flush line before and after administration. Can be given on same line as sodium chloride 0.9%.
31 | Heparin | Monitor PTT. Do not administer other drugs through same IV line.
32 | High Risk | This is a HIGH RISK medication. 2nd RN to verify prior to initiating therapy.
33 | High Risk Central | This is a HIGH RISK medication that should be infused via a CENTRAL LINE. 2nd RN to verify prior to initiating therapy.
34 | Hypersensitivity | Risk of hypersensitivity. Check vital signs frequently during infusion.
35 | IFEX | Dose is based on Ifosfamide (IFEX) content
36 | In-line filter | Infusion requires the use of an in-line filter.
37 | INR > 6 | INR > 6 ** Max dose of Kcentra should not exceed 5000 units **
38 | INR 2-4 | INR 2 - 4 ** Max dose of Kcentra should not exceed 2500 units **
39 | INR 4-6 | INR 4 - 6 ** Max dose of Kcentra should not exceed 3500 units **
40 | Insulin Maternity | 75 ml/hr = 0.75 u/hr 100 ml/hr = 1 u/hr 125 ml/hr = 1.25 u/hr 150 ml/hr = 1.5 u/hr
41 | Levophed | *****Must have physician order for rates greater than 60 mcg/min*****
42 | Low Binding 0.2or1.2 | This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
43 | Low Binding Filter | This medication must be infused using a 0.2- micron, low protein binding filter.
44 | Low Binding Tubing | * Use Low Binding Tubing *
45 | MAX DOSE 180 mg | MAXIMUM dose of brentuximab is 180 mg.
46 | Midline Warning | Do NOT infuse through a midline catheter.
47 | Monitoring required | Monitoring module required.
48 | MVI | 1 unit = 1 dose
49 | Neuromuscular BA | *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
50 | No Dextrose | Do NOT administer through lines containing dextrose.
51 | NO FILTER | Do NOT use an in-line filter for administration.
52 | Non-PVC | This medication must be infused using Non-PVC tubing.
53 | NS only | Do not dilute with D5W Dilute with NS 0.9% only
54 | oritavacin | Contraindicated to run ****Heparin**** within 48 hours of Oritavacin administration. May increase risk of bleeding with oral anticoagulant therapy.
55 | Patient Lie Down | Advise patient to lie down during and 30 minutes after infusion to prevent hypotension.
56 | PenG-Syringe | 1 unit in pump = 1,000 units on label. Eg. 200 units in pump = 200,000 units on drug label.
57 | Phenylephrine | *****Must have physician order for rates greater than 300 mcg/min*****
58 | Premedicate | Premedicate before each infusion.
59 | Protect From Light | Protect this drug from light.
60 | SCr 2.0-4.0 | SCr 2.0 - 4.0
61 | SCr<2.0 | SCr < 2.0
62 | Siltuximab | Use PVC with DEHP or PU containing administration set containing a 0.2 micron filter.
63 | Stress Test | For STRESS TEST only.
64 | Vesicant | Drug is a vesicant. Check blood return.
65 | Vesicant No Filter | Drug is a vesicant. Check blood return. Do not use an in-line filter.
66 | Vincristine Lipo | Use dedicated line. Do not filter.

## Master Therapies
Master Therapies | Master Therapies
Line Item | Therapy
1 | ..Less than 3 kg
2 | ..Standard
3 | ..Std. Cardiac/PVD
4 | .Double Concentrated
5 | .Induction
6 | .INR 2 to < 4
7 | .INR 4 - 6
8 | .MG Dosing
9 | .mg, pt weight <40kg
10 | .mg, pt weight>=40kg
11 | .Standard & wt>121kg
12 | .Standard Dose
13 | .Standard NON STEMI
14 | .STANDARD STEMI PCI
15 | .Standard wt < 115kg
16 | .Standard wt >=50 kg
17 | .Std Cardiac or PVD
18 | .Test Dose
19 | .VTE (DVT, PE)
20 | .VTE/DVT/PE
21 | .Weight >= 30 kg
22 | .wt <= 100 kg
23 | .wt > 100 kg
24 | .wt less than 100kg
25 | 0.2 mg / 1 mL
26 | 1 mg/1mL
27 | 1 mg/mL *HIGH CONC*
28 | 10 mEq/L
29 | 15 mEq/L
30 | 1st Dose <=100kg
31 | 1st Dose >100kg
32 | 20 mEq/L
33 | 2nd Dose <=100kg
34 | 2nd Dose >100kg
35 | 3 - 9.9 kg
36 | 30 mEq/L
37 | 3rd Dose <=100kg
38 | 3rd Dose >100kg
39 | 40 mEq/L
40 | 5 mg/1mL *HIGH CONC*
41 | Anavip
42 | Anesthesia
43 | Anti-Convulsant
44 | Anticonvulsant Thera
45 | AUC 2-3 q week
46 | AUC 3-7.5 q3weeks
47 | Based on AUC
48 | Based on BSA
49 | Bleeding
50 | Bleeding/Hemophilia
51 | BSA Based Dose
52 | Cardiac Arrest
53 | CENTRAL < 20kg
54 | CENTRAL >=20kg
55 | Central IV Occlusion
56 | Central Line
57 | Cerebrovascular
58 | CrCl <= 60
59 | CrCl > 60
60 | Crofab
61 | Cycle 1
62 | Cycle 2
63 | Cyclic
64 | Diabetes Insipidus
65 | DVT / Ischemic Limb
66 | Echo Stress Test
67 | ESRD
68 | Every 2 weeks
69 | Every 3 weeks
70 | First Dose
71 | GI Bleed
72 | GM dosing
73 | gram, pt wt < 40kg
74 | gram, pt wt >= 40kg
75 | Hepatic Dose
76 | High Dose
77 | High Dose - GRAMS
78 | Induction
79 | INR > 6
80 | Insulin Maternity
81 | Kimsyra gram
82 | Kimyrsa mg
83 | Load <67kg WtBasedGM
84 | Load <67kg WtBasedMG
85 | Load >67kg WtBasedGM
86 | Load >67kg WtBasedMG
87 | Load NonWtBased MG
88 | Load NonWtBasedGM
89 | Loading Dose
90 | LoDose & >= 153 kg
91 | LoDose & CrCl < 30
92 | LoDose & CrCl >= 30
93 | LoDose,Renal&>=153kg
94 | Maintenance Dose
95 | Maintenance Dose GM
96 | Maintenance Dose MG
97 | Non-Anesthesia
98 | Non-Based Based Dose
99 | OB/L&D
100 | Orbactiv gram
101 | Orbactiv mg
102 | PERIPH < 20kg
103 | PERIPH >=20kg
104 | Peripheral Line
105 | Postpartum
106 | pt age < = 12 years
107 | pt age > = 12 years
108 | pt BSA < 1 m2
109 | pt BSA >=1
110 | pt weight < 20 kg
111 | pt weight < 30 kg
112 | pt weight < 40 kg
113 | pt weight < 50 kg
114 | pt weight < 60 kg
115 | pt weight >= 20 kg
116 | pt weight >= 30 kg
117 | pt weight >= 40 kg
118 | pt weight >= 50 kg
119 | pt weight >= 60 kg
120 | Pulmonary Embolism
121 | Renal Dose
122 | Renal Dose CrCl < 30
123 | Renal Dose wt >121kg
124 | Renal NON STEMI
125 | Renal STEMI PCI
126 | Rheumatoid Arthritis
127 | Second Dose
128 | Sedation
129 | Shock
130 | Spinal bolus
131 | Stress Test
132 | Stroke < 101 kg
133 | Stroke 101kg or more
134 | Subsequent Doses
135 | Third Dose
136 | Undergoing PCI
137 | Weekly
138 | Weight < 30 kg
139 | Weight < 50 kg
140 | Weight >= 115 kg
141 | Weight above 110kg
142 | Weight Based
143 | Weight over 80kg
144 | wt < or = 1.5 kg
145 | WT = 10 - 19.9 kg
146 | WT = 20 - 39.9 kg
147 | WT = 40 kg and over
148 | wt 1.5 - 2.5 kg
149 | wt above 2.5 kg
150 | wt more than 100kg

## Channel Labels
Channel Labels | Channel Labels | Channel Labels | Channel Labels | Channel Labels
Line Item | Channel Labels are not used | Channel Labels are not used | Channel Labels are not used | Channel Labels are not used

## Master Syringe List
Master Syringe List | Master Syringe List | Master Syringe List
Line Item | Syringe Type | Included?
1 | BD Plastipak 1 mL | False
2 | BD Plastipak 3 mL | True
3 | BD Plastipak 5 mL | True
4 | BD Plastipak 10 mL | True
5 | BD Plastipak 20 mL | True
6 | BD Plastipak 30 mL | True
7 | BD Plastipak 50 and 60 mL | True
8 | Monoject 3 mL | True
9 | Monoject 6 mL | False
10 | Monoject 12 mL | False
11 | Monoject 20 mL | True
12 | Monoject 35 mL | False
13 | Monoject 60 mL | False
14 | Terumo 3 mL | True
15 | Terumo 5 mL | True
16 | Terumo 10 mL | True
17 | Terumo 20 mL | True
18 | Terumo 30 mL | True
19 | Terumo 60 mL | True

## Master PCA Syringe List
Master PCA Syringe List | Master PCA Syringe List | Master PCA Syringe List
Line Item | PCA Syringe Type | Included?
1 | BD Plastipak 20 mL | True
2 | BD Plastipak 30 mL | True
3 | BD Plastipak 50 and 60 mL | True
4 | Monoject 20 mL | True
5 | Monoject 35 mL | True
6 | Monoject 60 mL | True
7 | Terumo 20 mL | True
8 | Terumo 30 mL | True
9 | Terumo 60 mL | True
10 | IMS 30 mL Prefill | True