## Data Set Properties
Data Set Properties
Data Set Name: | Conway2024v62is
Hospital Name: | Conway Medical Center
Status: | Released
Data Set ID: | 02d39432c-R

## ContBolus Non-Anesth Drugs
Continuous/Bolus Non-Anesthesia Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount (- - = custom concentration) | Drug Amount Units | Diluent Volume (- - = custom concentration) | Diluent Volume Units | Final Concentration | Final Concentration Units | Module Availability | Dosing Units | Dose Limits Soft Min | Dose Limits Soft Max | Dose Limits Hard Max | Dose Initial Value | Bolus Dose Units | Bolus Dose Limits Soft Min | Bolus Dose Limits Soft Max | Bolus Dose Limits Hard Max | Bolus Dose Limits Initial Value | Bolus Dose Admin. Rate Units | Bolus Dose Admin. Rate Limits Soft Min | Bolus Dose Admin. Rate Limits Soft Max | Bolus Dose Admin. Rate Limits Hard Max | Bolus Dose Admin. Rate Initial Value | Conc. Limit Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max | Clinical Advisory
1 | Adult | adenosine (CATH LAB) | .Standard wt < 115kg | 90 | mg | 90 | mL | 1 | mg/mL | Pump | mcg/kg/min | 135 | 145 |  | 140 | [10 empty] | n/a
2 | Adult | adenosine (CATH LAB) | Weight >= 115 kg | 180 | mg | 90 | mL | 2 | mg/mL | Pump | mcg/kg/min | 135 | 145 |  | 140 | [10 empty] | n/a |  |  |  | Double Concentrated
3 | Adult | alteplase -stroke | Stroke < 101 kg | - - | mg | - - | mL |  | mg/mL | Pump | mg/kg/h | 0.7 | 0.9 |  | 0.81 | [10 empty] | mg/mL | 0.85 | 0.99 | 1.01 | Alteplase wt based
4 | Adult | alteplase -stroke | Stroke 101kg or more | 81 | mg | 81 | mL | 1 | mg/mL | Pump | mg/h | 71 | 90 |  | 81 | [10 empty] | n/a
5 | Adult | amiodarone Drip |  | 360 | mg | 200 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.51 | [11 empty] | n/a |  |  |  | In-line filter
6 | Adult | amiodarone Drip |  | 450 | mg | 250 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.51 | [11 empty] | n/a |  |  |  | In-line filter
7 | Adult | amiodarone Drip |  | 900 | mg | 500 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.51 | [11 empty] | n/a |  |  |  | In-line filter
8 | Adult | argatroban | .Standard Dose | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 10 | 11.5 | 2 | [10 empty] | n/a
9 | Adult | argatroban | .Standard Dose | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 10 | 11.5 | 2 | [10 empty] | n/a
10 | Adult | argatroban | Hepatic Dose | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.1 | 5 | 10.1 | 0.5 | [10 empty] | n/a
11 | Adult | argatroban | Hepatic Dose | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.1 | 5 | 10.1 | 0.5 | [10 empty] | n/a
12 | Adult | argatroban | Undergoing PCI | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 15 | 30 | 34 | 25 | mcg/kg | 149 | 351 |  |  | mcg/kg/min | 29 | 117 |  |  | n/a
13 | Adult | argatroban | Undergoing PCI | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 15 | 30 | 34 | 25 | mcg/kg | 149 | 351 |  |  | mcg/kg/min | 29 | 117 |  |  | n/a
14 | Adult | atracurium |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 0.5 | 29.5 | 39 |  | mg/kg | 0.07 | 0.5 | 0.58 |  | mg/kg/min | 0.07 | 0.25 | 0.29 |  | mcg/mL | 200 | 5000 | 5100 | Neuromuscular BA
15 | Adult | bivalirudin | .Standard Dose | 250 | mg | 50 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.1 | 1.8 | 2 | 1.75 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.76 |  | 0.75 | n/a
16 | Adult | bivalirudin | .Standard Dose | 500 | mg | 100 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.1 | 1.8 | 2 | 1.75 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.76 |  | 0.75 | n/a
17 | Adult | bivalirudin | .Standard Dose | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.1 | 1.8 | 2 | 1.75 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.76 |  | 0.75 | n/a
18 | Adult | bivalirudin | ESRD | 250 | mg | 50 | mL | 5 | mg/mL | Pump | mcg/kg/h | 0.1 | 1.2 | 1.4 | 0.25 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
19 | Adult | bivalirudin | ESRD | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/h | 0.1 | 1.2 | 1.4 | 0.25 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
20 | Adult | bivalirudin | Renal Dose CrCl < 30 | 250 | mg | 50 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.24 | 1.2 | 1.4 | 1 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
21 | Adult | bivalirudin | Renal Dose CrCl < 30 | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.24 | 1.2 | 1.4 | 1 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
22 | Adult | bivalirudin | Renal Dose CrCl < 30 | 500 | mg | 100 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.24 | 1.2 | 1.4 | 1 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
23 | Adult | bumetanide continuos |  | 24 | mg | 96 | mL | 0.25 | mg/mL | Pump | mg/h | 0.1 | 2 |  | 0.5 | mg | 0.25 | 4 |  |  | mg/min | 0.25 |  | 0.5 |  | n/a |  |  |  | Protect From Light
24 | Adult | cangrelor |  | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 2 | 4.2 | 4.6 | 4 | mcg/kg | 26 |  | 35 | 30 | mcg/kg/min | 26 |  | 35 |  | n/a
25 | Adult | cisatracurium |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | mcg/kg/min | 0.5 | 5 | 7 | 3 | [10 empty] | n/a |  |  |  | Neuromuscular BA
26 | Adult | conivaptan |  | 20 | mg | 100 | mL | 0.2 | mg/mL | Pump | mg/h | 0.6 | 1.7 | 2 | 0.8 | mg | 10 | 20 |  | 20 | mg/min | 0.333 | 0.7 |  | 0.667 | n/a
27 | Adult | conivaptan |  | 20 | mg | 250 | mL | 0.08 | mg/mL | Pump | mg/h | 0.6 | 1.7 | 2 | 0.8 | mg | 10 | 20 |  | 20 | mg/min | 0.333 | 0.7 |  | 0.667 | n/a
28 | Adult | dexmedetomidine |  | 200 | mcg | 50 | mL | 4 | mcg/mL | Pump | mcg/kg/h | 0.05 | 0.7 | 1.5 | 0.2 | mcg/kg | 0.9 | 1.1 | 2 | 1 | mcg/kg/min | 0.01 | 0.11 | 0.2 | 0.1 | n/a
29 | Adult | dexmedetomidine |  | 400 | mcg | 100 | mL | 4 | mcg/mL | Pump | mcg/kg/h | 0.05 | 0.7 | 1.5 | 0.2 | mcg/kg | 0.9 | 1.1 | 2 | 1 | mcg/kg/min | 0.01 | 0.11 | 0.2 | 0.1 | n/a
30 | Adult | dilTIAZem |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | mg/h | 0.5 | 15 | 25 | 5 | [10 empty] | n/a
31 | Adult | dilTIAZem |  | 125 | mg | 125 | mL | 1 | mg/mL | Pump | mg/h | 0.5 | 15 | 25 | 5 | [10 empty] | n/a
32 | Adult | DOBUTamine | ..Standard | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a
33 | Adult | DOBUTamine | ..Standard | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a
34 | Adult | DOBUTamine | .Double Concentrated | 500 | mg | 250 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a |  |  |  | Double Concentrated
35 | Adult | DOBUTamine | .Double Concentrated | 1000 | mg | 500 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a |  |  |  | Double Concentrated
36 | Adult | DOBUTamine | Echo Stress Test | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 4.9 | 50 |  | 5 | [10 empty] | mcg/mL | 900 |  | 5000
37 | Adult | DOBUTamine | Stress Test | 500 | mg | 250 | mL | 2 | mg/mL | Pump | mcg/kg/min | 5 | 50 |  | 10 | [10 empty] | n/a
38 | Adult | DOPamine | ..Standard | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 0.9 | 20 | 30 | [11 empty] | n/a
39 | Adult | DOPamine | .Double Concentrated | 800 | mg | 250 | mL | 3.2 | mg/mL | Pump | mcg/kg/min | 0.9 | 20 | 30 | [11 empty] | n/a |  |  |  | Double Concentrated
40 | Adult | DOPamine | Renal Dose | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 0.9 | 5 | 6 | 1 | [10 empty] | n/a
41 | Adult | doxapram |  | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mg/min | 0.1 | 6 | [12 empty] | n/a
42 | Adult | doxapram |  | - - | mg | - - | mL |  | mg/mL | Pump | mg/min | 0.1 | 6 | [12 empty] | mg/mL | 0.89 | 0.9 | 2.1
43 | Adult | epidural w/ fentanyl |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 0.01 | 12 | 15 |  | unit | 2 | 7 | 8 | 5 | unit/min | 4 | 14 | 16 | 10 | n/a |  |  |  | 1 unit = 1 mL
44 | Adult | epidural w/o fentany |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 0.01 | 12 | 15 |  | unit | 2 | 7 | 8 | 5 | unit/min | 4 | 14 | 16 | 10 | n/a |  |  |  | 1 unit = 1 mL
45 | Adult | EPINEPHrine | ..Standard | 5 | mg | 250 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 40 | [12 empty] | n/a |  |  |  | Dedicated IV Line
46 | Adult | EPINEPHrine | ..Standard | 10 | mg | 500 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 40 | [12 empty] | n/a |  |  |  | Dedicated IV Line
47 | Adult | EPINEPHrine | Cardiac Arrest | 30 | mg | 250 | mL | 0.12 | mg/mL | Pump | mcg/min | 0.33 | 40 |  | 33.33 | [10 empty] | n/a |  |  |  | Dedicated IV Line
48 | Adult | epoprostenol |  | 1.5 | mg | 50 | mL | 0.03 | mg/mL | Syringe | nanogram/kg/min | 5 | 51 |  | 50 | [10 empty] | n/a
49 | Adult | eptifibatide | .Standard & wt>121kg | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mg/h | 13 | 17 |  | 15 | [10 empty] | n/a |  |  |  | SCr<2.0
50 | Adult | eptifibatide | .Standard & wt>121kg | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mg/h | 13 | 17 |  | 15 | [10 empty] | n/a |  |  |  | SCr<2.0
51 | Adult | eptifibatide | .Standard Dose | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mcg/kg/min | 0.5 | 2.2 | 2.5 | 2 | [10 empty] | n/a |  |  |  | SCr<2.0
52 | Adult | eptifibatide | .Standard Dose | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 2.2 | 2.5 | 2 | [10 empty] | n/a |  |  |  | SCr<2.0
53 | Adult | eptifibatide | Renal Dose | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mcg/kg/min | 0.25 | 1.2 | 1.4 | 1 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
54 | Adult | eptifibatide | Renal Dose | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.25 | 1.2 | 1.4 | 1 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
55 | Adult | eptifibatide | Renal Dose wt >121kg | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mg/h | 6 | 8 | 9 | 7.5 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
56 | Adult | eptifibatide | Renal Dose wt >121kg | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mg/h | 6 | 8 | 9 | 7.5 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
57 | Adult | esmolol |  | 2500 | mg | 250 | mL | 10 | mg/mL | Pump | mcg/kg/min | 25 | 300 |  | 50 | mcg/kg | 100 | 500 |  |  | mcg/kg/min | 100 | 500 |  |  | n/a
58 | Adult | fenoldopam |  | 10 | mg | 250 | mL | 0.04 | mg/mL | Pump | mcg/kg/min | 0.03 | 1.6 | 1.8 | 0.1 | [10 empty] | n/a
59 | Adult | fentaNYL |  | 2500 | mcg | 100 | mL | 25 | mcg/mL | Pump | mcg/h | 0.0001 | 250 | 750 |  | mcg | 0.0001 | 250 | 500 |  | mcg/min | 0.0001 | 250 | 500 |  | n/a
60 | Adult | furosemide |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mg/h | 1 | 40 | [12 empty] | n/a |  |  |  | Protect From Light
61 | Adult | glucagon |  | 10 | mg | 50 | mL | 0.2 | mg/mL | Pump | mg/h | 0.1 | 10 | [12 empty] | n/a
62 | Adult | glucagon |  | 10 | mg | 100 | mL | 0.1 | mg/mL | Pump | mg/h | 0.1 | 10 | [12 empty] | n/a
63 | Adult | heparin | ..Std. Cardiac/PVD | 25000 | unit | 500 | mL | 50 | unit/mL | Pump | unit/kg/h | 1 | 100 |  | 12 | unit/kg | 24 |  | 62 |  | unit/kg/min | 12 |  | 31 |  | n/a |  |  |  | Heparin
64 | Adult | heparin | .VTE (DVT, PE) | 25000 | unit | 500 | mL | 50 | unit/mL | Pump | unit/kg/h | 1 | 100 |  | 18 | unit/kg | 24 |  | 82 |  | unit/kg/min | 12 |  | 41 |  | n/a |  |  |  | Heparin
65 | Adult | heparin | Cerebrovascular | 25000 | unit | 500 | mL | 50 | unit/mL | Pump | unit/kg/h | 1 | 100 |  | 10 | [10 empty] | n/a |  |  |  | Heparin
66 | Adult | insulin |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 1 | 20 | 30 |  | unit | 1 | 11 |  |  | unit/min | 0.5 | 15 |  | 10 | n/a |  |  |  | 2nd RN to verify
67 | Adult | insulin |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/kg/h | 0.001 | 0.3 | 0.8 |  | unit/kg | 0.001 | 0.3 |  | 0.1 | unit/min | 0.5 | 15 |  | 10 | n/a |  |  |  | 2nd RN to verify
68 | Adult | isoproterenol |  | 1 | mg | 250 | mL | 0.004 | mg/mL | Pump | mcg/min | 0.15 | 4 | 30 | 1 | [10 empty] | n/a
69 | Adult | isoproterenol |  | 2 | mg | 500 | mL | 0.004 | mg/mL | Pump | mcg/min | 0.15 | 4 | 30 | 1 | [10 empty] | n/a
70 | Adult | ketamine ANES |  | - - | mg | - - | mL |  | mg/mL | Pump | mg/kg/h | 0.05 | 0.5 | 1 | 0.1 | [10 empty] | mg/mL | 0.1 | 0.5 | 1.5
71 | Adult | labetalol |  | 200 | mg | 200 | mL | 1 | mg/mL | Pump | mg/min | 0.05 | 2 |  | 2 | [10 empty] | n/a
72 | Adult | lidocaine |  | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | mg/min | 0.5 | 4 | 4.6 | 2 | [10 empty] | n/a
73 | Adult | mag OB maintenance |  | 20 | gram | 500 | mL | 0.04 | gram/mL | Pump | gram/h | 0.5 | 3 | 5 | [11 empty] | n/a
74 | Adult | midazolam | Anti-Convulsant | 100 | mg | 100 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.0001 | 2 |  | 0.1 | mg/kg | 0.0001 | 0.21 |  | 0.2 | mg/min | 1.9 | 2.1 |  | 2 | n/a
75 | Adult | midazolam | Anti-Convulsant | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.0001 | 2 |  | 0.1 | mg/kg | 0.0001 | 0.21 |  | 0.2 | mg/min | 1.9 | 2.1 |  | 2 | n/a
76 | Adult | midazolam | Sedation | 100 | mg | 100 | mL | 1 | mg/mL | Pump | mg/h | 0.0001 | 5 | 10 |  | mg | 0.0001 | 1 | 5 |  | mg/min | 0.0001 | 1 | 5 |  | n/a
77 | Adult | milrinone |  | 40 | mg | 200 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 0.19 | 0.75 |  | 0.5 | mcg/kg | 25 | 50.1 |  | 50 | mcg/kg/min | 1 | 5.01 |  | 5 | n/a
78 | Adult | naloxone |  | 4 | mg | 100 | mL | 0.04 | mg/mL | Pump | mg/h | 0.25 | 2 | [12 empty] | n/a
79 | Adult | nesiritide |  | 1.5 | mg | 250 | mL | 0.006 | mg/mL | Pump | mcg/kg/min | 0.005 | 0.03 | 0.035 | 0.01 | [10 empty] | n/a
80 | Adult | niCARdipine |  | 20 | mg | 200 | mL | 0.1 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
81 | Adult | niCARdipine |  | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
82 | Adult | niCARdipine |  | 40 | mg | 200 | mL | 0.2 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
83 | Adult | niCARdipine |  | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
84 | Adult | niCARdipine |  | 50 | mg | 500 | mL | 0.1 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
85 | Adult | nitroglycerin | ..Standard | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mcg/min | 5 | 200 | 400 | 5 | [10 empty] | n/a
86 | Adult | nitroglycerin | .Double Concentrated | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mcg/min | 5 | 200 | 400 | 5 | [10 empty] | n/a
87 | Adult | nitroprusside |  | 50 | mg | 500 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.1 |  | 10 | 0.3 | [10 empty] | n/a |  |  |  | Protect From Light
88 | Adult | norepinephrine | ..Standard | 4 | mg | 250 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Levophed
89 | Adult | norepinephrine | ..Standard | 8 | mg | 500 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Levophed
90 | Adult | norepinephrine | .Double Concentrated | 16 | mg | 500 | mL | 0.032 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Double Concentrated
91 | Adult | octreotide |  | 600 | mcg | 240 | mL | 2.5 | mcg/mL | Pump | mcg/h | 25 | 50 |  |  | mcg | 25 | 150 |  | 50 | mcg/min | 1 | 50 |  | 10 | n/a
92 | Adult | oxytocin |  | 30 | unit | 500 | mL | 0.06 | unit/mL | Pump | milliunit/min | 1 | 41 |  |  | unit | 0.9 | 30.1 |  | 10 | unit/min | 0.1 | 1.01 |  | 0.999 | n/a
93 | Adult | pancuronium |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 0.01 | 2 | 3 |  | mcg/kg | 0.06 | 0.1 |  |  | mcg/kg/min | 0.06 | 1 |  |  | mcg/mL | 99 | 100 | 2001 | Neuromuscular BA
94 | Adult | pantoprazole drip |  | 40 | mg | 100 | mL | 0.4 | mg/mL | Pump | mg/h | 7 | 9 | 11 | 8 | [10 empty] | n/a
95 | Adult | pantoprazole drip |  | 96 | mg | 240 | mL | 0.4 | mg/mL | Pump | mg/h | 7 | 9 | 11 | 8 | [10 empty] | n/a
96 | Adult | phenylephrine |  | 10 | mg | 250 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Phenylephrine
97 | Adult | phenylephrine |  | 20 | mg | 500 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Phenylephrine
98 | Adult | phenylephrine |  | 40 | mg | 1000 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Phenylephrine
99 | Adult | phenylephrine ANES |  | 10 | mg | 101 | mL | 0.099 | mg/mL | Pump | mcg/min | 5 | 120 | 180 | [11 empty] | n/a
100 | Adult | phenylephrine ANES |  | 10 | mg | 251 | mL | 0.0398 | mg/mL | Pump | mcg/min | 5 | 120 | 180 | [11 empty] | n/a
101 | Adult | procainamide | Loading Dose | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | mg/min | 15 | 50 |  | 33.33 | [10 empty] | n/a
102 | Adult | procainamide | Maintenance Dose | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | mg/kg/min | 0.01 | 0.06 |  | 0.05 | [10 empty] | n/a
103 | Adult | propofol |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | mcg/kg/min | 1 | 55 | 100 | 5 | [10 empty] | n/a
104 | Adult | rocuronium |  | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 4 | 12 | 16 | 8 | [10 empty] | n/a |  |  |  | Neuromuscular BA
105 | Adult | tirofiban | CrCl <= 60 | 12.5 | mg | 250 | mL | 0.05 | mg/mL | Pump | mcg/kg/min | 0.065 |  | 0.085 | 0.075 | mcg/kg | 22 |  | 28 | 25 | mcg/kg/min | 4 |  | 6 | 5 | n/a |  |  |  | Bleeding Precautions
106 | Adult | tirofiban | CrCl > 60 | 12.5 | mg | 250 | mL | 0.05 | mg/mL | Pump | mcg/kg/min | 0.13 |  | 0.17 | 0.15 | mcg/kg | 22 |  | 28 | 25 | mcg/kg/min | 4 |  | 6 | 5 | n/a |  |  |  | Bleeding Precautions
107 | Adult | vasopressin | ..Standard | 40 | unit | 100 | mL | 0.4 | unit/mL | Pump | unit/min | 0.01 | 0.05 | 0.06 | 0.04 | [10 empty] | n/a
108 | Adult | vasopressin | .Double Concentrated | 80 | unit | 100 | mL | 0.8 | unit/mL | Pump | unit/min | 0.01 | 0.05 | 0.06 | 0.04 | [10 empty] | n/a |  |  |  | Double Concentrated
109 | Adult | vecuronium |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | mcg/kg/min | 0.01 | 1.3 | 1.67 | 1 | mg | 5 | 10 |  |  | mg/min | 5 | 20 |  |  | n/a |  |  |  | Neuromuscular BA
110 | Adult | zz amiodarone DRIP |  | 360 | mg | 200 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.41 | [11 empty] | n/a |  |  |  | Code only drugs
111 | Adult | zz amiodarone DRIP |  | 900 | mg | 500 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.41 | [11 empty] | n/a |  |  |  | Code only drugs
112 | Adult | zz DOBUTamine | ..Standard | 500 | mg | 250 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a |  |  |  | Code only drugs
113 | Adult | zz DOBUTamine | .Double Concentrated | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | [11 empty] | n/a |  |  |  | Code only drugs
114 | Adult | zz DOPamine | ..Standard | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 0.9 | 20 | 30 | [11 empty] | n/a |  |  |  | Code only drugs
115 | Adult | zz DOPamine | .Double Concentrated | 800 | mg | 250 | mL | 3.2 | mg/mL | Pump | mcg/kg/min | 0.9 | 20 | 30 | [11 empty] | n/a |  |  |  | Code only drugs
116 | Adult | zz EPINEPHrine |  | 5 | mg | 250 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 40 | [12 empty] | n/a |  |  |  | Code only drugs
117 | Adult | zz EPINEPHrine |  | 10 | mg | 500 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 40 | [12 empty] | n/a |  |  |  | Code only drugs
118 | Adult | zz lidocaine |  | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | mg/min | 0.5 | 4 | 4.6 | 2 | [10 empty] | n/a |  |  |  | Code only drugs
119 | Adult | zz nitroglycerin | ..Standard | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mcg/min | 5 | 200 | 400 | 5 | [10 empty] | n/a |  |  |  | Code only drugs
120 | Adult | zz nitroglycerin | .Double Concentrated | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mcg/min | 5 | 200 | 400 | 5 | [10 empty] | n/a
121 | Adult | zz nitroprusside |  | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 0.1 |  | 10 | 0.3 | [10 empty] | n/a |  |  |  | Code only drugs
122 | Adult | zz norepinephrine | ..Standard | 4 | mg | 250 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Code only drugs
123 | Adult | zz norepinephrine | ..Standard | 8 | mg | 500 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Code only drugs
124 | Adult | zz norepinephrine | .Double Concentrated | 16 | mg | 500 | mL | 0.032 | mg/mL | Pump | mcg/min | 0.3 | 80 | [12 empty] | n/a |  |  |  | Code only drugs
125 | Adult | zz phenylephrine |  | 10 | mg | 250 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Code only drugs
126 | Adult | zz phenylephrine |  | 20 | mg | 500 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Code only drugs
127 | Adult | zz phenylephrine |  | 40 | mg | 1000 | mL | 0.04 | mg/mL | Pump | mcg/min | 5 | 400 |  | 20 | [10 empty] | n/a |  |  |  | Code only drugs
128 | Nursery | No drugs exist in this section
129 | ONC Infusion Center | No drugs exist in this section
130 | OutPt InfusionCenter | No drugs exist in this section
131 | Pediatrics | alprostadil |  | 500 | mcg | 250 | mL | 2 | mcg/mL | Pump | mcg/kg/min | 0.05 | 0.1 | 0.4 | 0.05 | [10 empty] | n/a
132 | Pediatrics | alprostadil |  | - - | mcg | - - | mL |  | mcg/mL | Pump | mcg/kg/min | 0.05 | 0.1 | 0.4 | 0.05 | [10 empty] | mcg/mL | 0.5 | 1 | 20
133 | Pediatrics | aminophylline |  | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.2 | 1.1 |  |  | mg/kg | 0.001 |  | 6 |  | mg/kg/min | 0.19 |  | 0.36 | 0.2 | n/a
134 | Pediatrics | DOBUTamine |  | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | 2 | [10 empty] | n/a
135 | Pediatrics | DOPamine |  | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 2 |  | 20 | 2 | [10 empty] | n/a
136 | Pediatrics | EPINEPHrine |  | 5 | mg | 250 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 20 | 30 | [11 empty] | n/a
137 | Pediatrics | insulin |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 1 | 20 | 30 |  | unit | 1 | 11 |  |  | unit/min | 0.5 | 15 |  | 10 | n/a |  |  |  | 2nd RN to verify
138 | Pediatrics | insulin |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/kg/h | 0.001 | 0.3 | 0.8 |  | unit/kg | 0.001 | 0.3 |  | 0.1 | unit/min | 0.5 | 15 |  | 10 | n/a |  |  |  | 2nd RN to verify
139 | Pediatrics | lidocaine |  | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | mg/min | 0.5 | 4 | 4.6 | 2 | [10 empty] | n/a
140 | Pediatrics | milrinone |  | 20 | mg | 100 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 0.125 | 0.75 |  | 0.5 | mcg/kg | 25 | 50 |  |  | mcg/kg/min | 1 | 5 |  |  | n/a
141 | Pediatrics | naloxone |  | 2 | mg | 500 | mL | 0.004 | mg/mL | Pump | mg/h | 0.25 | 2 | [12 empty] | n/a
142 | Pediatrics | niCARdipine |  | 20 | mg | 200 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.01 | 4 | 5 | 0.5 | mcg/kg | 29 |  | 50 | 30 | mcg/kg/min | 29 |  | 50 | 30 | n/a
143 | Pediatrics | niCARdipine |  | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.01 | 4 | 5 | 0.5 | mcg/kg | 29 |  | 50 | 30 | mcg/kg/min | 29 |  | 50 | 30 | n/a
144 | Pediatrics | niCARdipine |  | 40 | mg | 200 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 0.01 | 4 | 5 | 0.5 | mcg/kg | 29 |  | 50 | 30 | mcg/kg/min | 29 |  | 50 | 30 | n/a
145 | Pediatrics | niCARdipine |  | 50 | mg | 500 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.01 | 4 | 5 | 0.5 | mcg/kg | 29 |  | 50 | 30 | mcg/kg/min | 29 |  | 50 | 30 | n/a
146 | Pediatrics | nitroglycerin |  | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mcg/min | 5 | 150 | 400 | 5 | [10 empty] | n/a
147 | Pediatrics | nitroprusside |  | 50 | mg | 500 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.1 | 8 | 8.1 | 1 | [10 empty] | n/a |  |  |  | Protect From Light
148 | Pediatrics | norepinephrine |  | 4 | mg | 250 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 45 | [12 empty] | n/a
149 | Pediatrics | norepinephrine |  | 8 | mg | 500 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 45 | [12 empty] | n/a
150 | Pediatrics | propofol |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | mcg/kg/min | 1 | 155 | 310 | [11 empty] | n/a
Note on AE4: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on AE5: Dose and volume to be infused need to be set according to weight based protocol to avoid potential over infusion of entire contents of bag.
Note on AE7: Infusion requires the use of an in-line filter.
Note on AE8: Infusion requires the use of an in-line filter.
Note on AE9: Infusion requires the use of an in-line filter.
Note on J16: Units Only concentration and Dosing Units have different units of measure
Note on AA16: Units Only concentration and Dosing Units have different units of measure
Note on AE16: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE25: Protect this drug from light.
Note on AE27: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE36: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on AE37: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on J38: Units Only concentration and Dosing Units have different units of measure
Note on AA38: Units Only concentration and Dosing Units have different units of measure
Note on AE41: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on AE45: 1 unit = 1 mL
Note on AE46: 1 unit = 1 mL
Note on AE47: Do not infuse any other drugs through this line.
Note on AE48: Do not infuse any other drugs through this line.
Note on AE49: Do not infuse any other drugs through this line.
Note on AE51: SCr < 2.0
Note on AE52: SCr < 2.0
Note on AE53: SCr < 2.0
Note on AE54: SCr < 2.0
Note on AE55: SCr 2.0 - 4.0
Note on AE56: SCr 2.0 - 4.0
Note on AE57: SCr 2.0 - 4.0
Note on AE58: SCr 2.0 - 4.0
Note on AE62: Protect this drug from light.
Note on AE65: Monitor PTT. Do not administer other drugs through same IV line.
Note on AE66: Monitor PTT. Do not administer other drugs through same IV line.
Note on AE67: Monitor PTT. Do not administer other drugs through same IV line.
Note on AE68: 2 RNs to verify and document concentration and programming prior to infusion.
Note on AE69: 2 RNs to verify and document concentration and programming prior to infusion.
Note on AE89: Protect this drug from light.
Note on AE90: *****Must have physician order for rates greater than 60 mcg/min*****
Note on AE91: *****Must have physician order for rates greater than 60 mcg/min*****
Note on AE92: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on J95: Units Only concentration and Dosing Units have different units of measure
Note on AA95: Units Only concentration and Dosing Units have different units of measure
Note on AE95: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE98: *****Must have physician order for rates greater than 300 mcg/min*****
Note on AE99: *****Must have physician order for rates greater than 300 mcg/min*****
Note on AE100: *****Must have physician order for rates greater than 300 mcg/min*****
Note on AE106: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE107: Monitor for signs of bleeding.
Note on AE108: Monitor for signs of bleeding.
Note on AE110: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on AE111: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE112: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE113: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE114: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE115: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE116: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE117: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE118: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE119: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE120: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE121: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE123: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE124: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE125: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE126: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE127: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE128: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE129: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE139: 2 RNs to verify and document concentration and programming prior to infusion.
Note on AE140: 2 RNs to verify and document concentration and programming prior to infusion.
Note on AE149: Protect this drug from light.

## ContBolus Anesth Drugs
Continuous/Bolus Anesthesia Only Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount (- - = custom concentration) | Drug Amount Units | Diluent Volume (- - = custom concentration) | Diluent Volume Units | Final Concentration | Final Concentration Units | Module Availability | Dosing Units | Dose Limits Soft Min | Dose Limits Soft Max | Dose Limits Hard Max | Dose Initial Value | Bolus Dose Units | Bolus Dose Limits Soft Min | Bolus Dose Limits Soft Max | Bolus Dose Limits Hard Max | Bolus Dose Limits Initial Value | Bolus Dose Admin. Rate Units | Bolus Dose Admin. Rate Limits Soft Min | Bolus Dose Admin. Rate Limits Soft Max | Bolus Dose Admin. Rate Limits Hard Max | Bolus Dose Admin. Rate Initial Value | Conc. Limit Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max
1 | Adult | ketamine ANES |  | 50 | mg | 50 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.05 | 0.5 | 1 | 0.1 | [10 empty] | n/a
2 | Adult | ketamine ANES |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.05 | 0.5 | 1 | 0.1 | [10 empty] | n/a
3 | Adult | phenylephrine ANES |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/min | 5 | 120 | 180 | [11 empty] | mcg/mL | 10 | 38 | 110
4 | Adult | propofol ANES |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | mcg/kg/min | 1 |  | 300 | 5 | mg | 10 | 70 |  | 30 | mg/min | 50 | 140 |  | 120 | n/a
5 | Adult | propofol ANES |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | mcg/kg/min | 1 |  | 300 | 5 | mg | 10 | 70 |  | 30 | mg/min | 50 | 140 |  | 120 | n/a
6 | Adult | remifentanil ANES |  | 1000 | mcg | 40 | mL | 25 | mcg/mL | Pump | mcg/kg/min | 0.025 | 2 | 4 | 0.05 | mcg/kg | 0.1 |  | 2 | 1 | mcg/min | 1 | 2000 | 4000 | 1000 | n/a
7 | Adult | remifentanil ANES |  | 2000 | mcg | 80 | mL | 25 | mcg/mL | Pump | mcg/kg/min | 0.025 | 2 | 4 | 0.05 | mcg/kg | 0.1 |  | 2 | 1 | mcg/min | 1 | 2000 | 4000 | 1000 | n/a
8 | Adult | remifentanil ANES |  | - - | mcg | - - | mL |  | mcg/mL | Pump | mcg/kg/min | 0.025 | 2 | 4 | 0.05 | mcg/kg | 0.1 |  | 2 | 1 | mcg/min | 1 | 2000 | 4000 | 1000 | mcg/mL | 8.9 | 9 | 250
9 | Adult | rocuronium ANES |  | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 4 | 12 | 16 | 8 | [10 empty] | n/a
10 | Adult | rocuronium ANES |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 4 | 12 | 16 | 8 | [10 empty] | mcg/mL | 890 | 900 | 1100
11 | Adult | SUFentanil ANES |  | 100 | mcg | 100 | mL | 1 | mcg/mL | Pump | mcg/kg/h | 0.1 | 1 | [12 empty] | n/a
12 | Adult | SUFentanil ANES |  | - - | mcg | - - | mL |  | mcg/mL | Pump | mcg/kg/h | 0.1 | 1 | [12 empty] | mcg/mL | 0.1 | 0.2 | 4
13 | Adult | vecuronium |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 0.01 | 1.3 | 1.67 | 1 | mg | 5 | 10 |  |  | mg/min | 5 | 20 |  |  | mcg/mL | 440 | 450 | 550
14 | Nursery | No drugs exist in this section
15 | ONC Infusion Center | No drugs exist in this section
16 | OutPt InfusionCenter | No drugs exist in this section
17 | Pediatrics | No drugs exist in this section
Note on J5: Units Only concentration and Dosing Units have different units of measure
Note on AA5: Units Only concentration and Dosing Units have different units of measure
Note on J12: Units Only concentration and Dosing Units have different units of measure
Note on AA12: Units Only concentration and Dosing Units have different units of measure
Note on J15: Units Only concentration and Dosing Units have different units of measure
Note on AA15: Units Only concentration and Dosing Units have different units of measure

## Intermittent Drugs
Intermittent Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount (- - = custom concentration) | Drug Amount Units | Diluent Volume (- - = custom concentration) | Diluent Volume Units | Final Concentration | Final Concentration Units | Module Availability | Available As | Duration Limits (hh:mm) Hard Min | Duration Limits (hh:mm) Soft Min | Duration Limits (hh:mm) Soft Max | Duration (hh:mm) Initial Value | Dosing Units | Dose Limits Soft Min | Dose Limits Soft Max | Dose Limits Hard Max | Conc. Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max | Clinical Advisory
1 | Adult | acetaminophen | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 500 | 1000 |  | n/a
2 | Adult | acetaminophen | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | gram | 0.5 | 1 |  | n/a
3 | Adult | acetaZOLAMIDE |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 00:40 | 00:30 | mg | 100 | 1000 |  | n/a
4 | Adult | acetylcysteine | 1st Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg/kg | 142.5 |  | 157.5 | mg/mL | 1 |  | 75
5 | Adult | acetylcysteine | 1st Dose >100kg | 15000 | mg | 275 | mL | 54.5455 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 13500 | 16500 | 18000 | n/a
6 | Adult | acetylcysteine | 2nd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg/kg | 47.5 |  | 52.5 | mg/mL | 1 |  | 10
7 | Adult | acetylcysteine | 2nd Dose >100kg | 5000 | mg | 525 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg | 4000 | 5500 | 6500 | n/a
8 | Adult | acetylcysteine | 3rd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg/kg | 95 |  | 105 | mg/mL | 1 |  | 10
9 | Adult | acetylcysteine | 3rd Dose >100kg | 10000 | mg | 1050 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg | 9000 | 11000 | 12500 | n/a
10 | Adult | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 100 | 2000 |  | mg/mL | 0.1 | 1 | 7
11 | Adult | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg/kg | 1 | 10.1 |  | mg/mL | 0.1 | 1 | 7
12 | Adult | albumin 25% |  | 5 | gram | 20 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:08 | 01:00 | 00:20 | gram | 4.9 | 25.1 |  | n/a
13 | Adult | albumin 25% |  | 12.5 | gram | 50 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:16 | 02:00 | 00:50 | gram | 4.9 | 25.1 |  | n/a
14 | Adult | albumin 25% |  | 25 | gram | 100 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:32 | 04:00 | 01:40 | gram | 4.9 | 25.1 |  | n/a
15 | Adult | alteplase-non-stroke | Central IV Occlusion | 2 | mg | 50 | mL | 0.04 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:00 | 00:15 | mg | 2 | 4 |  | n/a
16 | Adult | alteplase-non-stroke | Central IV Occlusion | 2 | mg | 100 | mL | 0.02 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:00 | 00:15 | mg | 2 | 4 |  | n/a
17 | Adult | alteplase-non-stroke | DVT / Ischemic Limb | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:00 |  | 24:00 |  | mg | 0.9 |  | 24 | mg/mL | 0.004 | 0.008 | 0.51
18 | Adult | alteplase-non-stroke | Pulmonary Embolism | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary Only |  | 01:45 | 02:15 | 02:00 | mg | 49 | 115 | 130 | n/a
19 | Adult | amikacin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 |  | mg/kg | 1 | 21 | 39 | mg/mL | 0.2 |  | 5.1
20 | Adult | aminocaproic acid | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 500 | 5000 |  | mg/mL | 1 | 5 | 50
21 | Adult | aminocaproic acid | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | gram | 0.5 | 5.7 |  | gram/mL | 0.001 | 0.005 | 0.05
22 | Adult | amiodarone *Loading* |  | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:08 | 00:15 | 00:10 | mg | 149 | 151 |  | n/a
23 | Adult | amiodarone *Loading* |  | 150 | mg | 103 | mL | 1.4563 | mg/mL | Pump | Primary and Secondary |  | 00:08 | 00:15 | 00:10 | mg | 149 | 151 |  | n/a
24 | Adult | ampho B (liposomal) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only |  | 00:55 | 04:00 | 02:00 | mg/kg | 1.9 | 10.5 |  | mg/mL | 0.8 | 0.9 | 2.1 | AmBisome
25 | Adult | amphotericin B |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 02:00 | 03:00 | 10:00 | 06:00 | mg/kg | 0.05 | 1 | 1.5 | mg/mL | 0.01 |  | 0.1
26 | Adult | ampicillin | .MG Dosing | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
27 | Adult | ampicillin | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
28 | Adult | ampicillin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
29 | Adult | ampicillin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
30 | Adult | ampicillin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
31 | Adult | ampicillin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2100 | 2400 | n/a
32 | Adult | ampicillin | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.1 | 2.4 | n/a
33 | Adult | ampicillin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.1 | 2.4 | n/a
34 | Adult | ampicillin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.1 | 2.4 | n/a
35 | Adult | ampicillin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.1 | 2.4 | n/a
36 | Adult | ampicillin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.1 | 2.4 | n/a
37 | Adult | ampicillin/sulbact | .MG Dosing | 1500 | mg | 50 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | mg | 1200 | 3400 |  | n/a
38 | Adult | ampicillin/sulbact | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | mg | 1200 | 3400 |  | n/a
39 | Adult | ampicillin/sulbact | .MG Dosing | 3000 | mg | 100 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 04:00 | 01:00 | mg | 1200 | 3400 |  | n/a
40 | Adult | ampicillin/sulbact | GM dosing | 1.5 | gram | 50 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | gram | 1.2 | 3.4 |  | n/a
41 | Adult | ampicillin/sulbact | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | gram | 1.2 | 3.4 |  | n/a
42 | Adult | ampicillin/sulbact | GM dosing | 3 | gram | 100 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 04:00 | 01:00 | gram | 1.2 | 3.4 |  | n/a
43 | Adult | andulafungin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 04:00 | 02:00 | mg | 50 | 200 |  | mg/mL | 0.1 | 0.35 | 0.5
44 | Adult | antivenin crotalidae | Anavip | 4 | unit | 250 | mL | 0.016 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 10.1 | 11 | n/a |  |  |  | Antivenin
45 | Adult | antivenin crotalidae | Anavip | 10 | unit | 250 | mL | 0.04 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 10.1 | 11 | n/a |  |  |  | Antivenin
46 | Adult | antivenin crotalidae | Anavip | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 10.1 | 11 | unit/mL | 0.003 |  | 0.045 | Antivenin
47 | Adult | antivenin crotalidae | Crofab | 4 | unit | 250 | mL | 0.016 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 6.1 | 13 | n/a |  |  |  | Antivenin
48 | Adult | antivenin crotalidae | Crofab | 6 | unit | 250 | mL | 0.024 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 6.1 | 13 | n/a |  |  |  | Antivenin
49 | Adult | antivenin crotalidae | Crofab | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 6.1 | 13 | unit/mL | 0.004 |  | 0.025 | Antivenin
50 | Adult | avibactam/ceftazidim |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 01:42 | 03:20 | 02:00 | gram | 0.93 |  | 2.8 | gram/mL | 0.009 |  | 0.051
51 | Adult | azithromycin |  | 250 | mg | 250 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:10 | 01:00 | mg | 210 | 575 |  | n/a
52 | Adult | azithromycin |  | 500 | mg | 250 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:10 | 01:00 | mg | 210 | 575 |  | n/a
53 | Adult | aztreonam | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
54 | Adult | aztreonam | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
55 | Adult | aztreonam | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg | 250 | 2000 |  | n/a
56 | Adult | aztreonam | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg | 250 | 2000 |  | n/a
57 | Adult | aztreonam | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.25 | 2 |  | n/a
58 | Adult | aztreonam | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.25 | 2 |  | n/a
59 | Adult | aztreonam | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | gram | 0.25 | 2 |  | n/a
60 | Adult | aztreonam | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 02:00 | gram | 0.25 | 2 |  | n/a
61 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 110 | mL | 6.3636 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:20 | 06:00 | 00:22 | mg | 699 | 701 |  | n/a
62 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 160 | mL | 4.375 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 06:00 | 00:31 | mg | 699 | 701 |  | n/a
63 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 210 | mL | 3.3333 | mg/mL | Pump | Primary and Secondary | 00:36 | 00:37 | 06:00 | 00:41 | mg | 699 | 701 |  | n/a
64 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 310 | mL | 2.2581 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 06:00 | 01:00 | mg | 699 | 701 |  | n/a
65 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 110 | mL | 6.3636 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:20 | 06:00 | 00:22 | mg | 699 | 701 |  | n/a
66 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 160 | mL | 4.375 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 06:00 | 00:31 | mg | 699 | 701 |  | n/a
67 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 210 | mL | 3.3333 | mg/mL | Pump | Primary and Secondary | 00:36 | 00:37 | 06:00 | 00:41 | mg | 699 | 701 |  | n/a
68 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 310 | mL | 2.2581 | mg/mL | Pump | Primary and Secondary | 00:59 | 01:00 | 06:00 | 01:10 | mg | 699 | 701 |  | n/a
69 | Adult | bamlanivimab |  | 700 | mg | 270 | mL | 2.5926 | mg/mL | Pump | Primary and Secondary | 01:00 |  | 10:00 | 01:00 | mg | 650 | 750 |  | n/a |  |  |  | Low Binding Filter
70 | Adult | caffeine citrate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 06:00 |  | mg | 1 | 1200 |  | mg/mL |  | 0.01 | 10
71 | Adult | caffeine/sod benzoa |  | 500 | mg | 1000 | mL | 0.5 | mg/mL | Pump | Primary and Secondary | 01:01 |  | 08:00 | 01:02 | mg | 50 | 501 | 1001 | n/a |  |  |  | 0.22 Filter Needed
72 | Adult | caffeine/sod benzoa |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 04:00 | 01:00 | mg | 50 | 501 | 1001 | mg/mL |  | 0.24 | 6 | 0.22 Filter Needed
73 | Adult | calcium chloride | .MG Dosing | 1000 | mg | 60 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | mg | 100 | 2000 |  | n/a
74 | Adult | calcium chloride | .MG Dosing | 1000 | mg | 110 | mL | 9.0909 | mg/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | mg | 100 | 2000 |  | n/a
75 | Adult | calcium chloride | .MG Dosing | 2000 | mg | 120 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 | 02:00 | mg | 100 | 2000 |  | n/a
76 | Adult | calcium chloride | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 |  | mg | 100 | 2000 |  | mg/mL |  | 0.1 | 20
77 | Adult | calcium chloride | GM dosing | 1 | gram | 60 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | gram | 0.1 | 10 |  | n/a
78 | Adult | calcium chloride | GM dosing | 1 | gram | 110 | mL | 0.0091 | gram/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | gram | 0.1 | 10 |  | n/a
79 | Adult | calcium chloride | GM dosing | 2 | gram | 120 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 | 02:00 | gram | 0.1 | 10 |  | n/a
80 | Adult | calcium chloride | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 |  | gram | 0.1 | 10 |  | gram/mL |  | 0.001 | 0.04
81 | Adult | calcium gluconate | .MG Dosing | 1000 | mg | 60 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | mg | 500 | 2400 |  | n/a
82 | Adult | calcium gluconate | .MG Dosing | 1000 | mg | 110 | mL | 9.0909 | mg/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | mg | 500 | 2400 |  | n/a
83 | Adult | calcium gluconate | .MG Dosing | 2000 | mg | 70 | mL | 28.5714 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | mg | 500 | 2400 |  | n/a
84 | Adult | calcium gluconate | .MG Dosing | 2000 | mg | 120 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | mg | 500 | 2400 |  | n/a
85 | Adult | calcium gluconate | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:25 | 06:00 |  | mg | 500 | 2400 |  | mg/mL | 1 |  | 40
86 | Adult | calcium gluconate | GM dosing | 1 | gram | 60 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
87 | Adult | calcium gluconate | GM dosing | 1 | gram | 110 | mL | 0.0091 | gram/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
88 | Adult | calcium gluconate | GM dosing | 2 | gram | 70 | mL | 0.0286 | gram/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
89 | Adult | calcium gluconate | GM dosing | 2 | gram | 120 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
90 | Adult | calcium gluconate | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 | 00:25 | 06:00 |  | gram | 0.5 | 2.4 |  | gram/mL | 0.001 |  | 0.04
91 | Adult | casirivimab/imdevima |  | 1200 | mg | 260 | mL | 4.6154 | mg/mL | Pump | Primary and Secondary | 00:48 | 00:49 | 04:00 | 01:00 | mg | 1190 | 1201 | 1210 | n/a |  |  |  | Low Binding Filter
92 | Adult | casirivimab/imdevima |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:18 | 00:19 | 04:00 |  | mg | 1190 | 1201 | 1210 | No |  |  |  | Low Binding Filter
93 | Adult | caspofungin |  | 35 | mg | 100 | mL | 0.35 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg | 35 | 90 |  | n/a |  |  |  | Dedicated IV Line
94 | Adult | caspofungin |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg | 35 | 90 |  | n/a |  |  |  | Dedicated IV Line
95 | Adult | caspofungin |  | 70 | mg | 250 | mL | 0.28 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg | 35 | 90 |  | n/a |  |  |  | Dedicated IV Line
96 | Adult | ceFAZolin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 |  | mg | 480 | 3100 |  | n/a
97 | Adult | ceFAZolin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 3100 |  | n/a
98 | Adult | ceFAZolin | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 3100 |  | n/a
99 | Adult | ceFAZolin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 3100 |  | n/a
100 | Adult | ceFAZolin | .MG Dosing | 3000 | mg | 100 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 3100 |  | n/a
101 | Adult | ceFAZolin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 3.1 |  | n/a
102 | Adult | ceFAZolin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 3.1 |  | n/a
103 | Adult | ceFAZolin | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 3.1 |  | n/a
104 | Adult | ceFAZolin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 3.1 |  | n/a
105 | Adult | ceFAZolin | GM dosing | 3 | gram | 100 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 3.1 |  | n/a
106 | Adult | cefepime | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
107 | Adult | cefepime | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
108 | Adult | cefepime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
109 | Adult | cefepime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
110 | Adult | cefepime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
111 | Adult | cefepime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
112 | Adult | cefepime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
113 | Adult | cefepime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
114 | Adult | cefepime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
115 | Adult | cefepime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
116 | Adult | cefiderocol | .MG Dosing | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
117 | Adult | cefiderocol | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
118 | Adult | cefiderocol | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
119 | Adult | cefiderocol | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
120 | Adult | cefiderocol | GM dosing | 0.75 | gram | 100 | mL | 0.0075 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
121 | Adult | cefiderocol | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
122 | Adult | cefiderocol | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
123 | Adult | cefiderocol | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
124 | Adult | cefotaxime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
125 | Adult | cefotaxime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
126 | Adult | cefotaxime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
127 | Adult | cefotaxime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
128 | Adult | cefotaxime | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 40
129 | Adult | cefotaxime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
130 | Adult | cefotaxime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
131 | Adult | cefotaxime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
132 | Adult | cefotaxime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
133 | Adult | cefotaxime | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.04
134 | Adult | cefoTEtan | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
135 | Adult | cefoTEtan | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
136 | Adult | cefoTEtan | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
137 | Adult | cefoTEtan | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
138 | Adult | cefoTEtan | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
139 | Adult | cefoTEtan | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
140 | Adult | cefoTEtan | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
141 | Adult | cefoTEtan | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
142 | Adult | cefOXitin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
143 | Adult | cefOXitin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
144 | Adult | cefOXitin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 20
145 | Adult | cefOXitin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
146 | Adult | cefOXitin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
147 | Adult | cefOXitin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.02
148 | Adult | ceftaroline |  | 200 | mg | 50 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
149 | Adult | ceftaroline |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
150 | Adult | ceftaroline |  | 300 | mg | 50 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
151 | Adult | ceftaroline |  | 300 | mg | 100 | mL | 3 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
152 | Adult | ceftaroline |  | 400 | mg | 50 | mL | 8 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
153 | Adult | ceftaroline |  | 400 | mg | 100 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
154 | Adult | ceftaroline |  | 600 | mg | 50 | mL | 12 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
155 | Adult | ceftaroline |  | 600 | mg | 100 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
156 | Adult | ceftaroline |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:07 | 04:00 | 01:00 | mg | 200 | 600 |  | mg/mL |  | 0.1 | 14
157 | Adult | cefTAZidime | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
158 | Adult | cefTAZidime | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
159 | Adult | cefTAZidime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
160 | Adult | cefTAZidime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
161 | Adult | cefTAZidime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
162 | Adult | cefTAZidime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
163 | Adult | cefTAZidime | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
164 | Adult | cefTAZidime | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
165 | Adult | cefTAZidime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
166 | Adult | cefTAZidime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
167 | Adult | cefTAZidime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
168 | Adult | cefTAZidime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
169 | Adult | ceftolozane/tazo | .MG Dosing | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 149 | 3010 |  | n/a
170 | Adult | ceftolozane/tazo | .MG Dosing | 375 | mg | 100 | mL | 3.75 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 149 | 3010 |  | n/a
171 | Adult | ceftolozane/tazo | .MG Dosing | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 149 | 3010 |  | n/a
172 | Adult | ceftolozane/tazo | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 149 | 3010 |  | n/a
173 | Adult | ceftolozane/tazo | .MG Dosing | 3000 | mg | 100 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 06:00 | 03:00 | mg | 149 | 3010 |  | n/a
174 | Adult | ceftolozane/tazo | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 149 | 3010 |  | mg/mL | 1 | 1.4 | 30.1
175 | Adult | ceftolozane/tazo | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | gram | 0.149 | 3.1 |  | n/a
176 | Adult | ceftolozane/tazo | GM dosing | 3 | gram | 100 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 06:00 | 03:00 | gram | 0.149 | 3.1 |  | n/a
177 | Adult | ceftolozane/tazo | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | gram | 0.149 | 3.1 |  | gram/mL | 0.001 |  | 0.031
178 | Adult | cefTRIAXone | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
179 | Adult | cefTRIAXone | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
180 | Adult | cefTRIAXone | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
181 | Adult | cefTRIAXone | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
182 | Adult | cefTRIAXone | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | mg/mL | 4.9 | 9 | 40
183 | Adult | cefTRIAXone | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
184 | Adult | cefTRIAXone | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
185 | Adult | cefTRIAXone | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
186 | Adult | cefTRIAXone | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
187 | Adult | cefTRIAXone | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | gram/mL | 0.004 | 0.009 | 0.04
188 | Adult | cefuroxime | .MG Dosing | 750 | mg | 50 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
189 | Adult | cefuroxime | .MG Dosing | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
190 | Adult | cefuroxime | .MG Dosing | 1500 | mg | 50 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
191 | Adult | cefuroxime | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
192 | Adult | cefuroxime | GM dosing | 0.75 | gram | 50 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
193 | Adult | cefuroxime | GM dosing | 0.75 | gram | 100 | mL | 0.0075 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
194 | Adult | cefuroxime | GM dosing | 1.5 | gram | 50 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
195 | Adult | cefuroxime | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
196 | Adult | chloramphenicol |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg/kg | 3 | 25 | 30 | mg/mL | 5 | 10 | 110
197 | Adult | chlorothiazide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 1000 | 2300 | mg/mL | 1 | 2 | 20
198 | Adult | chlorproMAZINE |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 02:00 | 00:15 | mg | 1 | 51 |  | mg/mL |  | 0.001 | 1.1 | Patient Lie Down
199 | Adult | ciprofloxacin |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
200 | Adult | ciprofloxacin |  | 400 | mg | 200 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
201 | Adult | clindamycin |  | 300 | mg | 50 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
202 | Adult | clindamycin |  | 300 | mg | 100 | mL | 3 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
203 | Adult | clindamycin |  | 600 | mg | 50 | mL | 12 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
204 | Adult | clindamycin |  | 600 | mg | 100 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
205 | Adult | clindamycin |  | 900 | mg | 50 | mL | 18 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
206 | Adult | clindamycin |  | 900 | mg | 100 | mL | 9 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
207 | Adult | colistimethate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 10 | 300 |  | mg/mL | 0.1 | 0.5 | 5
208 | Adult | colistimethate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 0.3 |  | 5 | mg/mL | 0.1 | 0.5 | 5
209 | Adult | conivaptan Loading |  | 20 | mg | 100 | mL | 0.2 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 15 | 25 |  | n/a
210 | Adult | dalbavancin | .MG Dosing | 375 | mg | 250 | mL | 1.5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
211 | Adult | dalbavancin | .MG Dosing | 500 | mg | 250 | mL | 2 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
212 | Adult | dalbavancin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
213 | Adult | dalbavancin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
214 | Adult | dalbavancin | .MG Dosing | 1125 | mg | 250 | mL | 4.5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
215 | Adult | dalbavancin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
216 | Adult | dalbavancin | GM dosing | 0.375 | gram | 250 | mL | 0.0015 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
217 | Adult | dalbavancin | GM dosing | 0.5 | gram | 250 | mL | 0.002 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
218 | Adult | dalbavancin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
219 | Adult | dalbavancin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
220 | Adult | dalbavancin | GM dosing | 1.125 | gram | 250 | mL | 0.0045 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
221 | Adult | dalbavancin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.51 | n/a |  |  |  | Flush Before/AfterD5
222 | Adult | dantrolene |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg/kg | 1 | 2.5 |  | mg/mL | 0.05 | 0.1 | 100
223 | Adult | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg | 100 | 2000 |  | mg/mL | 1 | 2 | 20
224 | Adult | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg/kg | 4 | 12.1 |  | mg/mL | 1 | 2 | 20
225 | Adult | deferoxamine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 04:00 | 08:00 | 24:00 |  | mg | 500 | 6000 |  | mg/mL | 1 |  | 95
226 | Adult | deferoxamine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 04:00 | 08:00 | 24:00 |  | mg/kg | 30 | 60 |  | mg/mL | 1 |  | 95
227 | Adult | desmopressin |  | - - | mcg | - - | mL |  | mcg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mcg | 0.1 | 22 |  | mcg/mL | 0.005 | 0.01 | 4
228 | Adult | desmopressin |  | - - | mcg | - - | mL |  | mcg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mcg/kg | 0.25 | 0.41 |  | mcg/mL | 0.005 | 0.01 | 4
229 | Adult | dexamethasone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:07 | 02:00 | 00:30 | mg | 1 | 20 |  | mg/mL | 0.01 |  | 10
230 | Adult | digoxin immune fab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 10 | 800 |  | mg/mL | 0.05 | 0.1 | 1
231 | Adult | doripenem |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 100 | 500 |  | n/a
232 | Adult | doripenem |  | 250 | mg | 100 | mL | 2.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 100 | 500 |  | n/a
233 | Adult | doripenem |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 100 | 500 |  | n/a
234 | Adult | doxycycline |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
235 | Adult | doxycycline |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
236 | Adult | ertapenem | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
237 | Adult | ertapenem | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
238 | Adult | ertapenem | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
239 | Adult | ertapenem | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
240 | Adult | ertapenem | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
241 | Adult | ertapenem | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
242 | Adult | ertapenem | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
243 | Adult | ertapenem | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
244 | Adult | erythromycin |  | 250 | mg | 100 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
245 | Adult | erythromycin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
246 | Adult | erythromycin |  | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:30 | 01:00 | mg | 100 | 1000 |  | n/a
247 | Adult | erythromycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | mg/mL | 0.05 | 1 | 5
248 | Adult | famotidine |  | 20 | mg | 50 | mL | 0.4 | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg | 8 | 44 | 50 | n/a
249 | Adult | famotidine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 02:00 |  | mg | 8 | 44 | 50 | mg/mL | 0.09 | 0.1 | 3.5
250 | Adult | fat emulsion 20% |  | 25 | gram | 125 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 11:00 | 24:00 |  | gram | 24 | 51 |  | n/a
251 | Adult | fat emulsion 20% |  | 50 | gram | 250 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 11:00 | 24:00 |  | gram | 24 | 51 |  | n/a
252 | Adult | ferric carboxymaltos |  | 750 | mg | 265 | mL | 2.8302 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg | 250 | 750 | 1501 | n/a |  |  |  | Hypersensitivity
253 | Adult | ferric carboxymaltos |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg | 250 | 750 | 1501 | mg/mL | 1.9 | 2 | 4 | Hypersensitivity
254 | Adult | ferric carboxymaltos |  | 750 | mg | 265 | mL | 2.8302 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg/kg | 5 | 16 |  | n/a |  |  |  | Hypersensitivity
255 | Adult | ferric carboxymaltos |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg/kg | 5 | 16 |  | mg/mL | 1.9 | 2 | 4 | Hypersensitivity
256 | Adult | ferumoxytol | .MG Dosing | 510 | mg | 117 | mL | 4.359 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 509 | 1021 | 1022 | n/a |  |  |  | Hypersensitivity
257 | Adult | ferumoxytol | .MG Dosing | 1020 | mg | 284 | mL | 3.5915 | mg/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | mg | 509 | 1021 | 1022 | n/a |  |  |  | Hypersensitivity
258 | Adult | ferumoxytol | GM dosing | 0.51 | gram | 117 | mL | 0.0044 | gram/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | gram | 0.509 | 1.021 | 1.022 | n/a |  |  |  | Hypersensitivity
259 | Adult | ferumoxytol | GM dosing | 1.02 | gram | 284 | mL | 0.0036 | gram/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | gram | 0.509 | 1.021 | 1.022 | n/a |  |  |  | Hypersensitivity
260 | Adult | fluconazole |  | 100 | mg | 50 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 04:00 | 00:30 | mg | 50 | 801 |  | n/a
261 | Adult | fluconazole |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 04:00 | 01:00 | mg | 50 | 801 |  | n/a
262 | Adult | fluconazole |  | 400 | mg | 200 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 01:00 | 04:00 | 02:00 | mg | 50 | 801 |  | n/a
263 | Adult | fluconazole |  | 800 | mg | 400 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 02:00 | 08:00 | 04:00 | mg | 50 | 801 |  | n/a
264 | Adult | folic acid |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 0.1 | 5 |  | mg/mL | 0.001 | 0.009 | 0.06
265 | Adult | fomepizole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 9 | 16 |  | mg/mL | 0.5 | 1 | 20
266 | Adult | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:11 | 01:00 | 00:20 | mg | 50 | 1500 |  | mg/mL | 0.5 | 1 | 30
267 | Adult | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:11 | 01:00 | 00:20 | mg/kg | 1 | 20 |  | mg/mL | 0.5 | 1 | 30
268 | Adult | ganciclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg/kg | 0.625 | 8 |  | mg/mL | 0.5 | 1 | 10 | In-line filter
269 | Adult | gentamicin |  | 80 | mg | 100 | mL | 0.8 | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 | 00:30 | mg | 10 | 1000 |  | n/a
270 | Adult | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 | 01:00 | mg | 10 | 1000 |  | mg/mL | 0.05 | 0.1 | 10
271 | Adult | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 |  | 01:20 | 01:00 | mg/kg | 0.5 | 10 |  | mg/mL | 0.05 | 0.1 | 10
272 | Adult | hetastarch 6% |  | 30 | gram | 500 | mL | 0.06 | gram/mL | Pump | Primary and Secondary |  | 00:31 | 24:00 |  | gram | 29 | 31 |  | n/a
273 | Adult | hydrocortisone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:30 | mg | 100 | 1000 |  | mg/mL | 0.5 | 1 | 20
274 | Adult | ibutilide | pt weight < 60 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:07 |  | 00:30 | 00:10 | mg/kg | 0.009 | 0.011 |  | mg/mL | 0.001 |  | 0.02
275 | Adult | ibutilide | pt weight >= 60 kg | 1 | mg | 60 | mL | 0.0167 | mg/mL | Pump | Primary and Secondary | 00:07 |  | 00:30 | 00:10 | mg | 0.4 | 1.25 |  | n/a
276 | Adult | imipenem/cilastatin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 199 | 1000 |  | n/a
277 | Adult | imipenem/cilastatin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 199 | 1000 |  | n/a
278 | Adult | imipenem/cilastatin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 01:00 | mg | 199 | 1000 |  | mg/mL | 1 | 2 | 5
279 | Adult | imipenem/cilastatin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 0.5 | 1 |  | n/a
280 | Adult | imipenem/cilastatin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | gram | 0.5 | 1 |  | n/a
281 | Adult | imipenem/cilastatin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 01:00 | gram | 0.5 | 1 |  | gram/mL | 0.001 | 0.003 | 0.006
282 | Adult | immune glob (IVIG) |  | 5 | gram | 50 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:00 |  | gram | 1 | 10.1 |  | n/a |  |  |  | Hypersensitivity
283 | Adult | immune glob (IVIG) |  | 10 | gram | 100 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:40 | 08:00 |  | gram | 1 | 10.1 |  | n/a |  |  |  | Hypersensitivity
284 | Adult | immune glob (IVIG) |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 02:00 | 36:00 |  | gram/kg | 1 | 2.4 |  | gram/mL | 0.01 | 0.05 | 0.1 | Hypersensitivity
285 | Adult | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg | 50 | 2000 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
286 | Adult | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg/kg | 2 | 10 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
287 | Adult | iron sucrose |  | 100 | mg | 105 | mL | 0.9524 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 08:00 | 00:30 | mg | 99 |  | 510 | n/a |  |  |  | Hypersensitivity
288 | Adult | iron sucrose |  | 200 | mg | 110 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 08:00 | 00:30 | mg | 99 |  | 510 | n/a |  |  |  | Hypersensitivity
289 | Adult | iron sucrose |  | 300 | mg | 265 | mL | 1.1321 | mg/mL | Pump | Primary and Secondary | 01:25 |  | 08:00 | 01:30 | mg | 99 |  | 510 | n/a |  |  |  | Hypersensitivity
290 | Adult | iron sucrose |  | 400 | mg | 270 | mL | 1.4815 | mg/mL | Pump | Primary and Secondary | 02:25 |  | 08:00 | 02:30 | mg | 99 |  | 510 | n/a |  |  |  | Hypersensitivity
291 | Adult | iron sucrose |  | 500 | mg | 275 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 03:25 |  | 08:00 | 04:00 | mg | 99 |  | 510 | n/a |  |  |  | Hypersensitivity
292 | Adult | itraconazole |  | 200 | mg | 60 | mL | 3.3333 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 100 | 400 |  | n/a |  |  |  | In-line filter
293 | Adult | itraconazole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 100 | 400 |  | mg/mL | 3.29 | 3.3 | 3.4 | In-line filter
294 | Adult | Kcentra (Factor IX) | .INR 2 to < 4 | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | unit/kg | 24 | 26 |  | unit/mL | 19 | 20 | 30 | INR 2-4
295 | Adult | Kcentra (Factor IX) | .INR 4 - 6 | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | unit/kg | 34 | 36 |  | unit/mL | 19 | 20 | 30 | INR 4-6
296 | Adult | Kcentra (Factor IX) | INR > 6 | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | unit/kg | 49 | 53 |  | unit/mL | 19 | 20 | 30 | INR > 6
297 | Adult | Kinevac (sincalide) |  | - - | mcg | - - | mL |  | mcg/mL | Pump | Primary and Secondary |  | 00:12 | 02:00 | 00:16 | mcg/kg | 0.018 | 0.13 |  | mcg/mL | 0.01 |  | 0.5
298 | Adult | lacosamide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 25 | 200 |  | mg/mL | 0.25 | 0.5 | 4.1
299 | Adult | levETIRAcetam |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:06 | 02:00 | 00:15 | mg | 250 | 2001 |  | mg/mL | 2 | 2.5 | 20.1
300 | Adult | levETIRAcetam |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:06 | 02:00 | 00:15 | mg/kg | 1 | 60 |  | mg/mL | 2 | 2.5 | 20.1
301 | Adult | levoFLOXacin |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
302 | Adult | levoFLOXacin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
303 | Adult | levoFLOXacin |  | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 01:15 | 02:00 | 01:30 | mg | 250 | 750 |  | n/a
304 | Adult | linezolid |  | 600 | mg | 300 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 04:00 | 01:00 | mg | 400 | 800 |  | n/a
305 | Adult | mag sulfate |  | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 02:00 | gram | 0.5 | 7 | 8 | n/a
306 | Adult | mag sulfate |  | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:30 | 06:00 | 04:00 | gram | 0.5 | 7 | 8 | n/a
307 | Adult | mag sulfate |  | 4 | gram | 100 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 09:00 |  | gram | 0.5 | 7 | 8 | n/a
308 | Adult | mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:30 | 12:00 |  | gram | 0.5 | 7 | 8 | gram/mL | 0.005 | 0.009 | 0.12
309 | Adult | mannitol 20% |  | 100 | gram | 500 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 00:31 | 04:00 |  | gram | 10 | 100 |  | n/a |  |  |  | In-line filter
310 | Adult | mannitol 20% |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:31 | 08:00 |  | gram | 10 | 100 |  | gram/mL | 0.15 | 0.18 | 0.21 | In-line filter
311 | Adult | mannitol 20% |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:31 | 08:00 |  | gram/kg | 0.1 | 2.1 |  | gram/mL | 0.15 | 0.18 | 0.21 | In-line filter
312 | Adult | meropenem | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
313 | Adult | meropenem | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
314 | Adult | meropenem | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
315 | Adult | meropenem | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
316 | Adult | meropenem | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
317 | Adult | meropenem | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
318 | Adult | meropenem | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
319 | Adult | meropenem | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
320 | Adult | meropenem | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
321 | Adult | meropenem | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
322 | Adult | meropenem/vaborbacta |  | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
323 | Adult | meropenem/vaborbacta |  | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
324 | Adult | meropenem/vaborbacta |  | 2 | gram | 250 | mL | 0.008 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
325 | Adult | meropenem/vaborbacta |  | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
326 | Adult | meropenem/vaborbacta |  | 4 | gram | 250 | mL | 0.016 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
327 | Adult | meropenem/vaborbacta |  | 4 | gram | 500 | mL | 0.008 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
328 | Adult | meropenem/vaborbacta |  | 4 | gram | 1000 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
329 | Adult | methocarbamol |  | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:10 |  | 04:00 |  | mg | 249 |  | 1001 | n/a
330 | Adult | methocarbamol |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 04:00 |  | mg | 249 |  | 1001 | mg/mL | 0.5 | 1 | 15
331 | Adult | methylPREDNISolone |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | n/a
332 | Adult | methylPREDNISolone |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | n/a
333 | Adult | methylPREDNISolone |  | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 125 | 2000 |  | n/a
334 | Adult | methylPREDNISolone |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 125 | 2000 |  | n/a
335 | Adult | methylPREDNISolone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | mg/mL | 0.5 | 1 | 25
336 | Adult | metroNIDAZOLE |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 250 | 1000 |  | n/a
337 | Adult | metroNIDAZOLE |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 250 | 1000 |  | n/a
338 | Adult | micafungin |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
339 | Adult | micafungin |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
340 | Adult | micafungin |  | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
341 | Adult | MVI + Thiame + Folic |  | 10 | unit | 1011 | mL | 0.0099 | unit/mL | Pump | Primary and Secondary | 01:01 |  | 24:00 | 08:06 | unit | 9.9 | 10 |  | n/a
342 | Adult | MVI Only |  | 10 | unit | 510 | mL | 0.0196 | unit/mL | Pump | Primary and Secondary | 00:31 |  | 24:00 | 04:05 | unit | 9.9 | 10.1 |  | n/a
343 | Adult | MVI Only |  | 10 | unit | 1010 | mL | 0.0099 | unit/mL | Pump | Primary and Secondary | 01:01 |  | 24:00 | 08:05 | unit | 9.9 | 10.1 |  | n/a
344 | Adult | nafcillin | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
345 | Adult | nafcillin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
346 | Adult | nafcillin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
347 | Adult | nafcillin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
348 | Adult | nafcillin | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | mg | 200 | 2000 |  | n/a
349 | Adult | nafcillin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | mg | 200 | 2000 |  | n/a
350 | Adult | nafcillin | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
351 | Adult | nafcillin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
352 | Adult | nafcillin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
353 | Adult | nafcillin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
354 | Adult | nafcillin | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | gram | 0.2 | 2 |  | n/a
355 | Adult | nafcillin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | gram | 0.2 | 2 |  | n/a
356 | Adult | ondansetron |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:30 | mg | 2 | 32 |  | mg/mL | 0.02 | 0.04 | 0.64
357 | Adult | pamidronate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 02:00 | 20:00 | 04:00 | mg | 29 | 99 | 110 | mg/mL | 0.001 |  | 0.4 | Dedicated IV Line
358 | Adult | pantoprazole bolus |  | 80 | mg | 100 | mL | 0.8 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:15 | mg | 40 | 80 |  | n/a
359 | Adult | penG 2.5 Mill Units |  | 2.5 | unit | 50 | mL | 0.05 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | unit | 2 | 3 |  | n/a |  |  |  | 1unit=1million
360 | Adult | penG 2.5 Mill Units |  | 2.5 | unit | 100 | mL | 0.025 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | unit | 2 | 3 |  | n/a |  |  |  | 1unit=1million
361 | Adult | penG 3 Million Units |  | 3 | unit | 50 | mL | 0.06 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | unit | 2.8 | 3.2 | 3.3 | n/a |  |  |  | 1unit=1million
362 | Adult | penG 5 Million Units |  | 5 | unit | 50 | mL | 0.1 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | unit | 4.5 | 5.5 |  | n/a |  |  |  | 1unit=1million
363 | Adult | penG 5 Million Units |  | 5 | unit | 100 | mL | 0.05 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | unit | 4.5 | 5.5 |  | n/a |  |  |  | 1unit=1million
364 | Adult | penG x Million Units |  | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | unit | 1 | 6 |  | unit/mL | 0.01 | 0.02 | 0.12 | 1unit=1million
365 | Adult | PHENobarbital |  | 130 | mg | 52 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:04 |  | 04:00 |  | mg | 5 | 270 |  | n/a
366 | Adult | PHENobarbital |  | 130 | mg | 102 | mL | 1.2745 | mg/mL | Pump | Primary and Secondary | 00:07 |  | 04:00 |  | mg | 5 | 270 |  | n/a
367 | Adult | PHENobarbital |  | 260 | mg | 54 | mL | 4.8148 | mg/mL | Pump | Primary and Secondary | 00:04 |  | 04:00 |  | mg | 5 | 270 |  | n/a
368 | Adult | PHENobarbital |  | 260 | mg | 104 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:07 |  | 04:00 |  | mg | 5 | 270 |  | n/a
369 | Adult | PHENobarbital |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:04 |  | 04:00 |  | mg | 5 | 270 |  | No
370 | Adult | phenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:30 | mg/kg | 0.5 | 20 |  | mg/mL | 0.5 | 1 | 10 | In-line filter
371 | Adult | phytonadione |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 01:00 | 00:30 | mg | 1 | 10 |  | mg/mL | 0.005 | 0.01 | 0.2
372 | Adult | piperacillin/tazo |  | 2.25 | gram | 50 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:30 | 00:30 | gram | 3 |  | 4.6 | n/a
373 | Adult | piperacillin/tazo |  | 2.25 | gram | 100 | mL | 0.0225 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:30 | 00:30 | gram | 3 |  | 4.6 | n/a
374 | Adult | piperacillin/tazo |  | 3.375 | gram | 50 | mL | 0.0675 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 | 04:00 | gram | 3 |  | 4.6 | n/a
375 | Adult | piperacillin/tazo |  | 3.375 | gram | 100 | mL | 0.0338 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 | 04:00 | gram | 3 |  | 4.6 | n/a
376 | Adult | piperacillin/tazo |  | 4.5 | gram | 50 | mL | 0.09 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
377 | Adult | piperacillin/tazo |  | 4.5 | gram | 100 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
378 | Adult | potassium acetate |  | 40 | mEq | 520 | mL | 0.0769 | mEq/mL | Pump | Primary and Secondary | 02:00 | 03:30 | 12:00 | 04:00 | mEq | 1 | 46 | 48 | n/a
379 | Adult | potassium acetate |  | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:45 |  | 12:00 |  | mEq | 1 | 46 | 48 | mEq/mL |  | 0.01 | 0.46
380 | Adult | potassium chloride | Central Line | 10 | mEq | 100 | mL | 0.1 | mEq/mL | Pump | Primary and Secondary | 00:14 | 00:45 | 06:00 | 01:00 | mEq | 5 |  | 41 | n/a
381 | Adult | potassium chloride | Central Line | 20 | mEq | 100 | mL | 0.2 | mEq/mL | Pump | Primary and Secondary | 00:25 | 01:30 | 08:00 | 02:00 | mEq | 5 |  | 41 | n/a
382 | Adult | potassium chloride | Central Line | 40 | mEq | 520 | mL | 0.0769 | mEq/mL | Pump | Primary and Secondary | 00:55 | 03:15 | 12:00 | 04:00 | mEq | 5 |  | 41 | n/a
383 | Adult | potassium chloride | Central Line | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:55 | 02:15 | 12:00 |  | mEq | 5 |  | 41 | mEq/mL | 0.01 | 0.02 | 0.21
384 | Adult | potassium chloride | Peripheral Line | 10 | mEq | 100 | mL | 0.1 | mEq/mL | Pump | Primary and Secondary | 00:25 | 00:45 | 06:00 | 01:00 | mEq | 5 |  | 41 | n/a
385 | Adult | potassium chloride | Peripheral Line | 40 | mEq | 520 | mL | 0.0769 | mEq/mL | Pump | Primary and Secondary | 03:15 |  | 12:00 | 04:00 | mEq | 5 |  | 41 | n/a
386 | Adult | potassium chloride | Peripheral Line | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:50 | 01:45 | 12:00 |  | mEq | 5 |  | 41 | mEq/mL | 0.04 | 0.07 | 0.11
387 | Adult | potassium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 01:00 | 02:00 | 09:00 |  | mmol | 15 | 40 |  | mmol/mL | 0.005 | 0.01 | 0.24
388 | Adult | pralidoxime | Load <67kg WtBasedGM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:15 | gram/kg | 0.001 |  | 0.034 | gram/mL |  | 0.009 | 0.051
389 | Adult | pralidoxime | Load <67kg WtBasedMG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:15 | mg/kg | 1 |  | 34 | mg/mL |  | 9 | 51
390 | Adult | pralidoxime | Load >67kg WtBasedGM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:20 | gram | 1 | 21 |  | gram/mL |  | 0.009 | 0.051
391 | Adult | pralidoxime | Load >67kg WtBasedMG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:20 | mg | 500 |  | 2300 | mg/mL |  | 9 | 51
392 | Adult | pralidoxime | Load NonWtBased MG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:20 | mg | 900 |  | 2300 | mg/mL |  | 9 | 51
393 | Adult | pralidoxime | Load NonWtBasedGM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:15 | gram | 0.9 |  | 2.3 | gram/mL |  | 0.009 | 0.051
394 | Adult | pralidoxime | Maintenance Dose GM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 |  | 24:00 |  | gram | 0.9 |  | 2.3 | gram/mL |  | 0.009 | 0.051
395 | Adult | pralidoxime | Maintenance Dose MG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 24:00 |  | mg | 900 |  | 2300 | mg/mL | 5 | 9 | 51
396 | Adult | promethazine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | mg | 6.25 | 25 |  | mg/mL | 0.031 |  | 1
397 | Adult | raNITIdine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg | 25 | 51 |  | mg/mL | 0.1 | 0.5 | 2.5
398 | Adult | remdesivir |  | 100 | mg | 230 | mL | 0.4348 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
399 | Adult | remdesivir |  | 100 | mg | 250 | mL | 0.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
400 | Adult | remdesivir |  | 200 | mg | 210 | mL | 0.9524 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
401 | Adult | remdesivir |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
402 | Adult | rifAMPin |  | 300 | mg | 250 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
403 | Adult | rifAMPin |  | 600 | mg | 500 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
404 | Adult | rifAMPin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | mg/mL | 0.1 | 0.5 | 6
405 | Adult | riTUXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:15 | 02:45 | 25:00 |  | mg | 250 | 1100 |  | mg/mL | 0.1 | 0.2 | 4 | Premedicate
406 | Adult | SMX/TMP |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:00 | 02:00 | 01:30 | mg/kg | 1.5 | 7 |  | mg/mL | 0.05 | 0.1 | 1.2 | 48 Hour Tubing
407 | Adult | sodium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 00:45 |  | 12:00 |  | mmol | 15 | 60 |  | mmol/mL | 0.005 | 0.01 | 0.25
408 | Adult | sotrovimab |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:22 | 00:23 | 04:00 | 00:30 | mg | 490 | 510 | 511 | n/a |  |  |  | Low Binding Filter
409 | Adult | sotrovimab |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:22 | 00:23 | 04:00 | 00:30 | mg | 490 | 510 | 511 | n/a |  |  |  | Low Binding Filter
410 | Adult | tedizolid phosphate |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 01:15 | 01:00 | mg | 175 | 225 | 250 | n/a
411 | Adult | thiamine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:15 | 00:30 | mg | 90 | 550 |  | mg/mL | 0.001 | 0.7 | 10.1
412 | Adult | ticarcillin/clav |  | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
413 | Adult | ticarcillin/clav |  | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
414 | Adult | ticarcillin/clav |  | 3.1 | gram | 50 | mL | 0.062 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
415 | Adult | ticarcillin/clav |  | 3.1 | gram | 100 | mL | 0.031 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
416 | Adult | tigecycline |  | 50 | mg | 50 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:15 | 00:30 | mg | 25 | 100 |  | n/a
417 | Adult | tigecycline |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:15 | 00:30 | mg | 25 | 100 |  | n/a
418 | Adult | tigecycline |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:15 | 00:30 | mg | 25 | 100 |  | n/a
419 | Adult | tobramycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 |  | 01:00 | 00:30 | mg/kg | 0.5 | 9 |  | mg/mL | 0.1 | 0.5 | 10
420 | Adult | tocilizumab |  | 800 | mg | 100 | mL | 8 | mg/mL | Pump | Primary and Secondary | 00:55 |  | 04:00 | 01:00 | mg | 161 |  | 801 | n/a |  |  |  | Dedicated IV Line
421 | Adult | tocilizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:55 |  | 04:00 | 01:00 | mg/kg | 3.9 |  | 8.1 | mg/mL | 1 | 1.1 | 8.1 | Dedicated IV Line
422 | Adult | tranexamic BOLUS | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:09 |  | 00:40 | 00:15 | mg | 500 | 2000 |  | n/a
423 | Adult | tranexamic BOLUS | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:07 | 00:20 | 00:40 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
424 | Adult | tranexamic BOLUS | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary | 00:10 |  | 00:40 | 00:15 | gram | 0.5 | 2 |  | n/a
425 | Adult | tranexamic BOLUS | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:07 | 00:20 | 00:40 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
426 | Adult | valproate sodium |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg | 250 | 2000 |  | mg/mL | 0.1 | 1 | 11
427 | Adult | valproate sodium |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg/kg | 1 | 40 |  | mg/mL | 0.1 | 1 | 11
428 | Adult | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 500 | 2000 |  | n/a
429 | Adult | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
430 | Adult | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
431 | Adult | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
432 | Adult | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
433 | Adult | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | mg | 500 | 2000 |  | n/a
434 | Adult | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg | 500 | 2000 |  | n/a
435 | Adult | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg | 500 | 2000 |  | n/a
436 | Adult | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
437 | Adult | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
438 | Adult | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
439 | Adult | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
440 | Adult | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | mg | 500 | 2000 |  | mg/mL | 1 | 2.9 | 5.1
441 | Adult | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg/kg | 5 |  | 30 | n/a
442 | Adult | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
443 | Adult | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
444 | Adult | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
445 | Adult | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
446 | Adult | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
447 | Adult | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
448 | Adult | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
449 | Adult | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
450 | Adult | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
451 | Adult | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
452 | Adult | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
453 | Adult | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | mg/kg | 5 |  | 30 | mg/mL | 1 | 2.9 | 5.1
454 | Adult | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.5 | 2 |  | n/a
455 | Adult | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
456 | Adult | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
457 | Adult | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
458 | Adult | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
459 | Adult | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | gram | 0.5 | 2 |  | n/a
460 | Adult | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram | 0.5 | 2 |  | n/a
461 | Adult | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram | 0.5 | 2 |  | n/a
462 | Adult | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
463 | Adult | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
464 | Adult | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
465 | Adult | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
466 | Adult | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.002 | 0.006
467 | Adult | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram/kg | 0.005 |  | 0.03 | n/a
468 | Adult | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
469 | Adult | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
470 | Adult | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
471 | Adult | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
472 | Adult | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
473 | Adult | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
474 | Adult | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
475 | Adult | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
476 | Adult | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
477 | Adult | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
478 | Adult | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
479 | Adult | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | gram/kg | 0.005 |  | 0.03 | gram/mL | 0.001 | 0.002 | 0.006
480 | Adult | voriconazole |  | 200 | mg | 50 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 02:00 | 01:00 | mg | 200 | 400 |  | n/a
481 | Adult | voriconazole |  | 400 | mg | 100 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 01:40 | 04:00 | 02:00 | mg | 200 | 400 |  | n/a
482 | Adult | voriconazole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 | 01:00 | 04:00 | 02:00 | mg/kg | 2 | 7 |  | mg/mL |  | 0.5 | 5
483 | Adult | zidovudine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 24:00 |  | mg/kg | 1 | 2.2 | 3 | mg/mL | 0.001 | 0.01 | 4.2
484 | Adult | zoledronic acid |  | 4 | mg | 100 | mL | 0.04 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | n/a
485 | Adult | zoledronic acid |  | 5 | mg | 100 | mL | 0.05 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | n/a
486 | Adult | zoledronic acid |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | mg/mL | 0.028 | 0.029 | 0.051
487 | Adult | zz amiodarone BOLUS |  | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:08 | 00:15 | 00:10 | mg | 149 | 151 |  | n/a |  |  |  | Code only drugs
488 | Adult | zz mag sulfate |  | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary | 00:07 | 00:14 | 02:00 | 00:20 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
489 | Adult | zz mag sulfate |  | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary | 00:07 | 00:14 | 02:00 | 00:20 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
490 | Adult | zz mag sulfate |  | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary | 00:14 | 00:28 | 04:00 | 00:40 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
491 | Adult | zz mag sulfate |  | 4 | gram | 100 | mL | 0.04 | gram/mL | Pump | Primary and Secondary | 00:28 | 00:56 | 08:00 | 01:20 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
492 | Adult | zz mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:14 | 00:28 | 08:00 |  | gram | 1 | 6 |  | gram/mL | 0.005 | 0.01 | 0.08 | Code only drugs
493 | Nursery | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:50 | 00:55 | 02:00 | 01:00 | mg/kg | 4 | 21 | 22 | mg/mL | 1 | 4 | 7
494 | Nursery | ampicillin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:07 | 00:08 | 01:10 | 00:30 | mg/kg | 24 | 101 | 105 | mg/mL |  | 10 | 105
495 | Nursery | ceFAZolin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:07 | 00:08 | 01:05 | 00:30 | mg/kg | 8 | 55 | 60 | mg/mL |  | 45 | 55
496 | Nursery | cefotaxime |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:12 | 00:13 | 01:05 | 00:30 | mg/kg | 12.5 | 80 | 81 | mg/mL |  | 45 | 55
497 | Nursery | cefTRIAXone |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:45 | 01:10 | 01:00 | mg/kg | 20 | 135 | 140 | mg/mL |  | 35 | 45 | CeftriSyringe
498 | Nursery | clindamycin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:08 | 00:09 | 01:05 | 00:30 | mg/kg | 3 | 7.1 | 7.2 | mg/mL |  | 13 | 17
499 | Nursery | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:25 | 02:15 | 00:30 | mg/kg | 0.5 | 7 | 7.1 | mg/mL |  | 3 | 7
500 | Nursery | nafcillin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:25 | 01:10 | 00:30 | mg/kg | 10 | 53 | 55 | mg/mL |  | 35 | 45
501 | Nursery | PenG Potass Syringe |  | - - | unit | - - | mL |  | unit/mL | Syringe | Primary Only | 00:15 |  | 02:00 | 00:30 | unit/kg | 20 | 52.5 | 152 | unit/mL | 10 |  | 150 | PenG-Syringe
502 | Nursery | piperacillin/tazo |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:25 | 04:15 | 00:30 | mg/kg | 49 | 125 | 130 | mg/mL |  | 15 | 85
503 | Nursery | vancomycin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:52 | 00:55 | 03:10 | 01:00 | mg/kg | 4 | 22 | 23 | mg/mL |  | 4.5 | 6.5
504 | ONC Infusion Center | abatacept |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
505 | ONC Infusion Center | abatacept |  | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
506 | ONC Infusion Center | abatacept |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
507 | ONC Infusion Center | aprepitant |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:30 | mg | 99 | 131 |  | n/a |  |  |  | DEHP
508 | ONC Infusion Center | aprepitant |  | 130 | mg | 100 | mL | 1.3 | mg/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:30 | mg | 99 | 131 |  | n/a |  |  |  | DEHP
509 | ONC Infusion Center | atezolizumab |  | 840 | mg | 250 | mL | 3.36 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 799 | 1681 |  | n/a
510 | ONC Infusion Center | atezolizumab |  | 1200 | mg | 250 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 799 | 1681 |  | n/a
511 | ONC Infusion Center | atezolizumab |  | 1680 | mg | 250 | mL | 6.72 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 799 | 1681 |  | n/a
512 | ONC Infusion Center | Bendamusti (Bendeka) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:08 | 00:15 | 00:10 | mg/m2 | 25 | 120 |  | mg/mL |  | 1.85 | 5.6
513 | ONC Infusion Center | Bendamusti (Treanda) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:20 | 00:30 | mg/m2 | 25 | 120 |  | mg/mL |  | 0.2 | 0.7
514 | ONC Infusion Center | bevacizumab | First Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:15 | 01:30 | 03:00 | 01:30 | mg/kg | 5 | 15 |  | mg/mL |  | 1 | 25
515 | ONC Infusion Center | bevacizumab | Second Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 01:30 | 01:00 | mg/kg | 5 | 15 |  | mg/mL |  | 1 | 25
516 | ONC Infusion Center | bevacizumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 01:30 | 00:30 | mg/kg | 5 | 15 |  | mg/mL |  | 1 | 25
517 | ONC Infusion Center | bleomycin | BSA Based Dose | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:05 | 00:15 | 00:10 | unit/m2 | 2 | 30 |  | unit/mL |  | 0.1 | 3
518 | ONC Infusion Center | bleomycin | Non-Based Based Dose | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:08 | 00:20 | 00:10 | unit | 1 | 60 | 61 | unit/mL |  | 0.1 | 3
519 | ONC Infusion Center | brentuximab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 00:40 | 00:30 | mg/kg | 1 | 1.8 | 1.81 | mg/mL |  | 0.4 | 1.8 | MAX DOSE 180 mg
520 | ONC Infusion Center | cabazitaxel |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:10 | 01:00 | mg/m2 | 15 | 25 |  | mg/mL |  | 0.1 | 0.26 | 0.22 Filter Needed
521 | ONC Infusion Center | CARBOplatin | Based on AUC | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 24:00 | 00:30 | mg | 100 | 800 |  | mg/mL |  | 0.4 | 5
522 | ONC Infusion Center | CARBOplatin | Based on BSA | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 | 00:30 | mg/m2 | 60 | 700 |  | mg/mL |  | 0.5 | 5
523 | ONC Infusion Center | carfilzomib |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:35 | 00:30 | mg/m2 | 15 | 70 |  | mg/mL |  | 0.3 | 1.6
524 | ONC Infusion Center | cemiplimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 00:35 | 00:30 | mg | 349 | 351 |  | mg/mL |  | 1 | 20 | In-line filter
525 | ONC Infusion Center | cetuximab | Loading Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:45 |  | 04:00 | 02:00 | mg/m2 | 400 | 500 |  | No |  |  |  | Low Binding Filter
526 | ONC Infusion Center | cetuximab | Maintenance Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg/m2 | 150 | 250 |  | No |  |  |  | Low Binding Filter
527 | ONC Infusion Center | CISplatin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 24:00 |  | mg/m2 | 10 | 100 |  | mg/mL |  | 0.04 | 0.5
528 | ONC Infusion Center | crizanlizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 04:30 | 00:30 | mg/kg | 4.5 | 5.5 | 5.51 | mg/mL | 1 | 1.6 | 5 | 0.22 Filter Needed
529 | ONC Infusion Center | cyclophosphamide | BSA Based Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:30 | 01:00 | mg/m2 | 100 | 2000 |  | mg/mL | 1.9 | 2 | 20
530 | ONC Infusion Center | cyclophosphamide | Weight Based | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:30 |  | mg/kg | 1 | 50 | 51 | mg/mL | 1.9 | 2 | 20
531 | ONC Infusion Center | dacarbazine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:40 |  | mg/m2 | 175 | 850 |  | mg/mL |  | 0.19 | 20.1
532 | ONC Infusion Center | daratumumab | First Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 02:20 | 02:30 | 20:00 |  | mg/kg | 7 | 16 | 16.1 | mg/mL |  | 0.4 | 4 | 0.22 Filter Needed
533 | ONC Infusion Center | daratumumab | Second Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 02:30 | 04:00 | 10:00 |  | mg/kg | 15 | 16 | 16.1 | mg/mL |  | 0.9 | 5 | 0.22 Filter Needed
534 | ONC Infusion Center | daratumumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 02:30 | 03:50 | 10:00 |  | mg/kg | 15 | 16 | 16.1 | mg/mL |  | 0.9 | 5 | 0.22 Filter Needed
535 | ONC Infusion Center | decitabine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 |  | mg/m2 | 10 | 22 |  | mg/mL | 0.089 | 0.09 | 1.1
536 | ONC Infusion Center | dexamethasone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:30 | 00:15 | mg | 1 | 21 |  | No
537 | ONC Infusion Center | DOCEtaxel |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg/m2 | 16 | 100 |  | mg/mL |  | 0.3 | 0.74 | DEHP
538 | ONC Infusion Center | DOXOrubicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 24:00 |  | mg/m2 | 3 | 90 |  | mg/mL |  | 0.01 | 5
539 | ONC Infusion Center | DOXOrubicin, liposo |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg/m2 | 2.5 | 55 |  | mg/mL |  | 0.1 | 0.36 | NO FILTER
540 | ONC Infusion Center | durvalumab | .Weight >= 30 kg | 1500 | mg | 280 | mL | 5.3571 | mg/mL | Pump | Primary Only |  | 00:55 | 08:00 | 01:00 | mg | 1499 |  | 1501 | n/a |  |  |  | Low Binding Filter
541 | ONC Infusion Center | durvalumab | Weight < 30 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only |  | 00:55 | 08:00 | 01:00 | mg/kg | 9 | 21 |  | mg/mL | 0.9 | 1 | 15 | Low Binding Filter
542 | ONC Infusion Center | eriBULin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:02 | 00:20 | 00:10 | mg/m2 | 0.7 | 1.4 |  | mg/mL |  | 0.001 | 0.5 | No Dextrose
543 | ONC Infusion Center | etoposide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 03:00 |  | mg/m2 | 35 | 750 |  | mg/mL |  | 0.2 | 0.4 | Non-PVC
544 | ONC Infusion Center | ferumoxytol | .MG Dosing | 510 | mg | 117 | mL | 4.359 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:30 | mg | 509 | 1021 | 1022 | n/a
545 | ONC Infusion Center | ferumoxytol | .MG Dosing | 1020 | mg | 284 | mL | 3.5915 | mg/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | mg | 509 | 1021 | 1022 | n/a
546 | ONC Infusion Center | ferumoxytol | GM dosing | 0.51 | gram | 117 | mL | 0.0044 | gram/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:30 | gram | 0.509 | 1.021 | 1.022 | n/a
547 | ONC Infusion Center | ferumoxytol | GM dosing | 1.02 | gram | 284 | mL | 0.0036 | gram/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | gram | 0.509 | 1.021 | 1.022 | n/a
548 | ONC Infusion Center | fluorouracil |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 24:00 |  | mg/m2 | 200 | 3000 |  | mg/mL |  | 0.1 | 50
549 | ONC Infusion Center | fosaprepitant |  | 150 | mg | 150 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 00:40 | 00:30 | mg | 149 | 151 |  | n/a
550 | ONC Infusion Center | gemcitabine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg/m2 | 300 | 1250 |  | mg/mL |  | 0.1 | 38
551 | ONC Infusion Center | golimumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:00 | 00:30 | mg/kg | 1.8 |  | 2.2 | mg/mL | 0.49 | 0.5 | 4 | Low Binding Filter
552 | ONC Infusion Center | ifosfamide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 25:00 |  | mg/m2 | 250 | 5000 |  | mg/mL |  | 0.6 | 80
553 | ONC Infusion Center | ipilimumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:25 | 01:35 | 01:30 | mg/kg | 1 | 10 |  | mg/mL |  | 1 | 2
554 | ONC Infusion Center | irinotecan |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:20 |  | 02:30 | 01:30 | mg/m2 | 20 | 600 |  | mg/mL |  | 0.12 | 2.8
555 | ONC Infusion Center | irinotecan liposomal |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:20 | 02:00 | 01:30 | mg/m2 | 34 | 71 |  | mg/mL | 0.05 | 0.1 | 0.6 | NO FILTER
556 | ONC Infusion Center | leucovorin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mg/m2 | 10 | 1000 |  | mg/mL |  | 0.05 | 20
557 | ONC Infusion Center | leucovorin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mg | 1 | 100 |  | mg/mL |  | 0.05 | 20
558 | ONC Infusion Center | mag sulfate |  | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 02:30 | 01:00 | gram | 0.9 | 1.1 |  | n/a
559 | ONC Infusion Center | mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:45 | 08:00 |  | gram | 0.9 | 1.1 |  | No
560 | ONC Infusion Center | mesna |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 25:00 |  | mg/m2 | 300 | 5000 |  | mg/mL |  | 0.01 | 20
561 | ONC Infusion Center | mesna |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 25:00 |  | mg | 1 | 1000 |  | mg/mL |  | 0.01 | 20
562 | ONC Infusion Center | mitoMYcin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg/m2 | 1 | 20 |  | mg/mL |  | 0.02 | 0.04
563 | ONC Infusion Center | mitoMYcin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg | 1 | 20 |  | mg/mL |  | 0.02 | 0.04
564 | ONC Infusion Center | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
565 | ONC Infusion Center | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
566 | ONC Infusion Center | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
567 | ONC Infusion Center | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
568 | ONC Infusion Center | nivolumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | mg/mL |  | 1 | 10 | In-line filter
569 | ONC Infusion Center | obinutuzumab | First Dose | 100 | mg | 104 | mL | 0.9615 | mg/mL | Pump | Primary and Secondary |  | 04:00 | 08:00 | 04:00 | mg | 99 | 100.1 |  | n/a
570 | ONC Infusion Center | obinutuzumab | Subsequent Doses | 900 | mg | 286 | mL | 3.1469 | mg/mL | Pump | Primary and Secondary |  | 02:00 | 60:00 |  | mg | 900 | 1000 |  | n/a
571 | ONC Infusion Center | obinutuzumab | Subsequent Doses | 1000 | mg | 290 | mL | 3.4483 | mg/mL | Pump | Primary and Secondary |  | 02:00 | 60:00 |  | mg | 900 | 1000 |  | n/a
572 | ONC Infusion Center | obinutuzumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 02:00 | 60:00 |  | mg | 900 | 1000 |  | mg/mL |  | 0.9 | 3.5
573 | ONC Infusion Center | ondansetron |  | 8 | mg | 54 | mL | 0.1481 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 4 | 16 |  | n/a
574 | ONC Infusion Center | ondansetron |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 |  | mg | 4 | 16 |  | mg/mL |  | 0.01 | 0.64
575 | ONC Infusion Center | oxaliplatin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:00 | 06:30 | 02:00 | mg/m2 | 20 | 130 |  | mg/mL |  | 0.2 | 1.2
576 | ONC Infusion Center | PACLitaxel |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 24:00 |  | mg/m2 | 40 | 225 |  | mg/mL |  | 0.3 | 1.2 | Filter and DEHP Free
577 | ONC Infusion Center | PACLitaxel protein-b |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/m2 | 49 | 261 |  | mg/mL |  | 4.9 | 5.1 | NO FILTER
578 | ONC Infusion Center | panitumumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg/kg | 3 | 6 |  | mg/mL |  | 2 | 10 | Low Binding Filter
579 | ONC Infusion Center | pembrolizumab |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 00:35 | 00:30 | mg | 100 | 250 |  | n/a |  |  |  | In-line filter
580 | ONC Infusion Center | pembrolizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 00:35 | 00:30 | mg | 100 | 250 |  | mg/mL |  | 1 | 10 | In-line filter
581 | ONC Infusion Center | PEMEtrexed |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:10 | mg/m2 | 240 | 550 | 910 | mg/mL |  | 3 | 13
582 | ONC Infusion Center | pertuzumab | Loading Dose | 840 | mg | 278 | mL | 3.0216 | mg/mL | Pump | Primary and Secondary | 01:00 |  | 02:00 | 01:00 | mg | 839 |  | 841 | n/a
583 | ONC Infusion Center | pertuzumab | Maintenance Dose | 420 | mg | 264 | mL | 1.5909 | mg/mL | Pump | Primary and Secondary | 00:30 |  | 02:00 | 00:30 | mg | 419 |  | 421 | n/a
584 | ONC Infusion Center | polatuzumab vedotin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 03:00 |  | mg/kg | 0.9 | 1.98 | 2 | mg/mL | 0.71 | 0.72 | 2.71 | Low Binding Filter
585 | ONC Infusion Center | ramucirumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 |  | mg/kg | 5 | 10 |  | mg/mL |  | 0.8 | 12 | Low Binding Filter
586 | ONC Infusion Center | raNITIdine |  | 50 | mg | 102 | mL | 0.4902 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:20 | mg | 25 | 50 |  | n/a
587 | ONC Infusion Center | riTUXimab | Based on BSA | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg/m2 | 250 | 500 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
588 | ONC Infusion Center | riTUXimab | Rheumatoid Arthritis | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg | 90 | 1200 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
589 | ONC Infusion Center | temozolomide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:20 | 02:00 | 01:30 | mg/m2 | 74 | 201 |  | mg/mL |  | 2.4 | 2.6 | Flush Before/AfterNS
590 | ONC Infusion Center | topotecan |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg/m2 | 0.2 | 1.5 | 4 | mg/mL |  | 0.01 | 0.5
591 | ONC Infusion Center | trabectedin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 02:45 | 02:50 | 04:00 | 03:00 | mg/m2 | 0.2 | 1.15 | 1.2 | mg/mL |  | 0.001 | 0.01 | 0.22 Filter Needed
592 | ONC Infusion Center | trastuzumab | Loading Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:00 | 03:00 | 01:30 | mg/kg | 4 | 8 |  | mg/mL |  | 0.4 | 8
593 | ONC Infusion Center | trastuzumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 2 | 6 |  | mg/mL |  | 0.4 | 8
594 | ONC Infusion Center | vinBLAStine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 01:00 | 00:15 | mg/m2 | 0.75 | 18.5 | 18.6 | mg/mL |  | 0.001 | 0.6 | Vesicant
595 | ONC Infusion Center | vinCRIStine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:20 | 00:10 | mg | 0.4 | 2 | 2.1 | mg/mL |  | 0.001 | 0.1 | Vesicant
596 | ONC Infusion Center | vinorelbine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:06 | 00:30 | 00:10 | mg/m2 | 5 | 35 |  | mg/mL |  | 0.5 | 2
597 | ONC Infusion Center | ziv-aflibercept |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:55 | 01:05 | 01:00 | mg/kg | 2 | 4 |  | mg/mL |  | 0.6 | 8 | 0.22 Filter Needed
598 | ONC Infusion Center | zoledronic acid |  | 4 | mg | 100 | mL | 0.04 | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 02:00 | 00:30 | mg | 3 | 4 |  | n/a
599 | OutPt InfusionCenter | abatacept |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
600 | OutPt InfusionCenter | abatacept |  | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
601 | OutPt InfusionCenter | abatacept |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
602 | OutPt InfusionCenter | abatacept |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg/kg | 9 |  | 11 | mg/mL | 2 | 3.8 | 10.2 | Low Binding 0.2or1.2
603 | OutPt InfusionCenter | albumin 25% |  | 5 | gram | 20 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:08 | 01:00 | 00:20 | gram | 4.9 | 25.1 |  | n/a
604 | OutPt InfusionCenter | albumin 25% |  | 12.5 | gram | 50 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:16 | 02:00 | 00:50 | gram | 4.9 | 25.1 |  | n/a
605 | OutPt InfusionCenter | albumin 25% |  | 25 | gram | 100 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:32 | 04:00 | 01:40 | gram | 4.9 | 25.1 |  | n/a
606 | OutPt InfusionCenter | amikacin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 |  | mg/kg | 1 | 21 | 39 | mg/mL | 0.2 |  | 5.1
607 | OutPt InfusionCenter | ampho B (liposomal) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only |  | 00:55 | 04:00 | 02:00 | mg/kg | 1.9 | 10.5 |  | mg/mL | 0.8 | 0.9 | 2.1 | AmBisome
608 | OutPt InfusionCenter | belimumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 00:55 |  | 08:00 | 01:00 | mg/kg | 9 |  | 11 | mg/mL | 1.6 |  | 8.5
609 | OutPt InfusionCenter | bezlotoxumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 | 00:51 | 03:00 | 01:00 | mg/kg | 9 | 11 | 11.1 | mg/mL | 0.8 | 0.9 | 10.1 | 0.2 to 5 micron filt
610 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
611 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
612 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
613 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
614 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | mg/mL | 5 | 10 | 40
615 | OutPt InfusionCenter | ceFAZolin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
616 | OutPt InfusionCenter | ceFAZolin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
617 | OutPt InfusionCenter | ceFAZolin | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
618 | OutPt InfusionCenter | ceFAZolin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
619 | OutPt InfusionCenter | ceFAZolin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | gram/mL | 0.005 | 0.01 | 0.04
620 | OutPt InfusionCenter | cefepime | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
621 | OutPt InfusionCenter | cefepime | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
622 | OutPt InfusionCenter | cefepime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
623 | OutPt InfusionCenter | cefepime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
624 | OutPt InfusionCenter | cefepime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
625 | OutPt InfusionCenter | cefepime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
626 | OutPt InfusionCenter | cefepime | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 4 | 4.9 | 40
627 | OutPt InfusionCenter | cefepime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
628 | OutPt InfusionCenter | cefepime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
629 | OutPt InfusionCenter | cefepime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
630 | OutPt InfusionCenter | cefepime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
631 | OutPt InfusionCenter | cefepime | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.004 | 0.009 | 0.04
632 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
633 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
634 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
635 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
636 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
637 | OutPt InfusionCenter | cefTAZidime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 249 | 2000 |  | n/a
638 | OutPt InfusionCenter | cefTAZidime | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
639 | OutPt InfusionCenter | cefTAZidime | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
640 | OutPt InfusionCenter | cefTAZidime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
641 | OutPt InfusionCenter | cefTAZidime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
642 | OutPt InfusionCenter | cefTAZidime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
643 | OutPt InfusionCenter | cefTAZidime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
644 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
645 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
646 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
647 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | n/a
648 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 499 | 2001 |  | mg/mL | 4.9 | 9 | 40
649 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
650 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
651 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
652 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | n/a
653 | OutPt InfusionCenter | cefTRIAXone | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.499 | 2.001 |  | gram/mL | 0.004 | 0.009 | 0.04
654 | OutPt InfusionCenter | ciprofloxacin |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
655 | OutPt InfusionCenter | ciprofloxacin |  | 400 | mg | 200 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
656 | OutPt InfusionCenter | clindamycin |  | 300 | mg | 50 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
657 | OutPt InfusionCenter | clindamycin |  | 300 | mg | 100 | mL | 3 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
658 | OutPt InfusionCenter | clindamycin |  | 600 | mg | 50 | mL | 12 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
659 | OutPt InfusionCenter | clindamycin |  | 600 | mg | 100 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
660 | OutPt InfusionCenter | clindamycin |  | 900 | mg | 50 | mL | 18 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
661 | OutPt InfusionCenter | clindamycin |  | 900 | mg | 100 | mL | 9 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
662 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 375 | mg | 100 | mL | 3.75 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
663 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 375 | mg | 250 | mL | 1.5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
664 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
665 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 500 | mg | 250 | mL | 2 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
666 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
667 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
668 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 1125 | mg | 250 | mL | 4.5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
669 | OutPt InfusionCenter | dalbavancin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 374 |  | 1501 | n/a |  |  |  | Flush Before/AfterD5
670 | OutPt InfusionCenter | dalbavancin | GM dosing | 0.375 | gram | 100 | mL | 0.0038 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
671 | OutPt InfusionCenter | dalbavancin | GM dosing | 0.375 | gram | 250 | mL | 0.0015 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
672 | OutPt InfusionCenter | dalbavancin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
673 | OutPt InfusionCenter | dalbavancin | GM dosing | 0.5 | gram | 250 | mL | 0.002 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
674 | OutPt InfusionCenter | dalbavancin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
675 | OutPt InfusionCenter | dalbavancin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
676 | OutPt InfusionCenter | dalbavancin | GM dosing | 1.125 | gram | 250 | mL | 0.0045 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
677 | OutPt InfusionCenter | dalbavancin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.374 |  | 1.501 | n/a
678 | OutPt InfusionCenter | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg | 100 | 2000 |  | mg/mL | 1 | 2 | 20
679 | OutPt InfusionCenter | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg/kg | 4 | 12.1 |  | mg/mL | 1 | 2 | 20
680 | OutPt InfusionCenter | dexamethasone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:30 | 00:15 | mg | 1 | 21 |  | No
681 | OutPt InfusionCenter | doxycycline |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
682 | OutPt InfusionCenter | doxycycline |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
683 | OutPt InfusionCenter | ertapenem | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
684 | OutPt InfusionCenter | ertapenem | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
685 | OutPt InfusionCenter | ertapenem | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
686 | OutPt InfusionCenter | ertapenem | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
687 | OutPt InfusionCenter | ertapenem | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
688 | OutPt InfusionCenter | ertapenem | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
689 | OutPt InfusionCenter | ertapenem | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
690 | OutPt InfusionCenter | ertapenem | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
691 | OutPt InfusionCenter | erythromycin |  | 250 | mg | 100 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
692 | OutPt InfusionCenter | erythromycin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
693 | OutPt InfusionCenter | erythromycin |  | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:30 | 01:00 | mg | 100 | 1000 |  | n/a
694 | OutPt InfusionCenter | erythromycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | mg/mL | 0.05 | 1 | 5
695 | OutPt InfusionCenter | famotidine |  | 20 | mg | 50 | mL | 0.4 | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg | 8 | 44 | 50 | n/a
696 | OutPt InfusionCenter | famotidine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 02:00 |  | mg | 8 | 44 | 50 | mg/mL | 0.09 | 0.1 | 3.5
697 | OutPt InfusionCenter | ferumoxytol | .MG Dosing | 510 | mg | 117 | mL | 4.359 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 509 | 1021 | 1022 | n/a |  |  |  | Hypersensitivity
698 | OutPt InfusionCenter | ferumoxytol | .MG Dosing | 1020 | mg | 284 | mL | 3.5915 | mg/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | mg | 509 | 1021 | 1022 | n/a |  |  |  | Hypersensitivity
699 | OutPt InfusionCenter | ferumoxytol | GM dosing | 0.51 | gram | 117 | mL | 0.0044 | gram/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | gram | 0.509 | 1.021 | 1.022 | n/a |  |  |  | Hypersensitivity
700 | OutPt InfusionCenter | ferumoxytol | GM dosing | 1.02 | gram | 284 | mL | 0.0036 | gram/mL | Pump | Primary and Secondary | 00:28 |  | 04:00 | 00:30 | gram | 0.509 | 1.021 | 1.022 | n/a |  |  |  | Hypersensitivity
701 | OutPt InfusionCenter | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 |  | mg | 10 | 1000 |  | mg/mL | 0.05 | 0.1 | 10
702 | OutPt InfusionCenter | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 |  | mg/kg | 0.5 | 10 |  | mg/mL | 0.05 | 0.1 | 10
703 | OutPt InfusionCenter | golimumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:00 | 00:30 | mg/kg | 1.8 |  | 2.2 | mg/mL | 0.49 | 0.5 | 4 | Low Binding Filter
704 | OutPt InfusionCenter | immune glob (IVIG) |  | 5 | gram | 50 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:00 |  | gram | 1 | 10.1 |  | n/a |  |  |  | Hypersensitivity
705 | OutPt InfusionCenter | immune glob (IVIG) |  | 10 | gram | 100 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:40 | 08:00 |  | gram | 1 | 10.1 |  | n/a |  |  |  | Hypersensitivity
706 | OutPt InfusionCenter | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg | 50 | 2000 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
707 | OutPt InfusionCenter | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg/kg | 2 | 10 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
708 | OutPt InfusionCenter | iron sucrose |  | 100 | mg | 105 | mL | 0.9524 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 08:00 | 00:30 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
709 | OutPt InfusionCenter | iron sucrose |  | 200 | mg | 110 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 08:00 | 00:30 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
710 | OutPt InfusionCenter | iron sucrose |  | 300 | mg | 265 | mL | 1.1321 | mg/mL | Pump | Primary and Secondary | 01:25 |  | 08:00 | 01:30 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
711 | OutPt InfusionCenter | iron sucrose |  | 400 | mg | 270 | mL | 1.4815 | mg/mL | Pump | Primary and Secondary | 02:25 |  | 08:00 | 02:30 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
712 | OutPt InfusionCenter | iron sucrose |  | 500 | mg | 275 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 03:25 |  | 08:00 | 04:00 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
713 | OutPt InfusionCenter | levoFLOXacin |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
714 | OutPt InfusionCenter | levoFLOXacin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
715 | OutPt InfusionCenter | levoFLOXacin |  | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 01:15 | 02:00 | 01:30 | mg | 250 | 750 |  | n/a
716 | OutPt InfusionCenter | mag sulfate |  | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 02:00 | gram | 0.9 | 7 | 8 | n/a
717 | OutPt InfusionCenter | mag sulfate |  | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:30 | 06:00 | 04:00 | gram | 0.9 | 7 | 8 | n/a
718 | OutPt InfusionCenter | mag sulfate |  | 4 | gram | 100 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 01:00 | 09:00 | 08:00 | gram | 0.9 | 7 | 8 | n/a
719 | OutPt InfusionCenter | mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:30 | 12:00 |  | gram | 0.9 | 7 | 8 | gram/mL | 0.005 | 0.009 | 0.12
720 | OutPt InfusionCenter | meropenem | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
721 | OutPt InfusionCenter | meropenem | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
722 | OutPt InfusionCenter | meropenem | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
723 | OutPt InfusionCenter | meropenem | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
724 | OutPt InfusionCenter | meropenem | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | mg | 500 | 2000 |  | n/a
725 | OutPt InfusionCenter | meropenem | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
726 | OutPt InfusionCenter | meropenem | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
727 | OutPt InfusionCenter | meropenem | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
728 | OutPt InfusionCenter | meropenem | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
729 | OutPt InfusionCenter | meropenem | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 03:10 | 00:30 | gram | 0.5 | 2 |  | n/a
730 | OutPt InfusionCenter | methylPREDNISolone |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | n/a
731 | OutPt InfusionCenter | methylPREDNISolone |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | n/a
732 | OutPt InfusionCenter | methylPREDNISolone |  | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 125 | 2000 |  | n/a
733 | OutPt InfusionCenter | methylPREDNISolone |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 04:00 | 01:00 | mg | 125 | 2000 |  | n/a
734 | OutPt InfusionCenter | methylPREDNISolone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 |  | mg | 125 | 2000 |  | mg/mL | 0.5 | 1 | 25
735 | OutPt InfusionCenter | micafungin |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
736 | OutPt InfusionCenter | micafungin |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
737 | OutPt InfusionCenter | micafungin |  | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg | 50 | 150 |  | n/a
738 | OutPt InfusionCenter | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
739 | OutPt InfusionCenter | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
740 | OutPt InfusionCenter | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
741 | OutPt InfusionCenter | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
742 | OutPt InfusionCenter | nivolumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | mg/mL |  | 1 | 10 | In-line filter
743 | OutPt InfusionCenter | ocrelizumab |  | 300 | mg | 260 | mL | 1.1538 | mg/mL | Pump | Primary and Secondary | 02:20 | 02:25 | 24:00 |  | mg | 290 | 610 | 611 | n/a |  |  |  | 0.22 Filter Needed
744 | OutPt InfusionCenter | ocrelizumab |  | 600 | mg | 520 | mL | 1.1538 | mg/mL | Pump | Primary and Secondary | 01:50 | 01:55 | 24:00 |  | mg | 290 | 610 | 611 | n/a |  |  |  | 0.22 Filter Needed
745 | OutPt InfusionCenter | ondansetron |  | 8 | mg | 54 | mL | 0.1481 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 4 | 16 |  | n/a
746 | OutPt InfusionCenter | ondansetron |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 |  | mg | 4 | 16 |  | mg/mL | 0.02 | 0.04 | 0.64
747 | OutPt InfusionCenter | oritavacin | Kimsyra gram | 1.2 | gram | 250 | mL | 0.0048 | gram/mL | Pump | Primary and Secondary | 00:55 |  | 01:15 | 01:00 | gram | 1.199 |  | 1.201 | n/a |  |  |  | Flush Before/After
748 | OutPt InfusionCenter | oritavacin | Kimyrsa mg | 1200 | mg | 250 | mL | 4.8 | mg/mL | Pump | Primary and Secondary | 00:55 |  | 01:15 | 01:00 | mg | 1199 |  | 1201 | n/a |  |  |  | Flush Before/After
749 | OutPt InfusionCenter | oritavacin | Orbactiv gram | 1.2 | gram | 250 | mL | 0.0048 | gram/mL | Pump | Primary and Secondary | 02:45 |  | 03:30 | 03:00 | gram | 1.199 |  | 1.201 | n/a |  |  |  | Flush Before/AfterD5
750 | OutPt InfusionCenter | oritavacin | Orbactiv mg | 1200 | mg | 1000 | mL | 1.2 | mg/mL | Pump | Primary and Secondary | 02:45 |  | 03:30 | 03:00 | mg | 1199 |  | 1201 | n/a |  |  |  | Flush Before/AfterD5
751 | OutPt InfusionCenter | piperacillin/tazo |  | 3.375 | gram | 50 | mL | 0.0675 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
752 | OutPt InfusionCenter | piperacillin/tazo |  | 3.375 | gram | 100 | mL | 0.0338 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
753 | OutPt InfusionCenter | piperacillin/tazo |  | 4.5 | gram | 50 | mL | 0.09 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
754 | OutPt InfusionCenter | piperacillin/tazo |  | 4.5 | gram | 100 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 3 |  | 4.6 | n/a
755 | OutPt InfusionCenter | potassium chloride |  | 10 | mEq | 100 | mL | 0.1 | mEq/mL | Pump | Primary and Secondary | 00:25 | 00:45 | 06:00 | 01:00 | mEq | 5 |  | 41 | n/a
756 | OutPt InfusionCenter | potassium chloride |  | 40 | mEq | 520 | mL | 0.0769 | mEq/mL | Pump | Primary and Secondary | 03:15 |  | 12:00 | 04:00 | mEq | 5 |  | 41 | n/a
757 | OutPt InfusionCenter | potassium chloride |  | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:50 | 01:45 | 12:00 |  | mEq | 5 |  | 41 | mEq/mL | 0.04 | 0.07 | 0.11
758 | OutPt InfusionCenter | promethazine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | mg | 6.25 | 25 |  | mg/mL | 1 | 1.25 | 25
759 | OutPt InfusionCenter | raNITIdine |  | 50 | mg | 102 | mL | 0.4902 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:20 | mg | 25 | 50 |  | n/a
760 | OutPt InfusionCenter | rifAMPin |  | 300 | mg | 250 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
761 | OutPt InfusionCenter | rifAMPin |  | 600 | mg | 500 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
762 | OutPt InfusionCenter | rifAMPin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | mg/mL | 0.1 | 0.5 | 6
763 | OutPt InfusionCenter | riTUXimab | Based on BSA | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg/m2 | 250 | 500 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
764 | OutPt InfusionCenter | riTUXimab | Rheumatoid Arthritis | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg | 90 | 1200 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
765 | OutPt InfusionCenter | spesolimab |  | 900 | mg | 100 | mL | 9 | mg/mL | Pump | Primary Only |  | 01:25 | 03:00 | 01:30 | mg | 899 |  | 901 | n/a |  |  |  | Low Binding Filter
766 | OutPt InfusionCenter | teprotumumab | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 | 00:55 | 08:00 |  | mg/kg | 7 | 21 | 22 | mg/mL | 2.9 | 3 | 18.1
767 | OutPt InfusionCenter | teprotumumab | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 | 00:55 | 08:00 |  | gram/kg | 0.007 | 0.021 | 0.022 | gram/mL | 0.002 | 0.003 | 0.019
768 | OutPt InfusionCenter | tocilizumab |  | 800 | mg | 100 | mL | 8 | mg/mL | Pump | Primary and Secondary | 00:55 |  | 04:00 | 01:00 | mg | 161 |  | 801 | n/a |  |  |  | Dedicated IV Line
769 | OutPt InfusionCenter | tocilizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:55 |  | 04:00 | 01:00 | mg/kg | 3.9 |  | 8.1 | mg/mL | 1 | 1.1 | 8.1 | Dedicated IV Line
770 | OutPt InfusionCenter | ustekinumab |  | 260 | mg | 250 | mL | 1.04 | mg/mL | Pump | Primary Only | 00:55 |  | 03:00 | 01:00 | mg | 259 |  | 521 | n/a |  |  |  | Low Binding Filter
771 | OutPt InfusionCenter | ustekinumab |  | 390 | mg | 250 | mL | 1.56 | mg/mL | Pump | Primary Only | 00:55 |  | 03:00 | 01:00 | mg | 259 |  | 521 | n/a |  |  |  | Low Binding Filter
772 | OutPt InfusionCenter | ustekinumab |  | 520 | mg | 250 | mL | 2.08 | mg/mL | Pump | Primary Only | 00:55 |  | 03:00 | 01:00 | mg | 259 |  | 521 | n/a |  |  |  | Low Binding Filter
773 | OutPt InfusionCenter | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 500 | 2000 |  | n/a
774 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
775 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
776 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
777 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
778 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | mg | 500 | 2000 |  | n/a
779 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg | 500 | 2000 |  | n/a
780 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg | 500 | 2000 |  | n/a
781 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
782 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
783 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
784 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
785 | OutPt InfusionCenter | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | mg | 500 | 2000 |  | mg/mL | 1 | 2.9 | 5.1
786 | OutPt InfusionCenter | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg/kg | 5 |  | 30 | n/a
787 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
788 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
789 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
790 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
791 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
792 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
793 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
794 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
795 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
796 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
797 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
798 | OutPt InfusionCenter | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 |  | mg/kg | 5 |  | 30 | mg/mL | 1 | 3 | 5
799 | OutPt InfusionCenter | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram | 0.5 | 2 |  | n/a
800 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
801 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
802 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
803 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
804 | OutPt InfusionCenter | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | gram | 0.5 | 2 |  | n/a
805 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram | 0.5 | 2 |  | n/a
806 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram | 0.5 | 2 |  | n/a
807 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
808 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
809 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
810 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
811 | OutPt InfusionCenter | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.002 | 0.006
812 | OutPt InfusionCenter | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | gram/kg | 0.005 |  | 0.03 | n/a
813 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
814 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
815 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
816 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
817 | OutPt InfusionCenter | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
818 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
819 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
820 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
821 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
822 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
823 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
824 | OutPt InfusionCenter | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | gram/kg | 0.005 |  | 0.03 | gram/mL | 0.001 | 0.002 | 0.006
825 | OutPt InfusionCenter | vedolizumab |  | 300 | mg | 250 | mL | 1.2 | mg/mL | Pump | Primary and Secondary | 00:28 |  | 02:00 | 00:30 | mg | 299 |  | 301 | n/a |  |  |  | Flush After NS-LR
826 | OutPt InfusionCenter | zoledronic acid |  | 4 | mg | 100 | mL | 0.04 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | n/a
827 | OutPt InfusionCenter | zoledronic acid |  | 5 | mg | 100 | mL | 0.05 | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | n/a
828 | OutPt InfusionCenter | zoledronic acid |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:14 |  | 02:00 | 00:15 | mg | 2.9 | 5.1 | 5.11 | mg/mL | 0.028 | 0.029 | 0.051
829 | Pediatrics | acetaminophen |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 999 | 1001 |  | n/a
830 | Pediatrics | acetaminophen |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg/kg | 7 |  | 16 | mg/mL | 9 |  | 11
831 | Pediatrics | acetylcysteine | 1st Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg/kg | 142.5 |  | 157.5 | mg/mL |  | 1 | 75
832 | Pediatrics | acetylcysteine | 1st Dose >100kg | 15000 | mg | 275 | mL | 54.5455 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 14999 |  | 15001 | n/a
833 | Pediatrics | acetylcysteine | 2nd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg/kg | 47.5 |  | 52.5 | mg/mL |  | 1 | 75
834 | Pediatrics | acetylcysteine | 2nd Dose >100kg | 5000 | mg | 525 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg | 4999 |  | 5001 | n/a
835 | Pediatrics | acetylcysteine | 3rd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg/kg | 95 |  | 105 | mg/mL |  | 1 | 75
836 | Pediatrics | acetylcysteine | 3rd Dose >100kg | 10000 | mg | 1050 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg | 9999 |  | 10001 | n/a
837 | Pediatrics | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 4 | 25 |  | mg/mL | 0.1 | 1 | 7
838 | Pediatrics | albumin 25% |  | 5 | gram | 20 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:08 | 01:00 | 00:20 | gram | 4.9 | 25.1 |  | n/a
839 | Pediatrics | albumin 25% |  | 12.5 | gram | 50 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:16 | 02:00 | 00:50 | gram | 4.9 | 25.1 |  | n/a
840 | Pediatrics | albumin 25% |  | 25 | gram | 100 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:32 | 04:00 | 01:40 | gram | 4.9 | 25.1 |  | n/a
841 | Pediatrics | ampicillin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:30 |  | mg/kg | 4 | 100 |  | mg/mL | 1 | 5 | 35
842 | Pediatrics | ampicillin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:30 | 00:30 | mg | 500 | 3000 |  | mg/mL | 1 | 5 | 35
843 | Pediatrics | ampicillin/sulbact | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 |  | mg/kg | 4 | 80 |  | mg/mL | 5 | 15 | 35
844 | Pediatrics | ampicillin/sulbact | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 |  | mg | 1000 | 3000 |  | mg/mL | 5 | 15 | 35
845 | Pediatrics | antivenin crotalidae |  | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 01:00 | 11:30 | 10:00 | unit | 1 | 6 |  | unit/mL | 0.004 |  | 0.024 | Antivenin
846 | Pediatrics | azithromycin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 3 | 14 |  | mg/mL | 0.1 | 1 | 2.2
847 | Pediatrics | azithromycin | pt weight >= 40 kg | 250 | mg | 250 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:40 | 02:00 | 01:00 | mg | 100 | 600 |  | n/a
848 | Pediatrics | azithromycin | pt weight >= 40 kg | 500 | mg | 250 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:40 | 02:00 | 01:00 | mg | 100 | 600 |  | n/a
849 | Pediatrics | aztreonam | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:15 | 02:00 |  | mg/kg | 25.5 | 50 |  | mg/mL | 4 | 5 | 40
850 | Pediatrics | aztreonam | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:15 | 02:00 |  | mg | 500 | 2000 |  | mg/mL | 4 | 5 | 40
851 | Pediatrics | caffeine citrate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 06:00 |  | mg | 1 | 1200 |  | mg/mL |  | 0.01 | 10
852 | Pediatrics | caffeine citrate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 06:00 |  | mg/kg | 1 | 22 |  | mg/mL |  | 0.01 | 10
853 | Pediatrics | caffeine/sod benzoa |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 04:00 | 01:00 | mg | 50 | 501 | 1001 | mg/mL |  | 0.24 | 6 | 0.22 Filter Needed
854 | Pediatrics | caffeine/sod benzoa |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 04:00 | 01:00 | mg/kg | 1 | 22 |  | mg/mL |  | 0.24 | 6 | 0.22 Filter Needed
855 | Pediatrics | ceFAZolin | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg/kg | 9 | 55 |  | mg/mL | 1 | 5 | 40
856 | Pediatrics | ceFAZolin | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
857 | Pediatrics | ceFAZolin | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram/kg | 0.009 | 0.055 |  | gram/mL | 0.001 | 0.005 | 0.04
858 | Pediatrics | ceFAZolin | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
859 | Pediatrics | cefepime | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 15 | 55 |  | mg/mL | 1 | 5 | 40
860 | Pediatrics | cefepime | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
861 | Pediatrics | cefepime | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | gram/kg | 0.015 | 0.055 |  | gram/mL | 0.001 | 0.005 | 0.04
862 | Pediatrics | cefepime | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
863 | Pediatrics | cefotaxime | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 22 | 80 |  | mg/mL | 5 | 10 | 55
864 | Pediatrics | cefotaxime | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 55
865 | Pediatrics | cefotaxime | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram/kg | 0.022 | 0.08 |  | gram/mL | 0.005 | 0.01 | 0.055
866 | Pediatrics | cefotaxime | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.055
867 | Pediatrics | cefOXitin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg/kg | 20 | 55 |  | mg/mL | 5 | 10 | 40
868 | Pediatrics | cefOXitin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 40
869 | Pediatrics | cefTAZidime | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | mg/kg | 25 | 67 | 135 | mg/mL | 1 | 5 | 50
870 | Pediatrics | cefTAZidime | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | mg | 250 | 2000 |  | mg/mL | 1 | 5 | 50
871 | Pediatrics | cefTAZidime | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | gram/kg | 0.025 | 0.067 | 0.135 | gram/mL | 0.001 | 0.005 | 0.05
872 | Pediatrics | cefTAZidime | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | gram | 0.25 | 2 |  | gram/mL | 0.001 | 0.005 | 0.05
873 | Pediatrics | cefTRIAXone | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg/kg | 21 | 105 |  | mg/mL | 5 | 10 | 50
874 | Pediatrics | cefTRIAXone | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 50
875 | Pediatrics | cefTRIAXone | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | gram/kg | 0.021 | 0.105 |  | gram/mL | 0.005 | 0.01 | 0.05
876 | Pediatrics | cefTRIAXone | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.05
877 | Pediatrics | cefuroxime | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg/kg | 18 | 55 |  | mg/mL | 5 | 7.5 | 30
878 | Pediatrics | cefuroxime | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg | 750 | 1500 |  | mg/mL | 5 | 7.5 | 30
879 | Pediatrics | chlorothiazide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 250 | 1000 |  | mg/mL | 0.5 | 1 | 30
880 | Pediatrics | chlorothiazide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg/kg | 0.4 | 11 |  | mg/mL | 0.5 | 1 | 30
881 | Pediatrics | clindamycin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 4 | 23 |  | mg/mL | 1 | 2 | 20
882 | Pediatrics | clindamycin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 300 | 2400 |  | mg/mL | 1 | 2 | 20
883 | Pediatrics | erythromycin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg/kg | 2.5 | 15 |  | mg/mL | 1 | 2 | 5.5
884 | Pediatrics | erythromycin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg | 100 | 1000 |  | mg/mL | 1 | 2 | 5.5
885 | Pediatrics | fluconazole | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 02:00 | 01:00 | mg/kg | 3 | 12 | 14 | mg/mL | 1.5 |  | 2.5
886 | Pediatrics | fluconazole | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg | 100 | 800 |  | mg/mL | 1.5 |  | 2.5
887 | Pediatrics | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:08 | 01:00 | 00:30 | mg | 100 | 1500 |  | mg/mL | 0.9 | 1 | 30
888 | Pediatrics | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:08 | 01:00 | 00:30 | mg/kg | 2 |  | 29 | mg/mL | 0.9 | 1 | 30
889 | Pediatrics | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:15 | 00:30 | mg/kg | 1 | 10 |  | mg/mL | 0.1 | 0.5 | 11
890 | Pediatrics | iron sucrose | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 01:35 |  | mg/kg | 0.4 | 7 | 7.3 | mg/mL | 0.9 | 1 | 2
891 | Pediatrics | iron sucrose | .mg, pt weight>=40kg | 100 | mg | 55 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 00:30 | mg | 90 |  | 315 | n/a
892 | Pediatrics | iron sucrose | .mg, pt weight>=40kg | 100 | mg | 105 | mL | 0.9524 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 00:30 | mg | 90 |  | 315 | n/a
893 | Pediatrics | iron sucrose | .mg, pt weight>=40kg | 200 | mg | 110 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary | 00:55 | 00:57 | 03:00 | 01:00 | mg | 90 |  | 315 | n/a
894 | Pediatrics | iron sucrose | .mg, pt weight>=40kg | 300 | mg | 265 | mL | 1.1321 | mg/mL | Pump | Primary and Secondary | 01:20 | 01:25 | 04:00 | 01:30 | mg | 90 |  | 315 | n/a
895 | Pediatrics | levETIRAcetam | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:08 | 00:10 | 01:00 |  | mg/kg | 4.3 | 57 | 115 | mg/mL | 0.5 | 1 | 30
896 | Pediatrics | levETIRAcetam | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:08 | 00:10 | 01:00 |  | mg | 200 | 3000 | 5000 | mg/mL | 0.5 | 1 | 30
897 | Pediatrics | mannitol 20% |  | 100 | gram | 500 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 00:31 | 02:00 |  | gram | 1 | 101 |  | n/a
898 | Pediatrics | mannitol 20% |  | 100 | gram | 500 | mL | 0.2 | gram/mL | Pump | Primary and Secondary | 00:25 | 00:31 | 02:00 |  | gram/kg | 0.2 | 2 |  | n/a |  |  |  | In-line filter
899 | Pediatrics | mannitol 20% |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:10 | 02:00 |  | gram/kg | 0.2 | 2 |  | gram/mL |  | 0.1 | 0.5 | In-line filter
900 | Pediatrics | meropenem | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | mg/kg | 10 | 40 |  | mg/mL | 1 | 5 | 40
901 | Pediatrics | meropenem | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
902 | Pediatrics | meropenem | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | gram/kg | 0.01 | 0.04 |  | gram/mL | 0.001 | 0.005 | 0.04
903 | Pediatrics | meropenem | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
904 | Pediatrics | metroNIDAZOLE |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 03:00 | 01:00 | mg | 250 | 500 |  | n/a
905 | Pediatrics | metroNIDAZOLE |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 03:00 | 01:00 | mg | 250 | 500 |  | n/a
906 | Pediatrics | metroNIDAZOLE |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 03:00 | 01:00 | mg/kg | 7 | 16 |  | mg/mL | 4 |  | 6
907 | Pediatrics | micafungin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 0.85 | 7 |  | mg/mL | 0.4 |  | 4
908 | Pediatrics | micafungin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg | 40 | 150 |  | mg/mL | 0.4 |  | 4
909 | Pediatrics | nafcillin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 10.6 | 57.5 |  | mg/mL | 1 | 5 | 40
910 | Pediatrics | nafcillin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
911 | Pediatrics | oxacillin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:08 | 01:30 |  | mg/kg | 5 | 50 |  | mg/mL | 5 | 10 | 40
912 | Pediatrics | oxacillin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:08 | 01:30 |  | mg | 500 | 2000 |  | mg/mL | 5 | 10 | 40
913 | Pediatrics | penG x Million Units |  | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | unit | 1 | 6 |  | unit/mL | 0.01 | 0.02 | 0.12 | 1unit=1million
914 | Pediatrics | pentamidine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 04:30 | 02:00 | mg/kg | 2.9 |  | 4.1 | mg/mL |  | 0.6 | 30
915 | Pediatrics | PHENobarbital |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:08 | 02:00 |  | mg/kg | 0.5 | 7.5 |  | mg/mL | 0.1 | 0.5 | 135 | High Risk
916 | Pediatrics | phenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:15 | 02:00 |  | mg/kg | 1 |  | 29 | mg/mL | 1 | 4 | 55 | In-line filter
917 | Pediatrics | piperacillin/tazo | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | mg/kg | 20 | 140 |  | mg/mL | 0.5 | 10 | 100
918 | Pediatrics | piperacillin/tazo | GM dosing | 2.25 | gram | 50 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
919 | Pediatrics | piperacillin/tazo | GM dosing | 2.25 | gram | 100 | mL | 0.0225 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
920 | Pediatrics | piperacillin/tazo | GM dosing | 3.375 | gram | 50 | mL | 0.0675 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
921 | Pediatrics | piperacillin/tazo | GM dosing | 3.375 | gram | 100 | mL | 0.0338 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
922 | Pediatrics | piperacillin/tazo | GM dosing | 4.5 | gram | 50 | mL | 0.09 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
923 | Pediatrics | piperacillin/tazo | GM dosing | 4.5 | gram | 100 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
924 | Pediatrics | piperacillin/tazo | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 04:30 |  | gram/kg | 0.025 | 0.14 |  | gram/mL | 0.005 | 0.01 | 0.1
925 | Pediatrics | potassium chloride |  | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:45 | 00:50 | 06:00 |  | mEq | 1 |  | 41 | mEq/mL | 0.04 | 0.05 | 0.2
926 | Pediatrics | potassium chloride |  | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:45 | 00:50 | 06:00 |  | mEq/kg | 0.2 | 1.5 | 2.5 | mEq/mL | 0.04 | 0.05 | 0.2 | High Risk
927 | Pediatrics | potassium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 01:30 | 03:00 | 12:30 |  | mmol | 5 | 20 | 40 | mmol/mL | 0.009 | 0.01 | 0.3
928 | Pediatrics | raNITIdine | pt weight < 50 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:10 | 01:00 | 00:30 | mg/kg | 0.4 | 2.5 |  | mg/mL | 0.1 | 0.5 | 3
929 | Pediatrics | raNITIdine | pt weight >= 50 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:10 | 01:00 |  | mg | 25 | 50 | 150 | mg/mL | 0.1 | 0.5 | 3
930 | Pediatrics | rifAMPin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 03:00 |  | mg/kg | 4.3 | 23 |  | mg/mL | 0.5 | 1 | 7
931 | Pediatrics | rifAMPin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 03:00 |  | mg | 300 | 600 |  | mg/mL | 0.5 | 1 | 7
932 | Pediatrics | SMX/TMP |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:30 | mg/kg | 1.7 | 8 | 10 | mg/mL | 0.01 | 0.04 | 1 | 48 Hour Tubing
933 | Pediatrics | sodium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 00:45 | 01:00 | 10:30 |  | mmol | 1 |  | 31 | mmol/mL | 0.009 | 0.01 | 0.3 | Calcium Precipitate
934 | Pediatrics | sodium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 00:45 | 01:00 | 10:30 |  | mmol/kg | 0.04 | 0.4 | 0.6 | mmol/mL | 0.009 | 0.01 | 0.3 | Calcium Precipitate
935 | Pediatrics | tobramycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:15 |  | mg/kg | 0.8 | 7 | 12 | mg/mL | 0.01 | 0.1 | 11
936 | Pediatrics | tocilizumab |  | 800 | mg | 100 | mL | 8 | mg/mL | Pump | Primary and Secondary | 00:55 |  | 04:00 | 01:00 | mg | 161 |  | 801 | n/a
937 | Pediatrics | tocilizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 00:55 |  | 04:00 | 01:00 | mg/kg | 3.9 |  | 12.1 | mg/mL | 1 | 1.1 | 8.1
938 | Pediatrics | valproate sodium |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 |  | mg/kg | 5 | 60 |  | mg/mL | 1 |  | 25
939 | Pediatrics | vancomycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:30 | 00:50 | 04:30 |  | mg/kg | 5 | 30 |  | mg/mL | 0.5 | 1 | 5
940 | Pediatrics | voriconazole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:10 |  | mg/kg | 2.6 | 9 |  | mg/mL | 0.1 | 0.4 | 5.5
941 | Pediatrics | zidovudine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 0.8 | 9 |  | mg/mL | 0.5 | 1 | 4
Note on Y26: Existing IV lines must be flushed with D5w prior to infusion. If unable, must start a new line. **An in-fine filter >= 1 micron may be used with infusion.
Note on Y46: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y47: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y48: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y49: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y50: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y51: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y71: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y73: Use a 0.22 micron in-line filter when administering this product.
Note on Y74: Use a 0.22 micron in-line filter when administering this product.
Note on Y93: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y94: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y95: Do not infuse any other drugs through this line.
Note on Y96: Do not infuse any other drugs through this line.
Note on Y97: Do not infuse any other drugs through this line.
Note on Y200: Advise patient to lie down during and 30 minutes after infusion to prevent hypotension.
Note on Y212: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y213: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y214: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y215: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y216: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y217: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y218: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y219: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y220: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y221: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y222: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y223: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y236: Protect this drug from light.
Note on Y237: Protect this drug from light.
Note on Y254: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y255: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y256: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y257: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y258: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y259: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y260: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y261: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y270: Infusion requires the use of an in-line filter.
Note on Y284: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y285: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y286: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y287: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y288: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y289: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y290: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y291: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y292: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y293: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y294: Infusion requires the use of an in-line filter.
Note on Y295: Infusion requires the use of an in-line filter.
Note on Y296: INR 2 - 4 ** Max dose of Kcentra should not exceed 2500 units **
Note on Y297: INR 4 - 6 ** Max dose of Kcentra should not exceed 3500 units **
Note on Y298: INR > 6 ** Max dose of Kcentra should not exceed 5000 units **
Note on Y311: Infusion requires the use of an in-line filter.
Note on Y312: Infusion requires the use of an in-line filter.
Note on Y313: Infusion requires the use of an in-line filter.
Note on Y359: Do not infuse any other drugs through this line.
Note on Y361: 1 unit = 1 million units
Note on Y362: 1 unit = 1 million units
Note on Y363: 1 unit = 1 million units
Note on Y364: 1 unit = 1 million units
Note on Y365: 1 unit = 1 million units
Note on Y366: 1 unit = 1 million units
Note on Y372: Infusion requires the use of an in-line filter.
Note on Y407: Premedicate before each infusion.
Note on Y408: Change tubing every 48 hours.
Note on Y410: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y411: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y422: Do not infuse any other drugs through this line.
Note on Y423: Do not infuse any other drugs through this line.
Note on Y489: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y490: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y491: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y492: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y493: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y494: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y499: Adminster over 60 minutes to decrease risk of bilirubin encephalopathy.
Note on Y503: 1 unit in pump = 1,000 units on label. Eg. 200 units in pump = 200,000 units on drug label.
Note on Y506: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y507: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y508: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y509: This medication must be infused using DEHP FREE tubing.
Note on Y510: This medication must be infused using DEHP FREE tubing.
Note on Y521: MAXIMUM dose of brentuximab is 180 mg.
Note on Y522: Use a 0.22 micron in-line filter when administering this product.
Note on Y526: Infusion requires the use of an in-line filter.
Note on Y527: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y528: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y530: Use a 0.22 micron in-line filter when administering this product.
Note on Y534: Use a 0.22 micron in-line filter when administering this product.
Note on Y535: Use a 0.22 micron in-line filter when administering this product.
Note on Y536: Use a 0.22 micron in-line filter when administering this product.
Note on Y539: This medication must be infused using DEHP FREE tubing.
Note on Y541: Do NOT use an in-line filter for administration.
Note on Y542: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y543: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y544: Do NOT administer through lines containing dextrose.
Note on Y545: This medication must be infused using Non-PVC tubing.
Note on Y553: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y557: Do NOT use an in-line filter for administration.
Note on Y566: Infusion requires the use of an in-line filter.
Note on Y567: Infusion requires the use of an in-line filter.
Note on Y568: Infusion requires the use of an in-line filter.
Note on Y569: Infusion requires the use of an in-line filter.
Note on Y570: Infusion requires the use of an in-line filter.
Note on Y578: Requires Low protein binding 0.2 micron filter and DEHP free bag
Note on Y579: Do NOT use an in-line filter for administration.
Note on Y580: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y581: Infusion requires the use of an in-line filter.
Note on Y582: Infusion requires the use of an in-line filter.
Note on Y586: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y587: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y589: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y590: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y591: Flush line before and after administration. Can be given on same line as sodium chloride 0.9%.
Note on Y593: Use a 0.22 micron in-line filter when administering this product.
Note on Y596: Drug is a vesicant. Check blood return.
Note on Y597: Drug is a vesicant. Check blood return.
Note on Y599: Use a 0.22 micron in-line filter when administering this product.
Note on Y601: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y602: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y603: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y604: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y609: Existing IV lines must be flushed with D5w prior to infusion. If unable, must start a new line. **An in-fine filter >= 1 micron may be used with infusion.
Note on Y611: Use a low-protein binding in-line or add-on filter between 0.2 and 5 microns.
Note on Y664: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y665: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y666: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y667: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y668: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y669: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y670: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y671: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y683: Protect this drug from light.
Note on Y684: Protect this drug from light.
Note on Y699: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y700: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y701: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y702: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y705: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y706: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y707: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y708: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y709: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y710: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y711: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y712: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y713: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y714: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y740: Infusion requires the use of an in-line filter.
Note on Y741: Infusion requires the use of an in-line filter.
Note on Y742: Infusion requires the use of an in-line filter.
Note on Y743: Infusion requires the use of an in-line filter.
Note on Y744: Infusion requires the use of an in-line filter.
Note on Y745: Use a 0.22 micron in-line filter when administering this product.
Note on Y746: Use a 0.22 micron in-line filter when administering this product.
Note on Y749: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w or NS.
Note on Y750: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w or NS.
Note on Y751: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y752: If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
Note on Y765: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y766: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y767: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y770: Do not infuse any other drugs through this line.
Note on Y771: Do not infuse any other drugs through this line.
Note on Y772: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y773: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y774: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y827: Following infusion, flush with 30 mL of sterile normal saline or lactated ringers.
Note on Y847: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y855: Use a 0.22 micron in-line filter when administering this product.
Note on Y856: Use a 0.22 micron in-line filter when administering this product.
Note on Y900: Infusion requires the use of an in-line filter.
Note on Y901: Infusion requires the use of an in-line filter.
Note on Y915: 1 unit = 1 million units
Note on Y917: This is a HIGH RISK medication. 2nd RN to verify prior to initiating therapy.
Note on Y918: Infusion requires the use of an in-line filter.
Note on Y928: This is a HIGH RISK medication. 2nd RN to verify prior to initiating therapy.
Note on Y934: Change tubing every 48 hours.
Note on Y935: Do not infuse Calcium and Phosphate though the same line.
Note on Y936: Do not infuse Calcium and Phosphate though the same line.

## Fluids
Fluids
Line Item | Profile | Fluid Name | Therapy Name | Module Availability | Supports Secondary (Y/N) | Rate Units | Rate Limits (mL/h) Soft Min | Rate Limits (mL/h) Soft Max | Rate Limits (mL/h) Hard Max | Clinical Advisory
1 | Adult | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 999
2 | Adult | .IVF + KCL | 10 mEq/L | Pump | Y | mL/h | 1 | 999
3 | Adult | .IVF + KCL | 15 mEq/L | Pump | Y | mL/h | 1 | 666 | 700
4 | Adult | .IVF + KCL | 20 mEq/L | Pump | Y | mL/h | 1 | 500 | 550
5 | Adult | .IVF + KCL | 30 mEq/L | Pump | Y | mL/h | 1 | 333 | 366
6 | Adult | .IVF + KCL | 40 mEq/L | Pump | Y | mL/h | 1 | 250 | 275
7 | Adult | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 1 | 999
8 | Adult | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 500 | 550
9 | Adult | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 1 | 250 | 275
10 | Adult | .Maint IVF+Additives |  | Pump | Y | mL/h | 1 |  | 999
11 | Adult | 3% Sodium Chloride |  | Pump | Y | mL/h | 1 | 50 | 100
12 | Adult | Bicarb Drip |  | Pump | Y | mL/h | 1 |  | 999
13 | Adult | Blood (whole) |  | Pump | N | mL/h | 1 | 301 |  | Dedicated IV Line
14 | Adult | Dextrose 10% |  | Pump | Y | mL/h | 1 | 150
15 | Adult | Fresh Frozen Plasma |  | Pump | N | mL/h | 1 | 201 |  | Dedicated IV Line
16 | Adult | Insulin Maternity |  | Pump | Y | mL/h | 70 | 160 | 200
17 | Adult | Parenteral Nutrition |  | Pump | N | mL/h | 19 | 95 |  | Midline Warning
18 | Adult | PRBCs |  | Pump | N | mL/h | 1 | 301 |  | Dedicated IV Line
19 | Nursery | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 15 | 25
20 | Nursery | Dextrose 10% |  | Pump | Y | mL/h | 1 | 15 | 25
21 | ONC Infusion Center | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 999
22 | ONC Infusion Center | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 1 | 999
23 | ONC Infusion Center | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 667
24 | ONC Infusion Center | .IVF + KCL 30 mEq/L |  | Pump | Y | mL/h | 1 | 334 | 667
25 | ONC Infusion Center | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 1 | 250 | 500
26 | ONC Infusion Center | .Maint IVF+Additives |  | Pump | Y | mL/h | 1 | 250
27 | ONC Infusion Center | Dextrose 10% |  | Pump | Y | mL/h | 15 | 200
28 | ONC Infusion Center | Rituximab (Rituxan) |  | Pump | Y | mL/h | 0.1 | 250
29 | OutPt InfusionCenter | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 999
30 | OutPt InfusionCenter | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 1 | 999
31 | OutPt InfusionCenter | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 667
32 | OutPt InfusionCenter | .IVF + KCL 30 mEq/L |  | Pump | Y | mL/h | 1 | 334 | 667
33 | OutPt InfusionCenter | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 1 | 250 | 500
34 | OutPt InfusionCenter | .Maint IVF+Additives |  | Pump | Y | mL/h | 1 | 250
35 | OutPt InfusionCenter | Dextrose 10% |  | Pump | Y | mL/h | 15 | 200
36 | OutPt InfusionCenter | PRBCs |  | Pump and Syringe | N | mL/h | 1 | 150 |  | Dedicated IV Line
37 | OutPt InfusionCenter | Rituximab (Rituxan) |  | Pump | Y | mL/h | 0.1 | 250
38 | Pediatrics | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 300
39 | Pediatrics | .IVF + KCL | 10 mEq/L | Pump | Y | mL/h | 1 | 999
40 | Pediatrics | .IVF + KCL | 15 mEq/L | Pump | Y | mL/h | 1 | 666 | 700
41 | Pediatrics | .IVF + KCL | 20 mEq/L | Pump | Y | mL/h | 1 | 500 | 550
42 | Pediatrics | .IVF + KCL | 30 mEq/L | Pump | Y | mL/h | 1 | 333 | 366
43 | Pediatrics | .IVF + KCL | 40 mEq/L | Pump | Y | mL/h | 1 | 250 | 275
44 | Pediatrics | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 0.1 | 200 | 999
45 | Pediatrics | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 200 | 500
46 | Pediatrics | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 0.1 | 200 | 275
47 | Pediatrics | Cryoprecipitate |  | Pump | N | mL/h | 0.1 | 300
48 | Pediatrics | Dextrose 10% |  | Pump | Y | mL/h | 1 | 50 | 150
49 | Pediatrics | Fresh Frozen Plasma |  | Pump | N | mL/h | 0.1 | 150 |  | Dedicated IV Line
50 | Pediatrics | Parenteral Nutrition |  | Pump | N | mL/h | 35 | 95 |  | Midline Warning
51 | Pediatrics | PRBCs |  | Pump | N | mL/h | 0.1 | 150 |  | Dedicated IV Line
Note on K15: Do not infuse any other drugs through this line.
Note on K17: Do not infuse any other drugs through this line.
Note on K19: Do NOT infuse through a midline catheter.
Note on K20: Do not infuse any other drugs through this line.
Note on K38: Do not infuse any other drugs through this line.
Note on K51: Do not infuse any other drugs through this line.
Note on K52: Do NOT infuse through a midline catheter.
Note on K53: Do not infuse any other drugs through this line.

## PCA Drugs
PCA Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount | Drug Amount Units | Diluent Volume | Diluent Volume Units | Final Concentration | Final Concentration Units | Max Accumulated Lockout | Dosing Units | Max Accumulated Dose Range Soft Min | Max Accumulated Dose Range Soft Max | Max Accumulated Dose Range Hard Max | Max Accumulated Dose Range Initial Value | PCA Dose Limits Soft Min | PCA Dose Limits Soft Max | PCA Dose Limits Hard Max | PCA Dose Limits Initial Value | Lock Out Interval (minutes) Hard Min | Lock Out Interval (minutes) Soft Min | Lock Out Interval (minutes) Soft Max | Lock Out Interval (minutes) Initial Value | Continuous Dose Limits Soft Min | Continuous Dose Limits Soft Max | Continuous Dose Limits Hard Max | Continuous Dose Limits Initial Value | Bolus Dose Limits Soft Min | Bolus Dose Limits Soft Max | Bolus Dose Limits Hard Max | Bolus Dose included in Accum Max Dose (Y/N) | Loading Dose Limits Soft Min | Loading Dose Limits Soft Max | Loading Dose Limits Hard Max | Conc. Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max | Supports PCA Pause Protocol (Y/N) | Clinical Advisory
1 | Adult | No drugs exist in this section
2 | Nursery | No drugs exist in this section
3 | ONC Infusion Center | No drugs exist in this section
4 | OutPt InfusionCenter | No drugs exist in this section
5 | Pediatrics | No drugs exist in this section

## PC Unit Config
| PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit
Line Item | Profile | Alarm Audio Profile | Minimum Audio Volume | Default Audio Volume | Anesthesia Mode | Battery Meter Display | Color Display | Key Click Audio | Limit Checking | Max Patient BSA (m2) | Max Patient Weight (kg) | Patient ID Entry | Preventive Maintenance Reminder | Tamper Resistant | Authorized User Mode | No Guardrails Basic Infusion Clinical Advisory | Clinical Advisory Message
1 | Adult | Profile 1 | 3 | 4 | Enabled | Enabled | Enabled | Enabled | Always | 3 | 500 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
2 | Nursery | Profile 1 | 2 | 3 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 7 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
3 | ONC Infusion Center | Profile 1 | 3 | 4 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 500 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
4 | OutPt InfusionCenter | Profile 1 | 3 | 4 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 500 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
5 | Pediatrics | Profile 1 | 3 | 4 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 250 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection

## Pump Config
| Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module
Line Item | Profile | Accumulated Air-In-Line | Air In Line (microliter) | Auto-Restart Attempts | KVO Rate Adjust (mL/h) | Max Occlusion Pressure (mmHg) | Default Pressure (mmHg) | Maximum Rate (mL/h) | Maximum VTBI (mL) | Pressure Mode Lock Status | Pressure Mode Selection | Secondary Infusion | SEC-PRI Audio Alert
1 | Adult | Enabled | 250 | 9 | 20 | 525 | 525 | 999 | 3000 | Unlocked | Pump | Enabled | Enabled
2 | Nursery | Enabled | 50 | 1 | 2 | 525 | 525 | 25 | 1200 | Unlocked | Pump | Enabled | Enabled
3 | ONC Infusion Center | Enabled | 250 | 9 | 20 | 525 | 525 | 999 | 3000 | Unlocked | Pump | Enabled | Enabled
4 | OutPt InfusionCenter | Enabled | 250 | 9 | 20 | 525 | 525 | 999 | 3000 | Unlocked | Pump | Enabled | Enabled
5 | Pediatrics | Enabled | 75 | 9 | 5 | 525 | 525 | 999 | 2200 | Unlocked | Pump | Enabled | Enabled

## Syringe Config
| Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module
Line Item | Profile | All Mode | Auto Occlusion Pressure | Back-Off After Occlusion | Fast-Start | KVO Option | KVO Rate Adjust (mL/h) | KVO Volume Adjust (%) | Near End of Infusion (Continuous, Fluids, Basic) | Alert Time (Continuous, Fluids, Basic) (minutes) | Near End of Infusion (Intermittent) | Alert Time (Intermittent) (minutes) | Occlusion Pressure Limit with Disc (mmHg) | Occlusion Pressure Limit No Disc | Priming
1 | Adult | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled
2 | Nursery | Enabled | Enabled | Enabled | Enabled | Enabled | 2 | 5 | Enabled | 2 | Enabled | 2 | 1000 | High | Enabled
3 | ONC Infusion Center | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled
4 | OutPt InfusionCenter | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled
5 | Pediatrics | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled

## Pump & Syringe Shared Config
| Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe)
Line Item | Profile | Delay Start Options | Delay Start Callback | Drug Calculation | Drug Calculation Bolus Dose | Pressure Dynamic Display | Volume/Duration
1 | Adult | Disabled | n/a | Enabled | Enabled | Enabled | Enabled
2 | Nursery | Disabled | n/a | Enabled | Enabled | Enabled | Enabled
3 | ONC Infusion Center | Disabled | n/a | Disabled | n/a | Enabled | Enabled
4 | OutPt InfusionCenter | Disabled | n/a | Disabled | n/a | Enabled | Enabled
5 | Pediatrics | Disabled | n/a | Enabled | Enabled | Enabled | Enabled

## PCA Config
| PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module
Line Item | Profile | Bolus Delivery Rate (mL/h) | Dose Request Cord (Profile) | Max Dose Limit | Time Window (h) | Module Location Enforcement | Near End of Infusion | Near End of Infusion Alert Time (%) | Pressure Limit | Priming | Security Access Level | Security Code
1 | Adult | 150 | 2 | Enabled | 1 | Enabled | Enabled | 5 | High | Enabled | 1 | n/a
2 | Nursery | 150 | 2 | Enabled | 4 | Enabled | Disabled | n/a | High | Enabled | 1 | n/a
3 | ONC Infusion Center | 150 | 2 | Enabled | 4 | Enabled | Disabled | n/a | High | Enabled | 1 | n/a
4 | OutPt InfusionCenter | 150 | 2 | Enabled | 4 | Enabled | Disabled | n/a | High | Enabled | 1 | n/a
5 | Pediatrics | 150 | 2 | Enabled | 1 | Enabled | Enabled | 5 | High | Enabled | 1 | n/a

## PCA Pause Protocol
| PCA Pause Protocol | PCA Pause Protocol | PCA Pause Protocol | PCA Pause Protocol | PCA Pause Protocol
Line Item | Profile | PCA Pause Protocol Enabled | Available For | Monitoring Module Attach Enforcement | PCA Pause Protocol
1 | Adult | Enabled | Etco2 | Disabled | PCA Infusion has paused due to a decline in respiratory status. Check Patient!
2 | Nursery | Disabled | n/a | n/a | n/a
3 | ONC Infusion Center | Disabled | n/a | n/a | n/a
4 | OutPt InfusionCenter | Disabled | n/a | n/a | n/a
5 | Pediatrics | Disabled | n/a | n/a | n/a

## PCA & Syringe Shared Config
| Shared Infusion Settings (PCA & Syringe) | Shared Infusion Settings (PCA & Syringe)
Line Item | Profile | Maximum Rate (mL/h) | NEOI Snooze | NEOI Snooze Time (minutes)
1 | Adult | 999 | Disabled | n/a
2 | Nursery | 25 | Enabled | 5
3 | ONC Infusion Center | 999 | Disabled | n/a
4 | OutPt InfusionCenter | 999 | Disabled | n/a
5 | Pediatrics | 999 | Disabled | n/a

## EtCO2 Config
| EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module
Line Item | Profile | Alarm Limit Type | EtCO2 High (mmHg) | EtCO2 Low (mmHg) | Respiratory Rate High (bpm) | Respiratory Rate Low (bpm) | No Breath Alarm (Seconds) | FiCO2 High (mmHg) | Waveform Time Scale (Seconds) | PCA Pause Resp. Rate Lower Limit (bpm) | PCA Pause Resp. Rate Initial Value (bpm)
1 | Adult | Adult | 60 | 10 | 35 | 7 | 30 | 8 | 10 | 5 | 6
2 | Nursery | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a
3 | ONC Infusion Center | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a
4 | OutPt InfusionCenter | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a
5 | Pediatrics | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a

## Master Drugs
Master Drugs | Master Drugs | Master Drugs | Master Drugs | Master Drugs
Line Item | Drug Name | Concentration | Used As | Aliases
1 | abatacept | 500 mg / 100 mL (5 mg / mL) | Intermittent | 4144595I
2 | abatacept | 750 mg / 100 mL (7.5 mg / mL) | Intermittent | 4144595I
3 | abatacept | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 4144595I
4 | abatacept | --- mg / --- mL | Intermittent | 4144595I
5 | abciximab | 7.2 mg / 250 mL (0.0288 mg / mL) |  | 2748001I
6 | abciximab | 7.4 mg / 250 mL (0.0296 mg / mL) |  | 2748001I
7 | abciximab | 9 mg / 250 mL (0.036 mg / mL) |  | 2748001I
8 | abciximab | --- mg / --- mL |  | 2748001I
9 | acetaminophen | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2748022I
10 | acetaminophen | --- mg / --- mL | Intermittent | 2748022I
11 | acetaminophen | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2748022I
12 | acetaZOLAMIDE | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2748400I
13 | acetaZOLAMIDE | --- mg / --- mL | Intermittent | 2748400I
14 | acetylcysteine | 5,000 mg / 525 mL (9.5238 mg / mL) | Intermittent | 2748462I
15 | acetylcysteine | 10,000 mg / 1,050 mL (9.5238 mg / mL) | Intermittent | 2748462I
16 | acetylcysteine | 15,000 mg / 275 mL (54.5455 mg / mL) | Intermittent | 2748462I
17 | acetylcysteine | --- mg / --- mL | Intermittent | 2748462I
18 | acyclovir | --- mg / --- mL | Intermittent | 2748487I
19 | adenosine (CATH LAB) | 60 mg / 50 mL (1.2 mg / mL) | Cont/Bolus or PCA | 2748529C
20 | adenosine (CATH LAB) | 90 mg / 90 mL (1 mg / mL) | Cont/Bolus or PCA | 2748529C
21 | adenosine (CATH LAB) | 180 mg / 90 mL (2 mg / mL) | Cont/Bolus or PCA | 2748529C
22 | ado-trastuzumab | --- mg / --- mL |  | 24501603I
23 | aflibercept | --- mg / --- mL |  | 19949112I
24 | agalsidase beta | --- mg / --- mL |  | 2748541I
25 | albumin 25% | 5 gram / 20 mL (0.25 gram / mL) | Intermittent | 2748593I
26 | albumin 25% | 12.5 gram / 50 mL (0.25 gram / mL) | Intermittent | 2748593I
27 | albumin 25% | 25 gram / 100 mL (0.25 gram / mL) | Intermittent | 2748593I
28 | albumin 25% | --- gram / --- mL | Intermittent | 2748593I
29 | albumin 5% | 12.5 gram / 250 mL (0.05 gram / mL) |  | 2748593
30 | albumin 5% | 25 gram / 500 mL (0.05 gram / mL) |  | 2748593
31 | aldesleukin | --- unit / --- mL |  | 2748678I
32 | alemtuzumab | --- mg / --- mL |  | 2748690I
33 | alfentanil | --- mcg / --- mL |  | 2748710I
34 | allopurinol | --- mg / --- mL |  | 2748758I
35 | alprostadil | 500 mcg / 250 mL (2 mcg / mL) | Cont/Bolus or PCA | 2748821C
36 | alprostadil | --- mcg / --- mL | Cont/Bolus or PCA | 2748821C
37 | alteplase | 100 mg / 100 mL (1 mg / mL) |  | 2748851
38 | alteplase | --- mg / --- mL |  | 2748851
39 | alteplase -stroke | 81 mg / 81 mL (1 mg / mL) | Cont/Bolus or PCA | 2748851C
40 | alteplase -stroke | 90 mg / 90 mL (1 mg / mL) | Cont/Bolus or PCA | 2748851C
41 | alteplase -stroke | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2748851C
42 | alteplase -stroke | --- mg / --- mL | Cont/Bolus or PCA | 2748851C
43 | alteplase-non-stroke | 2 mg / 50 mL (0.04 mg / mL) | Intermittent | 2748851I
44 | alteplase-non-stroke | 2 mg / 100 mL (0.02 mg / mL) | Intermittent | 2748851I
45 | alteplase-non-stroke | 100 mg / 100 mL (1 mg / mL) | Intermittent | 2748851I
46 | alteplase-non-stroke | --- mg / --- mL | Intermittent | 2748851I
47 | amifostine | --- mg / --- mL |  | 2748941I
48 | amikacin | 250 mg / 250 mL (1 mg / mL) | Intermittent | 2748946I
49 | amikacin | --- mg / --- mL | Intermittent | 2748946I
50 | aminocaproic acid | --- mg / --- mL | Intermittent | 2748973I
51 | aminocaproic acid | 5 gram / 250 mL (0.02 gram / mL) | Intermittent | 2748973I
52 | aminocaproic acid | --- gram / --- mL | Intermittent | 2748973I
53 | aminophylline | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2748995C
54 | aminophylline | 500 mg / 500 mL (1 mg / mL) | Cont/Bolus or PCA | 2748995C
55 | aminophylline | 1,000 mg / 500 mL (2 mg / mL) | Cont/Bolus or PCA | 2748995C
56 | aminophylline | --- mg / --- mL | Cont/Bolus or PCA | 2748995C
57 | amiodarone | 150 mg / 100 mL (1.5 mg / mL) |  | 2749009
58 | amiodarone | 450 mg / 250 mL (1.8 mg / mL) |  | 2749009
59 | amiodarone | 900 mg / 500 mL (1.8 mg / mL) |  | 2749009
60 | amiodarone *Loading* | 150 mg / 100 mL (1.5 mg / mL) | Intermittent | 2749009I
61 | amiodarone *Loading* | 150 mg / 103 mL (1.4563 mg / mL) | Intermittent | 2749009I
62 | amiodarone Drip | 360 mg / 200 mL (1.8 mg / mL) | Cont/Bolus or PCA | 2749009C
63 | amiodarone Drip | 450 mg / 250 mL (1.8 mg / mL) | Cont/Bolus or PCA | 2749009C
64 | amiodarone Drip | 900 mg / 500 mL (1.8 mg / mL) | Cont/Bolus or PCA | 2749009C
65 | ammonium chloride | --- mEq / --- mL |  | 2749089I
66 | ampho B (LIPID) | --- mg / --- mL |  | 2749273I
67 | ampho B (liposomal) | --- mg / --- mL | Intermittent | 2749278I
68 | ampho B cholesteryl | --- mg / --- mL |  | 2749266I
69 | amphotericin B | --- mg / --- mL | Intermittent | 2749258I
70 | ampicillin | 250 mg / 50 mL (5 mg / mL) | Intermittent | 2749283I
71 | ampicillin | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2749283I
72 | ampicillin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2749283I
73 | ampicillin | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2749283I
74 | ampicillin | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2749283I
75 | ampicillin | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2749283I
76 | ampicillin | --- mg / --- mL | Intermittent | 2749283I
77 | ampicillin | 0.5 gram / 50 mL (0.01 gram / mL) | Intermittent | 2749283I
78 | ampicillin | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2749283I
79 | ampicillin | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2749283I
80 | ampicillin | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2749283I
81 | ampicillin | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2749283I
82 | ampicillin | --- gram / --- mL | Intermittent | 2749283I
83 | ampicillin/sulbact | 1,500 mg / 50 mL (30 mg / mL) | Intermittent | 2749313I
84 | ampicillin/sulbact | 1,500 mg / 100 mL (15 mg / mL) | Intermittent | 2749313I
85 | ampicillin/sulbact | 3,000 mg / 100 mL (30 mg / mL) | Intermittent | 2749313I
86 | ampicillin/sulbact | --- mg / --- mL | Intermittent | 2749313I
87 | ampicillin/sulbact | 1.5 gram / 50 mL (0.03 gram / mL) | Intermittent | 2749313I
88 | ampicillin/sulbact | 1.5 gram / 100 mL (0.015 gram / mL) | Intermittent | 2749313I
89 | ampicillin/sulbact | 3 gram / 100 mL (0.03 gram / mL) | Intermittent | 2749313I
90 | andulafungin | --- mg / --- mL | Intermittent | 4144679I
91 | antithymocyte EQUIN | --- mg / --- mL |  | 19949425I
92 | antithymocyte RABBI | --- mg / --- mL |  | 19949425
93 | antivenin crotalidae | 4 unit / 250 mL (0.016 unit / mL) | Intermittent | 2749465I
94 | antivenin crotalidae | 6 unit / 250 mL (0.024 unit / mL) | Intermittent | 2749465I
95 | antivenin crotalidae | 10 unit / 250 mL (0.04 unit / mL) | Intermittent | 2749465I
96 | antivenin crotalidae | --- unit / --- mL | Intermittent | 2749465I
97 | aprepitant | 100 mg / 100 mL (1 mg / mL) | Intermittent | 2749846I
98 | aprepitant | 130 mg / 100 mL (1.3 mg / mL) | Intermittent | 2749846I
99 | argatroban | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2749866C
100 | argatroban | 500 mg / 500 mL (1 mg / mL) | Cont/Bolus or PCA | 2749866C
101 | arsenic trioxide | --- mg / --- mL |  | 2749895I
102 | asparaginase | --- unit / --- mL |  | 2750061I
103 | atezolizumab | 800 mg / 250 mL (3.2 mg / mL) | Intermittent | 2554527703I
104 | atezolizumab | 840 mg / 250 mL (3.36 mg / mL) | Intermittent | 2554527703I
105 | atezolizumab | 1,200 mg / 250 mL (4.8 mg / mL) | Intermittent | 2554527703I
106 | atezolizumab | 1,680 mg / 250 mL (6.72 mg / mL) | Intermittent | 2554527703I
107 | atracurium | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2750296C
108 | atracurium | --- mg / --- mL | Cont/Bolus or PCA | 2750296C
109 | avibactam/ceftazidim | --- gram / --- mL | Intermittent | 271758255I
110 | azaCITIDine | --- mg / --- mL |  | 3316777I
111 | azaTHIOprine | --- mg / --- mL |  | 2750417I
112 | azithromycin | 250 mg / 250 mL (1 mg / mL) | Intermittent | 2750452I
113 | azithromycin | 500 mg / 250 mL (2 mg / mL) | Intermittent | 2750452I
114 | azithromycin | --- mg / --- mL | Intermittent | 2750452I
115 | aztreonam | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2750475I
116 | aztreonam | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2750475I
117 | aztreonam | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2750475I
118 | aztreonam | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2750475I
119 | aztreonam | --- mg / --- mL | Intermittent | 2750475I
120 | aztreonam | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2750475I
121 | aztreonam | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2750475I
122 | aztreonam | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2750475I
123 | aztreonam | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2750475I
124 | bamlaniv/estesivimab | 700 mg / 110 mL (6.3636 mg / mL) | Intermittent | 1000000200I
125 | bamlaniv/estesivimab | 700 mg / 160 mL (4.375 mg / mL) | Intermittent | 1000000200I
126 | bamlaniv/estesivimab | 700 mg / 210 mL (3.3333 mg / mL) | Intermittent | 1000000200I
127 | bamlaniv/estesivimab | 700 mg / 310 mL (2.2581 mg / mL) | Intermittent | 1000000200I
128 | bamlaniv/estesivimab | --- mg / --- mL | Intermittent | 1000000200I
129 | bamlanivimab | 700 mg / 200 mL (3.5 mg / mL) | Intermittent | 3749717983I
130 | bamlanivimab | 700 mg / 270 mL (2.5926 mg / mL) | Intermittent | 3749717983I
131 | basiliximab | 20 mg / 50 mL (0.4 mg / mL) |  | 2750634I
132 | basiliximab | --- mg / --- mL |  | 2750634I
133 | belimumab | --- mg / --- mL | Intermittent | 19949906I
134 | belinostat | --- mg / --- mL |  | 45868857I
135 | Bendamusti (Bendeka) | --- mg / --- mL | Intermittent | 17213979bend
136 | Bendamusti (Treanda) | --- mg / --- mL | Intermittent | 17213979trea
137 | bendamustine HCl | --- mg / --- mL |  | 17213979I
138 | bevacizumab | --- mg / --- mL | Intermittent | 3316812I
139 | bezlotoxumab | --- mg / --- mL | Intermittent | 2557014467I
140 | bivalirudin | 250 mg / 50 mL (5 mg / mL) | Cont/Bolus or PCA | 2751435C
141 | bivalirudin | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2751435C
142 | bivalirudin | 250 mg / 500 mL (0.5 mg / mL) | Cont/Bolus or PCA | 2751435C
143 | bivalirudin | 500 mg / 100 mL (5 mg / mL) | Cont/Bolus or PCA | 2751435C
144 | bivalirudin | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA | 2751435C
145 | bivalirudin | 1,250 mg / 250 mL (5 mg / mL) | Cont/Bolus or PCA | 2751435C
146 | bleomycin | --- unit / --- mL | Intermittent | 2751447I
147 | bortezomib | --- mg / --- mL |  | 2751476I
148 | brentuximab | --- mg / --- mL | Intermittent | 19950376I
149 | bretylium | 1,000 mg / 250 mL (4 mg / mL) |  | 2751503I
150 | bretylium | 2,000 mg / 250 mL (8 mg / mL) |  | 2751503I
151 | bretylium | 2,000 mg / 500 mL (4 mg / mL) |  | 2751503I
152 | bumetanide | 10 mg / 100 mL (0.1 mg / mL) |  | 2751668I
153 | bumetanide | 12.5 mg / 50 mL (0.25 mg / mL) |  | 2751668I
154 | bumetanide | 12.5 mg / 250 mL (0.05 mg / mL) |  | 2751668I
155 | bumetanide | 25 mg / 100 mL (0.25 mg / mL) |  | 2751668I
156 | bumetanide | --- mg / --- mL |  | 2751668I
157 | bumetanide continuos | 24 mg / 96 mL (0.25 mg / mL) | Cont/Bolus or PCA | 2751668C
158 | bumetanide continuos | --- mg / --- mL | Cont/Bolus or PCA | 2751668C
159 | bupivacaine | --- mg / --- mL |  | 2751679I
160 | busulfan | --- mg / --- mL |  | 2751784I
161 | cabazitaxel | --- mg / --- mL | Intermittent | 19950393I
162 | caffeine citrate | --- mg / --- mL | Intermittent | 2751845I
163 | caffeine/sod benzoa | 500 mg / 1,000 mL (0.5 mg / mL) | Intermittent | 2751881I
164 | caffeine/sod benzoa | --- mg / --- mL | Intermittent | 2751881I
165 | calcium chloride | 1,000 mg / 60 mL (16.6667 mg / mL) | Intermittent | 2752076I
166 | calcium chloride | 1,000 mg / 110 mL (9.0909 mg / mL) | Intermittent | 2752076I
167 | calcium chloride | 2,000 mg / 120 mL (16.6667 mg / mL) | Intermittent | 2752076I
168 | calcium chloride | --- mg / --- mL | Intermittent | 2752076I
169 | calcium chloride | 1 gram / 60 mL (0.0167 gram / mL) | Intermittent | 2752076I
170 | calcium chloride | 1 gram / 110 mL (0.0091 gram / mL) | Intermittent | 2752076I
171 | calcium chloride | 2 gram / 120 mL (0.0167 gram / mL) | Intermittent | 2752076I
172 | calcium chloride | --- gram / --- mL | Intermittent | 2752076I
173 | calcium chloride | --- mEq / --- mL | Intermittent | 2752076I
174 | calcium gluconate | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752096I
175 | calcium gluconate | 1,000 mg / 60 mL (16.6667 mg / mL) | Intermittent | 2752096I
176 | calcium gluconate | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752096I
177 | calcium gluconate | 1,000 mg / 110 mL (9.0909 mg / mL) | Intermittent | 2752096I
178 | calcium gluconate | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752096I
179 | calcium gluconate | 2,000 mg / 70 mL (28.5714 mg / mL) | Intermittent | 2752096I
180 | calcium gluconate | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752096I
181 | calcium gluconate | 2,000 mg / 120 mL (16.6667 mg / mL) | Intermittent | 2752096I
182 | calcium gluconate | --- mg / --- mL | Intermittent | 2752096I
183 | calcium gluconate | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752096I
184 | calcium gluconate | 1 gram / 60 mL (0.0167 gram / mL) | Intermittent | 2752096I
185 | calcium gluconate | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752096I
186 | calcium gluconate | 1 gram / 110 mL (0.0091 gram / mL) | Intermittent | 2752096I
187 | calcium gluconate | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752096I
188 | calcium gluconate | 2 gram / 70 mL (0.0286 gram / mL) | Intermittent | 2752096I
189 | calcium gluconate | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752096I
190 | calcium gluconate | 2 gram / 120 mL (0.0167 gram / mL) | Intermittent | 2752096I
191 | calcium gluconate | --- gram / --- mL | Intermittent | 2752096I
192 | calcium gluconate | --- mEq / --- mL | Intermittent | 2752096I
193 | cangrelor | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 273246945C
194 | carbenicillin | --- mg / --- mL |  | 2752348I
195 | CARBOplatin | --- mg / --- mL | Intermittent | 2752528I
196 | carfilzomib | --- mg / --- mL | Intermittent | 24495877I
197 | carmustine | --- mg / --- mL |  | 2752551I
198 | casirivimab/imdevima | 1,200 mg / 260 mL (4.6154 mg / mL) | Intermittent | 3958645811I
199 | casirivimab/imdevima | 2,400 mg / 250 mL (9.6 mg / mL) | Intermittent | 3958645811I
200 | casirivimab/imdevima | --- mg / --- mL | Intermittent | 3958645811I
201 | caspofungin | 35 mg / 100 mL (0.35 mg / mL) | Intermittent | 2752634I
202 | caspofungin | 50 mg / 100 mL (0.5 mg / mL) | Intermittent | 2752634I
203 | caspofungin | 50 mg / 250 mL (0.2 mg / mL) | Intermittent | 2752634I
204 | caspofungin | 70 mg / 250 mL (0.28 mg / mL) | Intermittent | 2752634I
205 | caspofungin | --- mg / --- mL | Intermittent | 2752634I
206 | cefamandole | --- mg / --- mL |  | ???????????
207 | ceFAZolin | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2752689I
208 | ceFAZolin | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752689I
209 | ceFAZolin | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752689I
210 | ceFAZolin | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752689I
211 | ceFAZolin | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752689I
212 | ceFAZolin | 3,000 mg / 100 mL (30 mg / mL) | Intermittent | 2752689I
213 | ceFAZolin | --- mg / --- mL | Intermittent | 2752689I
214 | ceFAZolin | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752689I
215 | ceFAZolin | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752689I
216 | ceFAZolin | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752689I
217 | ceFAZolin | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752689I
218 | ceFAZolin | 3 gram / 100 mL (0.03 gram / mL) | Intermittent | 2752689I
219 | ceFAZolin | --- gram / --- mL | Intermittent | 2752689I
220 | cefepime | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2752723I
221 | cefepime | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2752723I
222 | cefepime | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752723I
223 | cefepime | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752723I
224 | cefepime | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752723I
225 | cefepime | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752723I
226 | cefepime | --- mg / --- mL | Intermittent | 2752723I
227 | cefepime | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752723I
228 | cefepime | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752723I
229 | cefepime | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752723I
230 | cefepime | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752723I
231 | cefepime | --- gram / --- mL | Intermittent | 2752723I
232 | cefiderocol | 750 mg / 100 mL (7.5 mg / mL) | Intermittent | 3602322021I
233 | cefiderocol | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 3602322021I
234 | cefiderocol | 1,500 mg / 100 mL (15 mg / mL) | Intermittent | 3602322021I
235 | cefiderocol | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 3602322021I
236 | cefiderocol | 0.75 gram / 100 mL (0.0075 gram / mL) | Intermittent | 3602322021I
237 | cefiderocol | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 3602322021I
238 | cefiderocol | 1.5 gram / 100 mL (0.015 gram / mL) | Intermittent | 3602322021I
239 | cefiderocol | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 3602322021I
240 | cefoperazone | --- mg / --- mL |  | ????????????
241 | cefotaxime | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752732I
242 | cefotaxime | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752732I
243 | cefotaxime | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752732I
244 | cefotaxime | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752732I
245 | cefotaxime | --- mg / --- mL | Intermittent | 2752732I
246 | cefotaxime | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752732I
247 | cefotaxime | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752732I
248 | cefotaxime | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752732I
249 | cefotaxime | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752732I
250 | cefotaxime | --- gram / --- mL | Intermittent | 2752732I
251 | cefoTEtan | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752747I
252 | cefoTEtan | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752747I
253 | cefoTEtan | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752747I
254 | cefoTEtan | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752747I
255 | cefoTEtan | --- mg / --- mL | Intermittent | 2752747I
256 | cefoTEtan | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752747I
257 | cefoTEtan | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752747I
258 | cefoTEtan | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752747I
259 | cefoTEtan | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752747I
260 | cefoTEtan | --- gram / --- mL | Intermittent | 2752747I
261 | cefOXitin | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752760I
262 | cefOXitin | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752760I
263 | cefOXitin | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752760I
264 | cefOXitin | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752760I
265 | cefOXitin | --- mg / --- mL | Intermittent | 2752760I
266 | cefOXitin | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752760I
267 | cefOXitin | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752760I
268 | cefOXitin | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752760I
269 | cefOXitin | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752760I
270 | cefOXitin | --- gram / --- mL | Intermittent | 2752760I
271 | ceftaroline | 200 mg / 50 mL (4 mg / mL) | Intermittent | 19950738I
272 | ceftaroline | 200 mg / 100 mL (2 mg / mL) | Intermittent | 19950738I
273 | ceftaroline | 200 mg / 250 mL (0.8 mg / mL) | Intermittent | 19950738I
274 | ceftaroline | 300 mg / 50 mL (6 mg / mL) | Intermittent | 19950738I
275 | ceftaroline | 300 mg / 100 mL (3 mg / mL) | Intermittent | 19950738I
276 | ceftaroline | 300 mg / 250 mL (1.2 mg / mL) | Intermittent | 19950738I
277 | ceftaroline | 400 mg / 50 mL (8 mg / mL) | Intermittent | 19950738I
278 | ceftaroline | 400 mg / 100 mL (4 mg / mL) | Intermittent | 19950738I
279 | ceftaroline | 400 mg / 250 mL (1.6 mg / mL) | Intermittent | 19950738I
280 | ceftaroline | 600 mg / 50 mL (12 mg / mL) | Intermittent | 19950738I
281 | ceftaroline | 600 mg / 100 mL (6 mg / mL) | Intermittent | 19950738I
282 | ceftaroline | 600 mg / 250 mL (2.4 mg / mL) | Intermittent | 19950738I
283 | ceftaroline | --- mg / --- mL | Intermittent | 19950738I
284 | cefTAZidime | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2752795I
285 | cefTAZidime | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2752795I
286 | cefTAZidime | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752795I
287 | cefTAZidime | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752795I
288 | cefTAZidime | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752795I
289 | cefTAZidime | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752795I
290 | cefTAZidime | --- mg / --- mL | Intermittent | 2752795I
291 | cefTAZidime | 0.5 gram / 50 mL (0.01 gram / mL) | Intermittent | 2752795I
292 | cefTAZidime | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2752795I
293 | cefTAZidime | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752795I
294 | cefTAZidime | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752795I
295 | cefTAZidime | 1.5 gram / 50 mL (0.03 gram / mL) | Intermittent | 2752795I
296 | cefTAZidime | 1.5 gram / 100 mL (0.015 gram / mL) | Intermittent | 2752795I
297 | cefTAZidime | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752795I
298 | cefTAZidime | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752795I
299 | cefTAZidime | --- gram / --- mL | Intermittent | 2752795I
300 | ceftizoxime | 1,000 mg / 50 mL (20 mg / mL) |  | 2752834I
301 | ceftizoxime | 2,000 mg / 50 mL (40 mg / mL) |  | 2752834I
302 | ceftizoxime | --- mg / --- mL |  | 2752834I
303 | ceftizoxime | 1 gram / 50 mL (0.02 gram / mL) |  | 2752834I
304 | ceftizoxime | 2 gram / 50 mL (0.04 gram / mL) |  | 2752834I
305 | ceftizoxime | --- gram / --- mL |  | 2752834I
306 | ceftolozane/tazo | 150 mg / 100 mL (1.5 mg / mL) | Intermittent | 202016343I
307 | ceftolozane/tazo | 375 mg / 100 mL (3.75 mg / mL) | Intermittent | 202016343I
308 | ceftolozane/tazo | 750 mg / 100 mL (7.5 mg / mL) | Intermittent | 202016343I
309 | ceftolozane/tazo | 1,500 mg / 100 mL (15 mg / mL) | Intermittent | 202016343I
310 | ceftolozane/tazo | 3,000 mg / 100 mL (30 mg / mL) | Intermittent | 202016343I
311 | ceftolozane/tazo | --- mg / --- mL | Intermittent | 202016343I
312 | ceftolozane/tazo | 1.5 gram / 100 mL (0.015 gram / mL) | Intermittent | 202016343I
313 | ceftolozane/tazo | 3 gram / 100 mL (0.03 gram / mL) | Intermittent | 202016343I
314 | ceftolozane/tazo | --- gram / --- mL | Intermittent | 202016343I
315 | cefTRIAXone | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752847I
316 | cefTRIAXone | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752847I
317 | cefTRIAXone | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752847I
318 | cefTRIAXone | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752847I
319 | cefTRIAXone | --- mg / --- mL | Intermittent | 2752847I
320 | cefTRIAXone | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752847I
321 | cefTRIAXone | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752847I
322 | cefTRIAXone | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752847I
323 | cefTRIAXone | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752847I
324 | cefTRIAXone | --- gram / --- mL | Intermittent | 2752847I
325 | cefuroxime | 750 mg / 50 mL (15 mg / mL) | Intermittent | 2752869I
326 | cefuroxime | 750 mg / 100 mL (7.5 mg / mL) | Intermittent | 2752869I
327 | cefuroxime | 1,500 mg / 50 mL (30 mg / mL) | Intermittent | 2752869I
328 | cefuroxime | 1,500 mg / 100 mL (15 mg / mL) | Intermittent | 2752869I
329 | cefuroxime | --- mg / --- mL | Intermittent | 2752869I
330 | cefuroxime | 0.75 gram / 50 mL (0.015 gram / mL) | Intermittent | 2752869I
331 | cefuroxime | 0.75 gram / 100 mL (0.0075 gram / mL) | Intermittent | 2752869I
332 | cefuroxime | 1.5 gram / 50 mL (0.03 gram / mL) | Intermittent | 2752869I
333 | cefuroxime | 1.5 gram / 100 mL (0.015 gram / mL) | Intermittent | 2752869I
334 | cefuroxime | --- gram / --- mL | Intermittent | 2752869I
335 | cemiplimab | --- mg / --- mL | Intermittent | 2970621255I
336 | cetuximab | --- mg / --- mL | Intermittent | 3316890I
337 | chloramphenicol | --- mg / --- mL | Intermittent | 2753038I
338 | chlorothiazide | --- mg / --- mL | Intermittent | 2753128I
339 | chlorproMAZINE | --- mg / --- mL | Intermittent | 2753638I
340 | cidofovir | --- mg / --- mL |  | 2753804I
341 | cimetidine | 300 mg / 50 mL (6 mg / mL) |  | 2753818I
342 | cimetidine | --- mg / --- mL |  | 2753818I
343 | ciprofloxacin | 200 mg / 100 mL (2 mg / mL) | Intermittent | 2753846I
344 | ciprofloxacin | 400 mg / 200 mL (2 mg / mL) | Intermittent | 2753846I
345 | ciprofloxacin | --- mg / --- mL | Intermittent | 2753846I
346 | cisatracurium | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2753896C
347 | cisatracurium | 100 mg / 250 mL (0.4 mg / mL) | Cont/Bolus or PCA | 2753896C
348 | cisatracurium | 200 mg / 200 mL (1 mg / mL) | Cont/Bolus or PCA | 2753896C
349 | cisatracurium | 200 mg / 250 mL (0.8 mg / mL) | Cont/Bolus or PCA | 2753896C
350 | cisatracurium | --- mg / --- mL | Cont/Bolus or PCA | 2753896C
351 | CISplatin | --- mg / --- mL | Intermittent | 2753903I
352 | cladribine | --- mg / --- mL |  | 2753956I
353 | clindamycin | 300 mg / 50 mL (6 mg / mL) | Intermittent | 2753995I
354 | clindamycin | 300 mg / 100 mL (3 mg / mL) | Intermittent | 2753995I
355 | clindamycin | 600 mg / 50 mL (12 mg / mL) | Intermittent | 2753995I
356 | clindamycin | 600 mg / 100 mL (6 mg / mL) | Intermittent | 2753995I
357 | clindamycin | 900 mg / 50 mL (18 mg / mL) | Intermittent | 2753995I
358 | clindamycin | 900 mg / 100 mL (9 mg / mL) | Intermittent | 2753995I
359 | clindamycin | --- mg / --- mL | Intermittent | 2753995I
360 | clodronate | --- mg / --- mL |  | 4333632I
361 | clofarabine | --- mg / --- mL |  | 3617188I
362 | cloxacillin | --- mg / --- mL |  | 2754214I
363 | CMV immunoglobulin | --- mg / --- mL |  | 2754840I
364 | colistimethate | --- mg / --- mL | Intermittent | 2754536I
365 | conivaptan | 20 mg / 100 mL (0.2 mg / mL) | Cont/Bolus or PCA | 4145032C
366 | conivaptan | 20 mg / 250 mL (0.08 mg / mL) | Cont/Bolus or PCA | 4145032C
367 | conivaptan | --- mg / --- mL | Cont/Bolus or PCA | 4145032C
368 | conivaptan Loading | 20 mg / 100 mL (0.2 mg / mL) | Intermittent | 4145032I
369 | cotrimoxazole | --- mg / --- mL |  | ????????
370 | crizanlizumab | --- mg / --- mL | Intermittent | 3488204201I
371 | cyclophosphamide | --- mg / --- mL | Intermittent | 2754749I
372 | cycloSPORINE | --- mg / --- mL |  | 2754780I
373 | cycloSPORINE cont | --- mg / --- mL |  | 2754780C
374 | cytarabine | --- mg / --- mL |  | 2754823I
375 | dacarbazine | --- mg / --- mL | Intermittent | 2754845I
376 | daclizumab | --- mg / --- mL |  | 2754853I
377 | DACTINomycin | --- mcg / --- mL |  | 2754858I
378 | dalbavancin | 375 mg / 100 mL (3.75 mg / mL) | Intermittent | 45868475I
379 | dalbavancin | 375 mg / 250 mL (1.5 mg / mL) | Intermittent | 45868475I
380 | dalbavancin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 45868475I
381 | dalbavancin | 500 mg / 250 mL (2 mg / mL) | Intermittent | 45868475I
382 | dalbavancin | 750 mg / 250 mL (3 mg / mL) | Intermittent | 45868475I
383 | dalbavancin | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 45868475I
384 | dalbavancin | 1,125 mg / 250 mL (4.5 mg / mL) | Intermittent | 45868475I
385 | dalbavancin | 1,500 mg / 500 mL (3 mg / mL) | Intermittent | 45868475I
386 | dalbavancin | --- mg / --- mL | Intermittent | 45868475I
387 | dalbavancin | 0.375 gram / 100 mL (0.0038 gram / mL) | Intermittent | 45868475I
388 | dalbavancin | 0.375 gram / 250 mL (0.0015 gram / mL) | Intermittent | 45868475I
389 | dalbavancin | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 45868475I
390 | dalbavancin | 0.5 gram / 250 mL (0.002 gram / mL) | Intermittent | 45868475I
391 | dalbavancin | 0.75 gram / 250 mL (0.003 gram / mL) | Intermittent | 45868475I
392 | dalbavancin | 1 gram / 250 mL (0.004 gram / mL) | Intermittent | 45868475I
393 | dalbavancin | 1.125 gram / 250 mL (0.0045 gram / mL) | Intermittent | 45868475I
394 | dalbavancin | 1.5 gram / 500 mL (0.003 gram / mL) | Intermittent | 45868475I
395 | dalbavancin | --- gram / --- mL | Intermittent | 45868475I
396 | danaparoid | --- unit / --- mL |  | 4333829I
397 | dantrolene | --- mg / --- mL | Intermittent | 2754894I
398 | DAPTOmycin | --- mg / --- mL | Intermittent | 2754915I
399 | daratumumab | --- mg / --- mL | Intermittent | 275110833I
400 | DAUNOrubicin | --- mg / --- mL |  | 2754939I
401 | DAUNOrubicin liposo | --- mg / --- mL |  | 2754945I
402 | decitabine | --- mg / --- mL | Intermittent | 4145048I
403 | deferoxamine | --- mg / --- mL | Intermittent | 2754950I
404 | denileukin | --- mcg / --- mL |  | 2754986I
405 | desferoxamine | --- mg / --- mL |  | 2754950
406 | desmopressin | --- mcg / --- mL | Intermittent | 2755032I
407 | dexamethasone | --- mg / --- mL | Intermittent | 2755093I
408 | dexmedetomidine | 200 mcg / 50 mL (4 mcg / mL) | Cont/Bolus or PCA | 2755261C
409 | dexmedetomidine | 200 mcg / 100 mL (2 mcg / mL) | Cont/Bolus or PCA | 2755261C
410 | dexmedetomidine | 400 mcg / 100 mL (4 mcg / mL) | Cont/Bolus or PCA | 2755261C
411 | dexmedetomidine | --- mcg / --- mL | Cont/Bolus or PCA | 2755261C
412 | dexrazoxane | --- mg / --- mL |  | 2755283I
413 | dextran 40 | 50 gram / 500 mL (0.1 gram / mL) |  | 2755295I
414 | dextran 40 | --- gram / --- mL |  | 2755295I
415 | diazePAM | 50 mg / 50 mL (1 mg / mL) |  | 2755701I
416 | diazePAM | --- mg / --- mL |  | 2755701I
417 | digoxin | --- mcg / --- mL |  | 2755862I
418 | digoxin immune fab | --- mg / --- mL | Intermittent | 2755888I
419 | dihydroergotamine | --- mg / --- mL |  | 2755900I
420 | dilTIAZem | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2755935C
421 | dilTIAZem | 125 mg / 125 mL (1 mg / mL) | Cont/Bolus or PCA | 2755935C
422 | dilTIAZem | --- mg / --- mL | Cont/Bolus or PCA | 2755935C
423 | dimenhyDRINATE | --- mg / --- mL |  | 2756003I
424 | diphenhydrAMINE | --- mg / --- mL |  | 2756043I
425 | dipyridamole | --- mg / --- mL |  | 2756254I
426 | DOBUTamine | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2756310C
427 | DOBUTamine | 500 mg / 250 mL (2 mg / mL) | Cont/Bolus or PCA | 2756310C
428 | DOBUTamine | 500 mg / 500 mL (1 mg / mL) | Cont/Bolus or PCA | 2756310C
429 | DOBUTamine | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA | 2756310C
430 | DOBUTamine | 1,000 mg / 500 mL (2 mg / mL) | Cont/Bolus or PCA | 2756310C
431 | DOBUTamine | --- mg / --- mL | Cont/Bolus or PCA | 2756310C,2756310I
432 | Dobutamine Stress | --- mg / --- mL |  | 2756310
433 | DOCEtaxel | --- mg / --- mL | Intermittent | 2756319I
434 | dolasetron | --- mg / --- mL |  | 2756447I
435 | DOPamine | 400 mg / 250 mL (1.6 mg / mL) | Cont/Bolus or PCA | 2756465C
436 | DOPamine | 800 mg / 250 mL (3.2 mg / mL) | Cont/Bolus or PCA | 2756465C
437 | DOPamine | --- mg / --- mL | Cont/Bolus or PCA | 2756465C
438 | doripenem | 250 mg / 50 mL (5 mg / mL) | Intermittent | 4316756I
439 | doripenem | 250 mg / 100 mL (2.5 mg / mL) | Intermittent | 4316756I
440 | doripenem | 500 mg / 100 mL (5 mg / mL) | Intermittent | 4316756I
441 | doxaCURIUM | --- mg / --- mL |  | 2756488I
442 | doxapram | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2756493C
443 | doxapram | --- mg / --- mL | Cont/Bolus or PCA | 2756493C
444 | DOXOrubicin | --- mg / --- mL | Intermittent | 2756540I
445 | DOXOrubicin, liposo | --- mg / --- mL | Intermittent | 2756552I
446 | doxycycline | 100 mg / 100 mL (1 mg / mL) | Intermittent | 2756557I
447 | doxycycline | 200 mg / 250 mL (0.8 mg / mL) | Intermittent | 2756557I
448 | doxycycline | --- mg / --- mL | Intermittent | 2756557I
449 | droperidol | --- mg / --- mL |  | 2756618I
450 | drotrecogin | --- mg / --- mL |  | 2756633I
451 | durvalumab | 1,500 mg / 280 mL (5.3571 mg / mL) | Intermittent | 2581288283I
452 | durvalumab | --- mg / --- mL | Intermittent | 2581288283I
453 | eculizumab | 300 mg / 60 mL (5 mg / mL) |  | 4245702I
454 | eculizumab | 600 mg / 120 mL (5 mg / mL) |  | 4245702I
455 | eculizumab | 900 mg / 180 mL (5 mg / mL) |  | 4245702I
456 | eculizumab | 1,200 mg / 240 mL (5 mg / mL) |  | 4245702I
457 | eculizumab | --- mg / --- mL |  | 4245702I
458 | EDTA | --- mg / --- mL |  | edta
459 | enalaprilat | --- mcg / --- mL |  | 2757060I
460 | epidural w/ fentanyl | 100 unit / 100 mL (1 unit / mL) | Cont/Bolus or PCA | epiduralfent
461 | epidural w/o fentany | 100 unit / 100 mL (1 unit / mL) | Cont/Bolus or PCA | epidnofent
462 | EPINEPHrine | 4 mg / 250 mL (0.016 mg / mL) | Cont/Bolus or PCA | 2757140C
463 | EPINEPHrine | 5 mg / 250 mL (0.02 mg / mL) | Cont/Bolus or PCA | 2757140C
464 | EPINEPHrine | 8 mg / 250 mL (0.032 mg / mL) | Cont/Bolus or PCA | 2757140C
465 | EPINEPHrine | 10 mg / 500 mL (0.02 mg / mL) | Cont/Bolus or PCA | 2757140C
466 | EPINEPHrine | 16 mg / 250 mL (0.064 mg / mL) | Cont/Bolus or PCA | 2757140C
467 | EPINEPHrine | 30 mg / 250 mL (0.12 mg / mL) | Cont/Bolus or PCA | 2757140C
468 | EPINEPHrine | --- mg / --- mL | Cont/Bolus or PCA | 2757140C
469 | epiRUBicin | --- mg / --- mL |  | 2757216I
470 | epoprostenol | 500 mcg / 100 mL (5 mcg / mL) | Cont/Bolus or PCA | 2757252C
471 | epoprostenol | 1,000 mcg / 100 mL (10 mcg / mL) | Cont/Bolus or PCA | 2757252C
472 | epoprostenol | 3,000 mcg / 100 mL (30 mcg / mL) | Cont/Bolus or PCA | 2757252C
473 | epoprostenol | --- mcg / --- mL | Cont/Bolus or PCA | 2757252C
474 | epoprostenol | 1.5 mg / 50 mL (0.03 mg / mL) | Cont/Bolus or PCA | 2757252C
475 | eptifibatide | 75 mg / 100 mL (0.75 mg / mL) | Cont/Bolus or PCA | 2757273C
476 | eptifibatide | 200 mg / 100 mL (2 mg / mL) | Cont/Bolus or PCA | 2757273C
477 | eriBULin | --- mg / --- mL | Intermittent | 19952644I
478 | ertapenem | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2757309I
479 | ertapenem | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2757309I
480 | ertapenem | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2757309I
481 | ertapenem | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2757309I
482 | ertapenem | 0.5 gram / 50 mL (0.01 gram / mL) | Intermittent | 2757309I
483 | ertapenem | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2757309I
484 | ertapenem | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2757309I
485 | ertapenem | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2757309I
486 | erythromycin | 250 mg / 100 mL (2.5 mg / mL) | Intermittent | 2757314I
487 | erythromycin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2757314I
488 | erythromycin | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 2757314I
489 | erythromycin | --- mg / --- mL | Intermittent | 2757314I
490 | erythromycin | --- gram / --- mL | Intermittent | 2757314I
491 | esmolol | 2,500 mg / 250 mL (10 mg / mL) | Cont/Bolus or PCA | 2757437C
492 | esmolol | 5,000 mg / 250 mL (20 mg / mL) | Cont/Bolus or PCA | 2757437C
493 | esmolol | --- mg / --- mL | Cont/Bolus or PCA | 2757437C
494 | esomeprazole | --- mg / --- mL |  | 2757446C,2757446I
495 | ethanol 10% | --- gram / --- mL |  | 2757676I
496 | ethanol 5% | --- gram / --- mL |  | 2757676
497 | etidronate | --- mg / --- mL |  | 2757876I
498 | etomidate | --- mg / --- mL |  | 2757905I
499 | etoposide | --- mg / --- mL | Intermittent | 2757912I
500 | factor IX | --- unit / --- mL |  | 2757938I
501 | factor VIII | --- unit / --- mL |  | 24606459I
502 | factor VIIIa | --- unit / --- mL |  | 24606459
503 | famotidine | 20 mg / 50 mL (0.4 mg / mL) | Intermittent | 2757966I
504 | famotidine | --- mg / --- mL | Intermittent | 2757966I
505 | fat emulsion 10% | 10 gram / 100 mL (0.1 gram / mL) |  | 2757989
506 | fat emulsion 20% | 20 gram / 100 mL (0.2 gram / mL) | Intermittent | 2757989I
507 | fat emulsion 20% | 25 gram / 125 mL (0.2 gram / mL) | Intermittent | 2757989I
508 | fat emulsion 20% | 50 gram / 250 mL (0.2 gram / mL) | Intermittent | 2757989I
509 | fat emulsion 30% | 30 gram / 100 mL (0.3 gram / mL) |  | 2757989-30%
510 | fenoldopam | 10 mg / 250 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2758046C
511 | fentaNYL | 1,000 mcg / 100 mL (10 mcg / mL) | Cont/Bolus or PCA | 2758060C
512 | fentaNYL | 1,250 mcg / 25 mL (50 mcg / mL) | Cont/Bolus or PCA | 2758060C
513 | fentaNYL | 2,500 mcg / 100 mL (25 mcg / mL) | Cont/Bolus or PCA | 2758060C
514 | fentaNYL | 2,500 mcg / 250 mL (10 mcg / mL) | Cont/Bolus or PCA | 2758060C
515 | fentaNYL | --- mcg / --- mL | Cont/Bolus or PCA | 2758060C
516 | ferric carboxymaltos | 750 mg / 265 mL (2.8302 mg / mL) | Intermittent | 24606832I
517 | ferric carboxymaltos | --- mg / --- mL | Intermittent | 24606832I
518 | ferric gluconate Na | 125 mg / 100 mL (1.25 mg / mL) |  | 2758129I
519 | ferumoxytol | 510 mg / 117 mL (4.359 mg / mL) | Intermittent | 19952174I
520 | ferumoxytol | 1,020 mg / 284 mL (3.5915 mg / mL) | Intermittent | 19952174I
521 | ferumoxytol | 0.51 gram / 117 mL (0.0044 gram / mL) | Intermittent | 19952174I
522 | ferumoxytol | 1.02 gram / 284 mL (0.0036 gram / mL) | Intermittent | 19952174I
523 | filgrastim | --- mcg / --- mL |  | 2758220I
524 | floxuridine | --- mg / --- mL |  | 2758251I
525 | fluconazole | 100 mg / 50 mL (2 mg / mL) | Intermittent | 2758256I
526 | fluconazole | 200 mg / 100 mL (2 mg / mL) | Intermittent | 2758256I
527 | fluconazole | 400 mg / 200 mL (2 mg / mL) | Intermittent | 2758256I
528 | fluconazole | 800 mg / 400 mL (2 mg / mL) | Intermittent | 2758256I
529 | fluconazole | --- mg / --- mL | Intermittent | 2758256I
530 | fludarabine | --- mg / --- mL |  | 2758282I
531 | fluorouracil | --- mg / --- mL | Intermittent | 2758517C,2758517I
532 | folic acid | --- mg / --- mL | Intermittent | 2758687I
533 | fomepizole | --- mg / --- mL | Intermittent | 2758716I
534 | fosaprepitant | 150 mg / 150 mL (1 mg / mL) | Intermittent | 17213916I
535 | foscarnet | 100 mg / 50 mL (2 mg / mL) |  | 2758754I
536 | foscarnet | --- mg / --- mL |  | 2758754I
537 | fosphenytoin | --- mg / --- mL | Intermittent | 2758779I
538 | furosemide | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2758798C
539 | furosemide | 200 mg / 100 mL (2 mg / mL) | Cont/Bolus or PCA | 2758798C
540 | furosemide | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2758798C
541 | furosemide | 300 mg / 30 mL (10 mg / mL) | Cont/Bolus or PCA | 2758798C
542 | furosemide | 500 mg / 100 mL (5 mg / mL) | Cont/Bolus or PCA | 2758798C
543 | furosemide | --- mg / --- mL | Cont/Bolus or PCA | 2758798C
544 | ganciclovir | --- mg / --- mL | Intermittent | 2758863I
545 | gancyclovir | --- mg / --- mL |  | 2758863
546 | ganisetron | --- mcg / --- mL |  | 2759194
547 | gatifloxacin | 200 mg / 100 mL (2 mg / mL) |  | 2758883I
548 | gatifloxacin | 400 mg / 200 mL (2 mg / mL) |  | 2758883I
549 | gatifloxacin | --- mg / --- mL |  | 2758883I
550 | gemcitabine | --- mg / --- mL | Intermittent | 2758906I
551 | gemtuzumab | --- mg / --- mL |  | 2758923I
552 | gentamicin | 80 mg / 100 mL (0.8 mg / mL) | Intermittent | 2758928I
553 | gentamicin | 120 mg / 100 mL (1.2 mg / mL) | Intermittent | 2758928I
554 | gentamicin | --- mg / --- mL | Intermittent | 2758928I
555 | glucagon | 10 mg / 50 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2759022C
556 | glucagon | 10 mg / 100 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2759022C
557 | glucagon | --- mg / --- mL | Cont/Bolus or PCA | 2759022C
558 | golimumab | --- mg / --- mL | Intermittent | 19952385I
559 | granisetron | --- mg / --- mL |  | 2759194I
560 | haloperidol | --- mg / --- mL |  | 2759713I
561 | heparin | 100 unit / 1 mL (100 unit / mL) | Cont/Bolus or PCA | 2759747C
562 | heparin | 1,250 unit / 50 mL (25 unit / mL) | Cont/Bolus or PCA | 2759747C
563 | heparin | 2,500 unit / 25 mL (100 unit / mL) | Cont/Bolus or PCA | 2759747C
564 | heparin | 2,500 unit / 50 mL (50 unit / mL) | Cont/Bolus or PCA | 2759747C
565 | heparin | 5,000 unit / 50 mL (100 unit / mL) | Cont/Bolus or PCA | 2759747C
566 | heparin | 5,000 unit / 100 mL (50 unit / mL) | Cont/Bolus or PCA | 2759747C
567 | heparin | 25,000 unit / 250 mL (100 unit / mL) | Cont/Bolus or PCA | 2759747C
568 | heparin | 25,000 unit / 500 mL (50 unit / mL) | Cont/Bolus or PCA | 2759747C
569 | heparin | --- unit / --- mL | Cont/Bolus or PCA | 2759747C
570 | hetastarch 6% | 30 gram / 500 mL (0.06 gram / mL) | Intermittent | 2759815C,2759815I
571 | hydrALAZINE | --- mg / --- mL |  | 2759877I
572 | hydrochloric acid | 0.1 mEq / 1 mL (0.1 mEq / mL) |  | HCL
573 | hydrochloric acid | 0.2 mEq / 1 mL (0.2 mEq / mL) |  | HCL
574 | hydrochloric acid | --- mEq / --- mL |  | HCL
575 | hydrocortisone | --- mg / --- mL | Intermittent | 2760102I
576 | HYDROmorphone | 20 mg / 100 mL (0.2 mg / mL) |  | 2760421C
577 | HYDROmorphone | 100 mg / 100 mL (1 mg / mL) |  | 2760421C
578 | HYDROmorphone | --- mg / --- mL |  | 2760421C
579 | ibuprofen lysine | --- mg / --- mL |  | 2760660I
580 | ibutilide | 0.5 mg / 50 mL (0.01 mg / mL) | Intermittent | 2760741I
581 | ibutilide | 0.5 mg / 100 mL (0.005 mg / mL) | Intermittent | 2760741I
582 | ibutilide | 1 mg / 50 mL (0.02 mg / mL) | Intermittent | 2760741I
583 | ibutilide | 1 mg / 60 mL (0.0167 mg / mL) | Intermittent | 2760741I
584 | ibutilide | 1 mg / 100 mL (0.01 mg / mL) | Intermittent | 2760741I
585 | ibutilide | --- mg / --- mL | Intermittent | 2760741I
586 | IDArubicin | --- mg / --- mL |  | 2760752I
587 | Ifex / Mesna | --- mg / --- mL |  | 2760768I
588 | Ifex / Mesna | --- gram / --- mL |  | 2760768I
589 | ifosfamide | --- mg / --- mL | Intermittent | 2760761I
590 | ifosfamide | --- gram / --- mL | Intermittent | 2760761I
591 | iloprost | --- mcg / --- mL |  | 3617215I
592 | imiglucerase | --- unit / --- mL |  | 2760783I
593 | imipenem/cilastatin | 250 mg / 100 mL (2.5 mg / mL) | Intermittent | 2760792I
594 | imipenem/cilastatin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2760792I
595 | imipenem/cilastatin | 750 mg / 250 mL (3 mg / mL) | Intermittent | 2760792I
596 | imipenem/cilastatin | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 2760792I
597 | imipenem/cilastatin | --- mg / --- mL | Intermittent | 2760792I
598 | imipenem/cilastatin | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2760792I
599 | imipenem/cilastatin | 1 gram / 250 mL (0.004 gram / mL) | Intermittent | 2760792I
600 | imipenem/cilastatin | --- gram / --- mL | Intermittent | 2760792I
601 | immune glob (IVIG) | 5 gram / 50 mL (0.1 gram / mL) | Intermittent | 2760832I
602 | immune glob (IVIG) | 5 gram / 100 mL (0.05 gram / mL) | Intermittent | 2760832I
603 | immune glob (IVIG) | 10 gram / 100 mL (0.1 gram / mL) | Intermittent | 2760832I
604 | immune glob (IVIG) | 10 gram / 200 mL (0.05 gram / mL) | Intermittent | 2760832I
605 | immune glob (IVIG) | 20 gram / 200 mL (0.1 gram / mL) | Intermittent | 2760832I
606 | immune glob (IVIG) | --- gram / --- mL | Intermittent | 2760832I
607 | immune globulin Rho | --- gram / --- mL |  | 2768135I
608 | inamrinone | 100 mg / 100 mL (1 mg / mL) |  | 2760881I
609 | inamrinone | 300 mg / 60 mL (5 mg / mL) |  | 2760881I
610 | indomethacin | --- mg / --- mL |  | 2760908I
611 | inFLIXimab | --- mg / --- mL | Intermittent | 2760925I
612 | insulin | 10 unit / 1,000 mL (0.01 unit / mL) | Cont/Bolus or PCA | 2761029C
613 | insulin | 100 unit / 100 mL (1 unit / mL) | Cont/Bolus or PCA | 2761029C
614 | insulin | 250 unit / 250 mL (1 unit / mL) | Cont/Bolus or PCA | 2761029C
615 | insulin | --- unit / --- mL | Cont/Bolus or PCA | 2761029C
616 | interferon alpha-2b | --- unit / --- mL |  | 2761065I
617 | ipilimumab | --- mg / --- mL | Intermittent | 19952974I
618 | irinotecan | --- mg / --- mL | Intermittent | 2761340I
619 | irinotecan liposomal | --- mg / --- mL | Intermittent | 275110257I
620 | iron dextran | 25 mg / 50 mL (0.5 mg / mL) |  | 2761345I
621 | iron dextran | --- mg / --- mL |  | 2761345I
622 | iron sucrose | 100 mg / 55 mL (1.8182 mg / mL) | Intermittent | 2761376I
623 | iron sucrose | 100 mg / 105 mL (0.9524 mg / mL) | Intermittent | 2761376I
624 | iron sucrose | 200 mg / 110 mL (1.8182 mg / mL) | Intermittent | 2761376I
625 | iron sucrose | 300 mg / 265 mL (1.1321 mg / mL) | Intermittent | 2761376I
626 | iron sucrose | 400 mg / 270 mL (1.4815 mg / mL) | Intermittent | 2761376I
627 | iron sucrose | 500 mg / 275 mL (1.8182 mg / mL) | Intermittent | 2761376I
628 | iron sucrose | --- mg / --- mL | Intermittent | 2761376I
629 | isoproterenol | 1 mg / 250 mL (0.004 mg / mL) | Cont/Bolus or PCA | 2761419C
630 | isoproterenol | 2 mg / 250 mL (0.008 mg / mL) | Cont/Bolus or PCA | 2761419C
631 | isoproterenol | 2 mg / 500 mL (0.004 mg / mL) | Cont/Bolus or PCA | 2761419C
632 | isoproterenol | --- mg / --- mL | Cont/Bolus or PCA | 2761419C
633 | itraconazole | 200 mg / 60 mL (3.3333 mg / mL) | Intermittent | 2761517I
634 | itraconazole | --- mg / --- mL | Intermittent | 2761517I
635 | ixabepilone | --- mg / --- mL |  | 4316936I
636 | Kcentra (Factor IX) | --- unit / --- mL | Intermittent | 24608305I
637 | ketamine | --- mg / --- mL |  | 2761553I
638 | ketamine ANES | 50 mg / 50 mL (1 mg / mL) | Cont/Bolus or PCA | 2761553C
639 | ketamine ANES | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2761553C
640 | ketamine ANES | --- mg / --- mL | Cont/Bolus or PCA | 2761553C
641 | Kinevac (sincalide) | --- mcg / --- mL | Intermittent | 2768854I
642 | Kinevac (sincalide) | --- mg / --- mL | Intermittent | 2768854I
643 | labetalol | 200 mg / 200 mL (1 mg / mL) | Cont/Bolus or PCA | 2761620C
644 | labetalol | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2761620C
645 | labetalol | 500 mg / 250 mL (2 mg / mL) | Cont/Bolus or PCA | 2761620C
646 | labetalol | --- mg / --- mL | Cont/Bolus or PCA | 2761620C
647 | lacosamide | --- mg / --- mL | Intermittent | 19952919I
648 | lepirudin | 100 mg / 250 mL (0.4 mg / mL) |  | 2761819I
649 | lepirudin | --- mg / --- mL |  | 2761819I
650 | leucovorin | --- mg / --- mL | Intermittent | 2761829I
651 | levETIRAcetam | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2761878I
652 | levETIRAcetam | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2761878I
653 | levETIRAcetam | 1,500 mg / 100 mL (15 mg / mL) | Intermittent | 2761878I
654 | levETIRAcetam | --- mg / --- mL | Intermittent | 2761878I
655 | levETIRAcetam | --- gram / --- mL | Intermittent | 2761878I
656 | levobupivicaine | --- mg / --- mL |  | 2761905I
657 | levOCARNitine | --- mg / --- mL |  | 2761919I
658 | levoFLOXacin | 250 mg / 50 mL (5 mg / mL) | Intermittent | 2761938I
659 | levoFLOXacin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2761938I
660 | levoFLOXacin | 750 mg / 150 mL (5 mg / mL) | Intermittent | 2761938I
661 | levoFLOXacin | --- mg / --- mL | Intermittent | 2761938I
662 | LEVOleucovorin | --- mg / --- mL |  | 17213770I
663 | levothyroxine | --- mcg / --- mL |  | 2761978I
664 | lidocaine | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA | 2762076C
665 | lidocaine | 2,000 mg / 500 mL (4 mg / mL) | Cont/Bolus or PCA | 2762076C
666 | lidocaine | --- mg / --- mL | Cont/Bolus or PCA | 2762076C
667 | linezolid | 200 mg / 100 mL (2 mg / mL) | Intermittent | 2762196I
668 | linezolid | 600 mg / 300 mL (2 mg / mL) | Intermittent | 2762196I
669 | linezolid | --- mg / --- mL | Intermittent | 2762196I
670 | liothyronine | --- mcg / --- mL |  | 2762207I
671 | LORazepam | 20 mg / 100 mL (0.2 mg / mL) |  | 2762369C
672 | LORazepam | 40 mg / 100 mL (0.4 mg / mL) |  | 2762369C
673 | LORazepam | 120 mg / 60 mL (2 mg / mL) |  | 2762369C
674 | LORazepam | --- mg / --- mL |  | 2762369C
675 | mag OB maintenance | 20 gram / 500 mL (0.04 gram / mL) | Cont/Bolus or PCA | 2762580C
676 | mag sulfate | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2762580I
677 | mag sulfate | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2762580I
678 | mag sulfate | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2762580I
679 | mag sulfate | 4 gram / 100 mL (0.04 gram / mL) | Intermittent | 2762580I
680 | mag sulfate | --- gram / --- mL | Intermittent | 2762580I
681 | magnesium sulfate | 1 gram / 50 mL (0.02 gram / mL) |  | 2762580
682 | magnesium sulfate | 4 gram / 100 mL (0.04 gram / mL) |  | 2762580
683 | magnesium sulfate | 40 gram / 500 mL (0.08 gram / mL) |  | 2762580
684 | magnesium sulfate | --- gram / --- mL |  | 2762580
685 | mannitol 20% | 50 gram / 250 mL (0.2 gram / mL) | Intermittent | 2762620I
686 | mannitol 20% | 100 gram / 500 mL (0.2 gram / mL) | Intermittent | 2762620I
687 | mannitol 20% | --- gram / --- mL | Intermittent | 2762620I
688 | mannitol 10% | 100 gram / 1,000 mL (0.1 gram / mL) |  | 2762620
689 | mannitol 10% | --- gram / --- mL |  | 2762620
690 | mannitol 15% | 75 gram / 500 mL (0.15 gram / mL) |  | 2762620-15
691 | mannitol 15% | --- gram / --- mL |  | 2762620-15
692 | mannitol 20% | 50 gram / 250 mL (0.2 gram / mL) |  | 2762620-20
693 | mannitol 20% | --- gram / --- mL |  | 2762620-20
694 | mannitol 25% | 12.5 gram / 50 mL (0.25 gram / mL) |  | 2762620-12.5
695 | mannitol 25% | --- gram / --- mL |  | 2762620-12.5
696 | mannitol 5% | 50 gram / 1,000 mL (0.05 gram / mL) |  | 2762620-5
697 | mannitol 5% | --- gram / --- mL |  | 2762620-5
698 | mechlorethamine | --- mg / --- mL |  | 2762671I
699 | melphalan | --- mg / --- mL |  | 2762749I
700 | meperidine | 10 mg / 1 mL (10 mg / mL) |  | 2762823I
701 | meperidine | 500 mg / 50 mL (10 mg / mL) |  | 2762823C
702 | meperidine | --- mg / --- mL |  | 2762823C
703 | meropenem | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2762894I
704 | meropenem | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2762894I
705 | meropenem | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2762894I
706 | meropenem | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2762894I
707 | meropenem | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2762894I
708 | meropenem | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2762894I
709 | meropenem | --- mg / --- mL | Intermittent | 2762894I
710 | meropenem | 0.5 gram / 50 mL (0.01 gram / mL) | Intermittent | 2762894I
711 | meropenem | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2762894I
712 | meropenem | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2762894I
713 | meropenem | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2762894I
714 | meropenem | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2762894I
715 | meropenem | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2762894I
716 | meropenem | --- gram / --- mL | Intermittent | 2762894I
717 | meropenem/vaborbacta | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2670044615I
718 | meropenem/vaborbacta | 1 gram / 250 mL (0.004 gram / mL) | Intermittent | 2670044615I
719 | meropenem/vaborbacta | 2 gram / 250 mL (0.008 gram / mL) | Intermittent | 2670044615I
720 | meropenem/vaborbacta | 2 gram / 500 mL (0.004 gram / mL) | Intermittent | 2670044615I
721 | meropenem/vaborbacta | 4 gram / 250 mL (0.016 gram / mL) | Intermittent | 2670044615I
722 | meropenem/vaborbacta | 4 gram / 500 mL (0.008 gram / mL) | Intermittent | 2670044615I
723 | meropenem/vaborbacta | 4 gram / 1,000 mL (0.004 gram / mL) | Intermittent | 2670044615I
724 | meropenem/vaborbacta | --- gram / --- mL | Intermittent | 2670044615I
725 | mesna | --- mg / --- mL | Intermittent | 2762918I
726 | methadone | --- mg / --- mL |  | 2763004C
727 | methocarbamol | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 2763080I
728 | methocarbamol | --- mg / --- mL | Intermittent | 2763080I
729 | methohexitol | --- mg / --- mL |  | 2763092I
730 | methotrexate | --- mg / --- mL |  | 2763101I
731 | methotrexate | --- gram / --- mL |  | 2763101I
732 | methyldopa | --- mg / --- mL |  | 2763256I
733 | methylPREDNISolone | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2763323I
734 | methylPREDNISolone | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2763323I
735 | methylPREDNISolone | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2763323I
736 | methylPREDNISolone | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2763323I
737 | methylPREDNISolone | --- mg / --- mL | Intermittent | 2763323I
738 | metoclopramide | --- mg / --- mL |  | 2763400I
739 | metoprolol | --- mg / --- mL |  | 2763424C,2763424I
740 | metroNIDAZOLE | 250 mg / 50 mL (5 mg / mL) | Intermittent | 2763443I
741 | metroNIDAZOLE | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2763443I
742 | metroNIDAZOLE | --- mg / --- mL | Intermittent | 2763443I
743 | mezlocillin | --- mg / --- mL |  | unknown
744 | micafungin | 50 mg / 100 mL (0.5 mg / mL) | Intermittent | 3830619I
745 | micafungin | 100 mg / 100 mL (1 mg / mL) | Intermittent | 3830619I
746 | micafungin | 150 mg / 100 mL (1.5 mg / mL) | Intermittent | 3830619I
747 | micafungin | --- mg / --- mL | Intermittent | 3830619I
748 | miconazole | --- mg / --- mL |  | 19953615I
749 | midazolam | 50 mg / 50 mL (1 mg / mL) | Cont/Bolus or PCA | 2763568C
750 | midazolam | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2763568C
751 | midazolam | 125 mg / 125 mL (1 mg / mL) | Cont/Bolus or PCA | 2763568C
752 | midazolam | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2763568C
753 | midazolam | --- mg / --- mL | Cont/Bolus or PCA | 2763568C
754 | milrinone | 20 mg / 100 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2763605C
755 | milrinone | 40 mg / 200 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2763605C
756 | milrinone | --- mg / --- mL | Cont/Bolus or PCA | 2763605C
757 | minocycline | --- mg / --- mL |  | 2763628I
758 | mitoMYcin | --- mg / --- mL | Intermittent | 2763694I
759 | mitoXANTRONE | --- mg / --- mL |  | 2763708I
760 | morphine | 1 mg / 1 mL (1 mg / mL) |  | 2763791I
761 | morphine | 30 mg / 30 mL (1 mg / mL) |  | 2763791C,2763791I
762 | morphine | 50 mg / 50 mL (1 mg / mL) |  | 2763791C
763 | morphine | 100 mg / 100 mL (1 mg / mL) |  | 2763791C
764 | morphine | --- mg / --- mL |  | 2763791C
765 | moxifloxacin | 400 mg / 250 mL (1.6 mg / mL) |  | 2763901I
766 | moxifloxacin | --- mg / --- mL |  | 2763901I
767 | muromonab | --- mg / --- mL |  | 2764672I
768 | MVI + Thiame + Folic | 10 unit / 1,011 mL (0.0099 unit / mL) | Intermittent | 1000000002I
769 | MVI Only | 10 unit / 510 mL (0.0196 unit / mL) | Intermittent | 1000000003I
770 | MVI Only | 10 unit / 1,010 mL (0.0099 unit / mL) | Intermittent | 1000000003I
771 | mycophenolate | --- mg / --- mL |  | 3320344I
772 | nafcillin | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2764713I
773 | nafcillin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2764713I
774 | nafcillin | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2764713I
775 | nafcillin | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2764713I
776 | nafcillin | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2764713I
777 | nafcillin | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2764713I
778 | nafcillin | --- mg / --- mL | Intermittent | 2764713I
779 | nafcillin | 0.5 gram / 50 mL (0.01 gram / mL) | Intermittent | 2764713I
780 | nafcillin | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2764713I
781 | nafcillin | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2764713I
782 | nafcillin | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2764713I
783 | nafcillin | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2764713I
784 | nafcillin | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2764713I
785 | nafcillin | --- gram / --- mL | Intermittent | 2764713I
786 | nalbuphine | --- mg / --- mL |  | 2764738I
787 | naloxone | 2 mg / 500 mL (0.004 mg / mL) | Cont/Bolus or PCA | 2764760C
788 | naloxone | 4 mg / 100 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2764760C
789 | naloxone | --- mg / --- mL | Cont/Bolus or PCA | 2764760C
790 | nelarabine | --- mg / --- mL |  | 3976209I
791 | nesiritide | 1.5 mg / 250 mL (0.006 mg / mL) | Cont/Bolus or PCA | 2764982C
792 | niCARdipine | 20 mg / 200 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2765030C
793 | niCARdipine | 25 mg / 250 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2765030C
794 | niCARdipine | 40 mg / 200 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2765030C
795 | niCARdipine | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2765030C
796 | niCARdipine | 50 mg / 500 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2765030C
797 | niCARdipine | 500 mg / 250 mL (2 mg / mL) | Cont/Bolus or PCA | 2765030C
798 | niCARdipine | --- mg / --- mL | Cont/Bolus or PCA | 2765030C
799 | nitroglycerin | 25 mg / 250 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2765191C
800 | nitroglycerin | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2765191C
801 | nitroglycerin | --- mg / --- mL | Cont/Bolus or PCA | 2765191C
802 | nitroprusside | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2765267C
803 | nitroprusside | 50 mg / 500 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2765267C
804 | nitroprusside | 100 mg / 250 mL (0.4 mg / mL) | Cont/Bolus or PCA | 2765267C
805 | nitroprusside | --- mg / --- mL | Cont/Bolus or PCA | 2765267C
806 | nivolumab | 240 mg / 100 mL (2.4 mg / mL) | Intermittent | 202016817I
807 | nivolumab | 480 mg / 100 mL (4.8 mg / mL) | Intermittent | 202016817I
808 | nivolumab | --- mg / --- mL | Intermittent | 202016817I
809 | norepinephrine | 4 mg / 250 mL (0.016 mg / mL) | Cont/Bolus or PCA | 2765336C
810 | norepinephrine | 8 mg / 250 mL (0.032 mg / mL) | Cont/Bolus or PCA | 2765336C
811 | norepinephrine | 8 mg / 500 mL (0.016 mg / mL) | Cont/Bolus or PCA | 2765336C
812 | norepinephrine | 16 mg / 250 mL (0.064 mg / mL) | Cont/Bolus or PCA | 2765336C
813 | norepinephrine | 16 mg / 500 mL (0.032 mg / mL) | Cont/Bolus or PCA | 2765336C
814 | norepinephrine | --- mg / --- mL | Cont/Bolus or PCA | 2765336C
815 | obinutuzumab | 100 mg / 100 mL (1 mg / mL) | Intermittent | 25920903I
816 | obinutuzumab | 100 mg / 104 mL (0.9615 mg / mL) | Intermittent | 25920903I
817 | obinutuzumab | 900 mg / 250 mL (3.6 mg / mL) | Intermittent | 25920903I
818 | obinutuzumab | 900 mg / 286 mL (3.1469 mg / mL) | Intermittent | 25920903I
819 | obinutuzumab | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 25920903I
820 | obinutuzumab | 1,000 mg / 290 mL (3.4483 mg / mL) | Intermittent | 25920903I
821 | obinutuzumab | --- mg / --- mL | Intermittent | 25920903I
822 | ocrelizumab | 300 mg / 260 mL (1.1538 mg / mL) | Intermittent | 2581289359I
823 | ocrelizumab | 600 mg / 520 mL (1.1538 mg / mL) | Intermittent | 2581289359I
824 | octreotide | 500 mcg / 100 mL (5 mcg / mL) | Cont/Bolus or PCA | 2765439C
825 | octreotide | 600 mcg / 240 mL (2.5 mcg / mL) | Cont/Bolus or PCA | 2765439C
826 | octreotide | 1,000 mcg / 250 mL (4 mcg / mL) | Cont/Bolus or PCA | 2765439C
827 | octreotide | 1,250 mcg / 250 mL (5 mcg / mL) | Cont/Bolus or PCA | 2765439C
828 | octreotide | --- mcg / --- mL | Cont/Bolus or PCA | 2765439C
829 | ofatumumab | --- mg / --- mL |  | 19954113I
830 | ondansetron | 4 mg / 50 mL (0.08 mg / mL) | Intermittent | 2765642I
831 | ondansetron | 8 mg / 50 mL (0.16 mg / mL) | Intermittent | 2765642I
832 | ondansetron | 8 mg / 54 mL (0.1481 mg / mL) | Intermittent | 2765642I
833 | ondansetron | 32 mg / 50 mL (0.64 mg / mL) | Intermittent | 2765642I
834 | ondansetron | --- mg / --- mL | Intermittent | 2765642I
835 | oritavacin | 1,200 mg / 250 mL (4.8 mg / mL) | Intermittent | 155136453I
836 | oritavacin | 1,200 mg / 1,000 mL (1.2 mg / mL) | Intermittent | 155136453I
837 | oritavacin | 1.2 gram / 250 mL (0.0048 gram / mL) | Intermittent | 155136453I
838 | oritavacin | 1.2 gram / 1,000 mL (0.0012 gram / mL) | Intermittent | 155136453I
839 | oxacillin | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2765734I
840 | oxacillin | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2765734I
841 | oxacillin | --- mg / --- mL | Intermittent | 2765734I
842 | oxacillin | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2765734I
843 | oxacillin | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2765734I
844 | oxaliplatin | --- mg / --- mL | Intermittent | 2765754I
845 | oxytocin | 10 unit / 1,000 mL (0.01 unit / mL) | Cont/Bolus or PCA | 2765930C
846 | oxytocin | 20 unit / 1,000 mL (0.02 unit / mL) | Cont/Bolus or PCA | 2765930C
847 | oxytocin | 30 unit / 500 mL (0.06 unit / mL) | Cont/Bolus or PCA | 2765930C
848 | oxytocin | 30 unit / 1,000 mL (0.03 unit / mL) | Cont/Bolus or PCA | 2765930C
849 | oxytocin | 40 unit / 1,000 mL (0.04 unit / mL) | Cont/Bolus or PCA | 2765930C
850 | oxytocin | --- unit / --- mL | Cont/Bolus or PCA | 2765930C
851 | PACLitaxel | --- mg / --- mL | Intermittent | 2765935I
852 | PACLitaxel protein-b | --- mg / --- mL | Intermittent | 3830361I
853 | pamidronate | 90 mg / 250 mL (0.36 mg / mL) | Intermittent | 2765957I
854 | pamidronate | --- mg / --- mL | Intermittent | 2765957I
855 | pancuronium | --- mg / --- mL | Cont/Bolus or PCA | 2766061C
856 | panitumumab | --- mg / --- mL | Intermittent | 4246362I
857 | pantoprazole | 40 mg / 100 mL (0.4 mg / mL) |  | 2766068
858 | pantoprazole | 80 mg / 100 mL (0.8 mg / mL) |  | 2766068
859 | pantoprazole | 96 mg / 240 mL (0.4 mg / mL) |  | 2766068
860 | pantoprazole | --- mg / --- mL |  | 2766068
861 | pantoprazole bolus | 40 mg / 100 mL (0.4 mg / mL) | Intermittent | 2766068I
862 | pantoprazole bolus | 80 mg / 100 mL (0.8 mg / mL) | Intermittent | 2766068I
863 | pantoprazole drip | 40 mg / 100 mL (0.4 mg / mL) | Cont/Bolus or PCA | 2766068C
864 | pantoprazole drip | 96 mg / 240 mL (0.4 mg / mL) | Cont/Bolus or PCA | 2766068C
865 | pantoprazole drip | --- mg / --- mL | Cont/Bolus or PCA | 2766068C
866 | papaverine | 300 mg / 300 mL (1 mg / mL) |  | 2766098I
867 | papaverine | --- mg / --- mL |  | 2766098I
868 | pegaspargase | --- unit / --- mL |  | 2766160I
869 | pembrolizumab | 200 mg / 100 mL (2 mg / mL) | Intermittent | 114589861I
870 | pembrolizumab | --- mg / --- mL | Intermittent | 114589861I
871 | PEMEtrexed | --- mg / --- mL | Intermittent | 3320387I
872 | penG 2.5 Mill Units | 2.5 unit / 50 mL (0.05 unit / mL) | Intermittent | 19955177I2.5
873 | penG 2.5 Mill Units | 2.5 unit / 100 mL (0.025 unit / mL) | Intermittent | 19955177I2.5
874 | penG 3 Million Units | 3 unit / 50 mL (0.06 unit / mL) | Intermittent | 19955177I-3
875 | penG 5 Million Units | 5 unit / 50 mL (0.1 unit / mL) | Intermittent | 19955177I-5
876 | penG 5 Million Units | 5 unit / 100 mL (0.05 unit / mL) | Intermittent | 19955177
877 | PenG Potass Syringe | --- unit / --- mL | Intermittent | 19955177I
878 | penG x Million Units | --- unit / --- mL | Intermittent | 19955177-x
879 | pentamidine | --- mg / --- mL | Intermittent | 2766284I
880 | PENTobarbital | 2,500 mg / 500 mL (5 mg / mL) |  | 2766302I
881 | PENTobarbital | --- mg / --- mL |  | 2766302I
882 | pentostatin | --- mg / --- mL |  | 2766314I
883 | pertuzumab | 240 mg / 250 mL (0.96 mg / mL) | Intermittent | 24607638I
884 | pertuzumab | 420 mg / 250 mL (1.68 mg / mL) | Intermittent | 24607638I
885 | pertuzumab | 420 mg / 264 mL (1.5909 mg / mL) | Intermittent | 24607638I
886 | pertuzumab | 840 mg / 250 mL (3.36 mg / mL) | Intermittent | 24607638I
887 | pertuzumab | 840 mg / 278 mL (3.0216 mg / mL) | Intermittent | 24607638I
888 | PHENobarbital | 130 mg / 52 mL (2.5 mg / mL) | Intermittent | 2766460I
889 | PHENobarbital | 130 mg / 102 mL (1.2745 mg / mL) | Intermittent | 2766460I
890 | PHENobarbital | 260 mg / 54 mL (4.8148 mg / mL) | Intermittent | 2766460I
891 | PHENobarbital | 260 mg / 104 mL (2.5 mg / mL) | Intermittent | 2766460I
892 | PHENobarbital | --- mg / --- mL | Intermittent | 2766460I
893 | phentolamine | --- mg / --- mL |  | 2766536I
894 | phenylephrine | 10 mg / 250 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2766541C
895 | phenylephrine | 10 mg / 251 mL (0.0398 mg / mL) | Cont/Bolus or PCA | 2766541C
896 | phenylephrine | 20 mg / 250 mL (0.08 mg / mL) | Cont/Bolus or PCA | 2766541C
897 | phenylephrine | 20 mg / 500 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2766541C
898 | phenylephrine | 20 mg / 502 mL (0.0398 mg / mL) | Cont/Bolus or PCA | 2766541C
899 | phenylephrine | 40 mg / 250 mL (0.16 mg / mL) | Cont/Bolus or PCA | 2766541C
900 | phenylephrine | 40 mg / 1,000 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2766541C
901 | phenylephrine | --- mg / --- mL | Cont/Bolus or PCA | 2766541C
902 | phenylephrine ANES | 10 mg / 100 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2766541
903 | phenylephrine ANES | 10 mg / 101 mL (0.099 mg / mL) | Cont/Bolus or PCA | 2766541
904 | phenylephrine ANES | 10 mg / 250 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2766541
905 | phenylephrine ANES | 10 mg / 251 mL (0.0398 mg / mL) | Cont/Bolus or PCA | 2766541
906 | phenylephrine ANES | --- mg / --- mL | Cont/Bolus or PCA | 2766541
907 | phenytoin | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 2766638I
908 | phenytoin | --- mg / --- mL | Intermittent | 2766638I
909 | phytonadione | --- mg / --- mL | Intermittent | 2766694I
910 | piperacillin | 2.25 gram / 50 mL (0.045 gram / mL) |  | 2766817I
911 | piperacillin | 3.37 gram / 50 mL (0.0674 gram / mL) |  | 2766817I
912 | piperacillin | 4.5 gram / 50 mL (0.09 gram / mL) |  | 2766817I
913 | piperacillin/tazo | --- mg / --- mL | Intermittent | 2766819I
914 | piperacillin/tazo | 2.25 gram / 50 mL (0.045 gram / mL) | Intermittent | 2766819I
915 | piperacillin/tazo | 2.25 gram / 100 mL (0.0225 gram / mL) | Intermittent | 2766819I
916 | piperacillin/tazo | 3.375 gram / 50 mL (0.0675 gram / mL) | Intermittent | 2766819I
917 | piperacillin/tazo | 3.375 gram / 100 mL (0.0338 gram / mL) | Intermittent | 2766819I
918 | piperacillin/tazo | 4.5 gram / 50 mL (0.09 gram / mL) | Intermittent | 2766819I
919 | piperacillin/tazo | 4.5 gram / 100 mL (0.045 gram / mL) | Intermittent | 2766819I
920 | piperacillin/tazo | --- gram / --- mL | Intermittent | 2766819I
921 | plicamycin | --- mcg / --- mL |  | 2766901I
922 | polatuzumab vedotin | --- mg / --- mL | Intermittent | 3173912591I
923 | porfirmer | --- mg / --- mL |  | 2767007I
924 | potassium acetate | 10 mEq / 100 mL (0.1 mEq / mL) | Intermittent | 2767012I
925 | potassium acetate | 20 mEq / 100 mL (0.2 mEq / mL) | Intermittent | 2767012I
926 | potassium acetate | 30 mEq / 100 mL (0.3 mEq / mL) | Intermittent | 2767012I
927 | potassium acetate | 40 mEq / 100 mL (0.4 mEq / mL) | Intermittent | 2767012I
928 | potassium acetate | 40 mEq / 520 mL (0.0769 mEq / mL) | Intermittent | 2767012I
929 | potassium acetate | --- mEq / --- mL | Intermittent | 2767012I
930 | potassium chloride | 10 mEq / 50 mL (0.2 mEq / mL) | Intermittent | 1000000100I
931 | potassium chloride | 10 mEq / 100 mL (0.1 mEq / mL) | Intermittent | 1000000100I
932 | potassium chloride | 20 mEq / 100 mL (0.2 mEq / mL) | Intermittent | 1000000100I
933 | potassium chloride | 40 mEq / 100 mL (0.4 mEq / mL) | Intermittent | 1000000100I
934 | potassium chloride | 40 mEq / 520 mL (0.0769 mEq / mL) | Intermittent | 1000000100I
935 | potassium chloride | --- mEq / --- mL | Intermittent | 1000000100I
936 | potassium phosphate | --- mEq / --- mL | Intermittent | 2767185I
937 | potassium phosphate | --- mmol / --- mL | Intermittent | 2767185I
938 | pralidoxime | --- mg / --- mL | Intermittent | 2767273I
939 | pralidoxime | --- gram / --- mL | Intermittent | 2767273I
940 | procainamide | 1,000 mg / 50 mL (20 mg / mL) | Cont/Bolus or PCA | 2767500C
941 | procainamide | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA | 2767500C
942 | procainamide | 2,000 mg / 250 mL (8 mg / mL) | Cont/Bolus or PCA | 2767500C
943 | procainamide | 2,000 mg / 500 mL (4 mg / mL) | Cont/Bolus or PCA | 2767500C
944 | prochlorperazine | --- mg / --- mL |  | 2767546I
945 | promethazine | 25 mg / 51 mL (0.4902 mg / mL) | Intermittent | 2767590I
946 | promethazine | --- mg / --- mL | Intermittent | 2767590I
947 | propofol | 500 mg / 50 mL (10 mg / mL) | Cont/Bolus or PCA | 2767660C
948 | propofol | 1,000 mg / 100 mL (10 mg / mL) | Cont/Bolus or PCA | 2767660C
949 | propofol | --- mg / --- mL | Cont/Bolus or PCA | 2767660C
950 | propofol ANES | 10 mg / 1 mL (10 mg / mL) | Cont/Bolus or PCA | 2767660I
951 | propofol ANES | 500 mg / 50 mL (10 mg / mL) | Cont/Bolus or PCA | 2767660I
952 | propofol ANES | 1,000 mg / 100 mL (10 mg / mL) | Cont/Bolus or PCA | 2767660I
953 | propofol ANES | --- mg / --- mL | Cont/Bolus or PCA | 2767660I
954 | protamine sulfate | --- mg / --- mL |  | 2767711I
955 | quiNIDine | --- mg / --- mL |  | 2767995I
956 | ramucirumab | --- mg / --- mL | Intermittent | 38821169I
957 | raNITIdine | 50 mg / 50 mL (1 mg / mL) | Intermittent | 2768055I
958 | raNITIdine | 50 mg / 100 mL (0.5 mg / mL) | Intermittent | 2768055I
959 | raNITIdine | 50 mg / 102 mL (0.4902 mg / mL) | Intermittent | 2768055I
960 | raNITIdine | 150 mg / 250 mL (0.6 mg / mL) | Intermittent | 2768055I
961 | raNITIdine | 250 mg / 250 mL (1 mg / mL) | Intermittent | 2768055I
962 | raNITIdine | --- mg / --- mL | Intermittent | 2768055I
963 | rasburicase | --- mg / --- mL |  | 2768080I
964 | remdesivir | 100 mg / 230 mL (0.4348 mg / mL) | Intermittent | 3520886709I
965 | remdesivir | 100 mg / 250 mL (0.4 mg / mL) | Intermittent | 3520886709I
966 | remdesivir | 200 mg / 210 mL (0.9524 mg / mL) | Intermittent | 3520886709I
967 | remdesivir | 200 mg / 250 mL (0.8 mg / mL) | Intermittent | 3520886709I
968 | remifentanil | 50 mcg / 1 mL (50 mcg / mL) |  | 2768094
969 | remifentanil | --- mcg / --- mL |  | 2768094
970 | remifentanil ANES | 50 mcg / 1 mL (50 mcg / mL) | Cont/Bolus or PCA | 2768094C
971 | remifentanil ANES | 1,000 mcg / 40 mL (25 mcg / mL) | Cont/Bolus or PCA | 2768094C
972 | remifentanil ANES | 2,000 mcg / 80 mL (25 mcg / mL) | Cont/Bolus or PCA | 2768094C
973 | remifentanil ANES | --- mcg / --- mL | Cont/Bolus or PCA | 2768094C
974 | reteplase | --- unit / --- mL |  | 2768131I
975 | rifAMPin | 300 mg / 250 mL (1.2 mg / mL) | Intermittent | 2768179I
976 | rifAMPin | 600 mg / 500 mL (1.2 mg / mL) | Intermittent | 2768179I
977 | rifAMPin | --- mg / --- mL | Intermittent | 2768179I
978 | ritodrine | --- mg / --- mL |  | 2768244I
979 | riTUXimab | 500 mg / 250 mL (2 mg / mL) | Intermittent | 2768259I
980 | riTUXimab | 1,000 mg / 500 mL (2 mg / mL) | Intermittent | 2768259I
981 | riTUXimab | --- mg / --- mL | Intermittent | 2768259I
982 | rocuronium | 500 mg / 500 mL (1 mg / mL) | Cont/Bolus or PCA | 2768289C
983 | rocuronium | --- mg / --- mL | Cont/Bolus or PCA | 2768289C
984 | rocuronium ANES | 500 mg / 500 mL (1 mg / mL) | Cont/Bolus or PCA | 2768289I
985 | rocuronium ANES | --- mg / --- mL | Cont/Bolus or PCA | 2768289I
986 | romiDEPsin | --- mg / --- mL |  | 19955002I
987 | ropivacaine | 200 mg / 100 mL (2 mg / mL) |  | 2768324C
988 | ropivacaine | 400 mg / 200 mL (2 mg / mL) |  | 2768324C
989 | ropivacaine | --- mg / --- mL |  | 2768324C
990 | salbutamol | --- mg / --- mL |  | salbutamol
991 | sargramostim | --- mcg / --- mL |  | 2768593I
992 | siltuximab | --- mg / --- mL |  | 45868529I
993 | SMX/TMP | --- mg / --- mL | Intermittent | 2769391I
994 | sodium bicarbonate | --- mEq / --- mL |  | 2768888
995 | sodium chloride 3% | 256.6 mEq / 500 mL (0.5132 mEq / mL) |  | 2778727
996 | sodium chloride 5% | 427.5 mEq / 500 mL (0.855 mEq / mL) |  | 2778729
997 | sodium phosphate | --- mmol / --- mL | Intermittent | 2769071I
998 | sotrovimab | 500 mg / 50 mL (10 mg / mL) | Intermittent | 4112134745I
999 | sotrovimab | 500 mg / 100 mL (5 mg / mL) | Intermittent | 4112134745I
1000 | spesolimab | 900 mg / 100 mL (9 mg / mL) | Intermittent | 4788045753I
1001 | streptozocin | --- mg / --- mL |  | 2769343I
1002 | succinylcholine | --- mg / --- mL |  | 2769352I
1003 | SUFentanil | --- mcg / --- mL |  | 2769370I
1004 | SUFentanil ANES | 100 mcg / 100 mL (1 mcg / mL) | Cont/Bolus or PCA | 2769370C
1005 | SUFentanil ANES | --- mcg / --- mL | Cont/Bolus or PCA | 2769370C
1006 | tacrolimus | --- mg / --- mL |  | 2769525I
1007 | tedizolid phosphate | 200 mg / 250 mL (0.8 mg / mL) | Intermittent | 45868623I
1008 | temozolomide | --- mg / --- mL | Intermittent | 2769599I
1009 | temsirolimus | --- mg / --- mL |  | 4246596I
1010 | teniposide | --- mg / --- mL |  | 2769615I
1011 | teprotumumab | --- mg / --- mL | Intermittent | 3602321791I
1012 | teprotumumab | --- gram / --- mL | Intermittent | 3602321791I
1013 | terbutaline | --- mg / --- mL |  | 2769660I
1014 | theophylline | 400 mg / 500 mL (0.8 mg / mL) |  | 2769843I
1015 | theophylline | 800 mg / 500 mL (1.6 mg / mL) |  | 2769843I
1016 | theophylline | --- mg / --- mL |  | 2769843I
1017 | thiamine | 100 mg / 50 mL (2 mg / mL) | Intermittent | 2769929I
1018 | thiamine | 100 mg / 100 mL (1 mg / mL) | Intermittent | 2769929I
1019 | thiamine | --- mg / --- mL | Intermittent | 2769929I
1020 | thiopental | --- mg / --- mL |  | 2769954I
1021 | thiotepa | --- mg / --- mL |  | 2769992I
1022 | ticarcillin | 3.1 gram / 100 mL (0.031 gram / mL) |  | 2770096I
1023 | ticarcillin/clav | --- mg / --- mL | Intermittent | 2770103I
1024 | ticarcillin/clav | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2770103I
1025 | ticarcillin/clav | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2770103I
1026 | ticarcillin/clav | 3.1 gram / 50 mL (0.062 gram / mL) | Intermittent | 2770103I
1027 | ticarcillin/clav | 3.1 gram / 100 mL (0.031 gram / mL) | Intermittent | 2770103I
1028 | tigecycline | 50 mg / 50 mL (1 mg / mL) | Intermittent | 3831121I
1029 | tigecycline | 50 mg / 100 mL (0.5 mg / mL) | Intermittent | 3831121I
1030 | tigecycline | 100 mg / 100 mL (1 mg / mL) | Intermittent | 3831121I
1031 | tirofiban | 12.5 mg / 250 mL (0.05 mg / mL) | Cont/Bolus or PCA | 2770172C
1032 | tobramycin | 80 mg / 50 mL (1.6 mg / mL) | Intermittent | 2770186I
1033 | tobramycin | 80 mg / 100 mL (0.8 mg / mL) | Intermittent | 2770186I
1034 | tobramycin | 120 mg / 100 mL (1.2 mg / mL) | Intermittent | 2770186I
1035 | tobramycin | --- mg / --- mL | Intermittent | 2770186I
1036 | tocilizumab | 800 mg / 100 mL (8 mg / mL) | Intermittent | 19955605I
1037 | tocilizumab | --- mg / --- mL | Intermittent | 19955605I
1038 | tolazoline | --- mg / --- mL |  | tolazoline
1039 | topotecan | --- mg / --- mL | Intermittent | 2770314I
1040 | torsemide | --- mg / --- mL |  | 2770324I
1041 | trabectedin | --- mg / --- mL | Intermittent | 275110021I
1042 | tranexamic BOLUS | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2770412I
1043 | tranexamic BOLUS | --- mg / --- mL | Intermittent | 2770412I
1044 | tranexamic BOLUS | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2770412I
1045 | tranexamic BOLUS | --- gram / --- mL | Intermittent | 2770412I
1046 | trastuzumab | --- mg / --- mL | Intermittent | 2770424I
1047 | trimetrexate | --- mg / --- mL |  | 2770672I
1048 | tromethamine | --- mEq / --- mL |  | 2770725I
1049 | urokinase | --- unit / --- mL |  | 2770879I
1050 | ustekinumab | 260 mg / 250 mL (1.04 mg / mL) | Intermittent | 19955913I
1051 | ustekinumab | 390 mg / 250 mL (1.56 mg / mL) | Intermittent | 19955913I
1052 | ustekinumab | 520 mg / 250 mL (2.08 mg / mL) | Intermittent | 19955913I
1053 | valproate sodium | --- mg / --- mL | Intermittent | 2770918I
1054 | valproic acid | --- mg / --- mL |  | 2770918
1055 | vancomycin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2770944I
1056 | vancomycin | 750 mg / 150 mL (5 mg / mL) | Intermittent | 2770944I
1057 | vancomycin | 750 mg / 250 mL (3 mg / mL) | Intermittent | 2770944I
1058 | vancomycin | 1,000 mg / 200 mL (5 mg / mL) | Intermittent | 2770944I
1059 | vancomycin | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 2770944I
1060 | vancomycin | 1,250 mg / 250 mL (5 mg / mL) | Intermittent | 2770944I
1061 | vancomycin | 1,500 mg / 300 mL (5 mg / mL) | Intermittent | 2770944I
1062 | vancomycin | 1,500 mg / 500 mL (3 mg / mL) | Intermittent | 2770944I
1063 | vancomycin | 1,750 mg / 350 mL (5 mg / mL) | Intermittent | 2770944I
1064 | vancomycin | 1,750 mg / 500 mL (3.5 mg / mL) | Intermittent | 2770944I
1065 | vancomycin | 2,000 mg / 400 mL (5 mg / mL) | Intermittent | 2770944I
1066 | vancomycin | 2,000 mg / 500 mL (4 mg / mL) | Intermittent | 2770944I
1067 | vancomycin | --- mg / --- mL | Intermittent | 2770944I
1068 | vancomycin | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2770944I
1069 | vancomycin | 0.75 gram / 150 mL (0.005 gram / mL) | Intermittent | 2770944I
1070 | vancomycin | 0.75 gram / 250 mL (0.003 gram / mL) | Intermittent | 2770944I
1071 | vancomycin | 1 gram / 200 mL (0.005 gram / mL) | Intermittent | 2770944I
1072 | vancomycin | 1 gram / 250 mL (0.004 gram / mL) | Intermittent | 2770944I
1073 | vancomycin | 1.25 gram / 250 mL (0.005 gram / mL) | Intermittent | 2770944I
1074 | vancomycin | 1.5 gram / 300 mL (0.005 gram / mL) | Intermittent | 2770944I
1075 | vancomycin | 1.5 gram / 500 mL (0.003 gram / mL) | Intermittent | 2770944I
1076 | vancomycin | 1.75 gram / 350 mL (0.005 gram / mL) | Intermittent | 2770944I
1077 | vancomycin | 1.75 gram / 500 mL (0.0035 gram / mL) | Intermittent | 2770944I
1078 | vancomycin | 2 gram / 400 mL (0.005 gram / mL) | Intermittent | 2770944I
1079 | vancomycin | 2 gram / 500 mL (0.004 gram / mL) | Intermittent | 2770944I
1080 | vancomycin | --- gram / --- mL | Intermittent | 2770944I
1081 | vasopressin | 0.02 unit / 1 mL (0.02 unit / mL) | Cont/Bolus or PCA | 2770987C
1082 | vasopressin | 0.04 unit / 1 mL (0.04 unit / mL) | Cont/Bolus or PCA | 2770987C
1083 | vasopressin | 1 unit / 1 mL (1 unit / mL) | Cont/Bolus or PCA | 2770987C
1084 | vasopressin | 4 unit / 100 mL (0.04 unit / mL) | Cont/Bolus or PCA | 2770987C
1085 | vasopressin | 20 unit / 100 mL (0.2 unit / mL) | Cont/Bolus or PCA | 2770987C
1086 | vasopressin | 40 unit / 100 mL (0.4 unit / mL) | Cont/Bolus or PCA | 2770987C
1087 | vasopressin | 80 unit / 100 mL (0.8 unit / mL) | Cont/Bolus or PCA | 2770987C
1088 | vasopressin | 100 unit / 100 mL (1 unit / mL) | Cont/Bolus or PCA | 2770987C
1089 | vasopressin | --- unit / --- mL | Cont/Bolus or PCA | 2770987C
1090 | vecuronium | 1 mg / 1 mL (1 mg / mL) | Cont/Bolus or PCA | 2770992C
1091 | vecuronium | 2 mg / 1 mL (2 mg / mL) | Cont/Bolus or PCA | 2770992C
1092 | vecuronium | 20 mg / 50 mL (0.4 mg / mL) | Cont/Bolus or PCA | 2770992C
1093 | vecuronium | 25 mg / 50 mL (0.5 mg / mL) | Cont/Bolus or PCA | 2770992C
1094 | vecuronium | 30 mg / 50 mL (0.6 mg / mL) | Cont/Bolus or PCA | 2770992C
1095 | vecuronium | 50 mg / 50 mL (1 mg / mL) | Cont/Bolus or PCA | 2770992C
1096 | vecuronium | 50 mg / 100 mL (0.5 mg / mL) | Cont/Bolus or PCA | 2770992C
1097 | vecuronium | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2770992C
1098 | vecuronium | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2770992C
1099 | vecuronium | --- mg / --- mL | Cont/Bolus or PCA | 2770992C
1100 | vedolizumab | 300 mg / 250 mL (1.2 mg / mL) | Intermittent | 45868365I
1101 | verapamil | --- mg / --- mL |  | 2771023I
1102 | vidarabine | --- mg / --- mL |  | vidarabine
1103 | vinBLAStine | --- mg / --- mL | Intermittent | 2771078I
1104 | vinCRIStine | --- mg / --- mL | Intermittent | 2771083I
1105 | vinCRIStine Liposom | --- mg / --- mL |  | 24608273I
1106 | vinorelbine | --- mg / --- mL | Intermittent | 2771090I
1107 | voriconazole | 200 mg / 50 mL (4 mg / mL) | Intermittent | 2771180I
1108 | voriconazole | 400 mg / 100 mL (4 mg / mL) | Intermittent | 2771180I
1109 | voriconazole | --- mg / --- mL | Intermittent | 2771180I
1110 | zidovudine | --- mg / --- mL | Intermittent | 2771278I
1111 | ziv-aflibercept | --- mg / --- mL | Intermittent | 24608076I
1112 | zoledronic acid | 4 mg / 100 mL (0.04 mg / mL) | Intermittent | 2771383I
1113 | zoledronic acid | 5 mg / 100 mL (0.05 mg / mL) | Intermittent | 2771383I
1114 | zoledronic acid | --- mg / --- mL | Intermittent | 2771383I
1115 | zz amiodarone BOLUS | 150 mg / 100 mL (1.5 mg / mL) | Intermittent
1116 | zz amiodarone DRIP | 360 mg / 200 mL (1.8 mg / mL) | Cont/Bolus or PCA
1117 | zz amiodarone DRIP | 900 mg / 500 mL (1.8 mg / mL) | Cont/Bolus or PCA
1118 | zz DOBUTamine | 2,000 mcg / 1 mL (2,000 mcg / mL) | Cont/Bolus or PCA
1119 | zz DOBUTamine | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA
1120 | zz DOBUTamine | 500 mg / 250 mL (2 mg / mL) | Cont/Bolus or PCA
1121 | zz DOBUTamine | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA
1122 | zz DOBUTamine | --- mg / --- mL | Cont/Bolus or PCA
1123 | zz DOPamine | 1,600 mcg / 1 mL (1,600 mcg / mL) | Cont/Bolus or PCA
1124 | zz DOPamine | 400 mg / 250 mL (1.6 mg / mL) | Cont/Bolus or PCA
1125 | zz DOPamine | 800 mg / 250 mL (3.2 mg / mL) | Cont/Bolus or PCA
1126 | zz DOPamine | --- mg / --- mL | Cont/Bolus or PCA
1127 | zz EPINEPHrine | 10 mcg / 1 mL (10 mcg / mL) | Cont/Bolus or PCA
1128 | zz EPINEPHrine | 4 mg / 250 mL (0.016 mg / mL) | Cont/Bolus or PCA
1129 | zz EPINEPHrine | 5 mg / 250 mL (0.02 mg / mL) | Cont/Bolus or PCA
1130 | zz EPINEPHrine | 8 mg / 250 mL (0.032 mg / mL) | Cont/Bolus or PCA
1131 | zz EPINEPHrine | 10 mg / 100 mL (0.1 mg / mL) | Cont/Bolus or PCA
1132 | zz EPINEPHrine | 10 mg / 500 mL (0.02 mg / mL) | Cont/Bolus or PCA
1133 | zz EPINEPHrine | 16 mg / 250 mL (0.064 mg / mL) | Cont/Bolus or PCA
1134 | zz lidocaine | 8 mg / 1 mL (8 mg / mL) | Cont/Bolus or PCA
1135 | zz lidocaine | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA
1136 | zz lidocaine | 2,000 mg / 500 mL (4 mg / mL) | Cont/Bolus or PCA
1137 | zz lidocaine | --- mg / --- mL | Cont/Bolus or PCA
1138 | zz lidocaine | 2 gram / 500 mL (0.004 gram / mL) | Cont/Bolus or PCA
1139 | zz mag sulfate | --- mg / --- mL | Intermittent
1140 | zz mag sulfate | 1 gram / 50 mL (0.02 gram / mL) | Intermittent
1141 | zz mag sulfate | 1 gram / 100 mL (0.01 gram / mL) | Intermittent
1142 | zz mag sulfate | 2 gram / 50 mL (0.04 gram / mL) | Intermittent
1143 | zz mag sulfate | 4 gram / 100 mL (0.04 gram / mL) | Intermittent
1144 | zz mag sulfate | 40 gram / 500 mL (0.08 gram / mL) | Intermittent
1145 | zz mag sulfate | --- gram / --- mL | Intermittent
1146 | zz nitroglycerin | 25 mg / 250 mL (0.1 mg / mL) | Cont/Bolus or PCA
1147 | zz nitroglycerin | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA
1148 | zz nitroglycerin | --- mg / --- mL | Cont/Bolus or PCA
1149 | zz nitroprusside | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA
1150 | zz nitroprusside | 100 mg / 250 mL (0.4 mg / mL) | Cont/Bolus or PCA
1151 | zz nitroprusside | --- mg / --- mL | Cont/Bolus or PCA
1152 | zz norepinephrine | 16 mcg / 1 mL (16 mcg / mL) | Cont/Bolus or PCA
1153 | zz norepinephrine | 4 mg / 250 mL (0.016 mg / mL) | Cont/Bolus or PCA
1154 | zz norepinephrine | 8 mg / 250 mL (0.032 mg / mL) | Cont/Bolus or PCA
1155 | zz norepinephrine | 8 mg / 500 mL (0.016 mg / mL) | Cont/Bolus or PCA
1156 | zz norepinephrine | 16 mg / 250 mL (0.064 mg / mL) | Cont/Bolus or PCA
1157 | zz norepinephrine | 16 mg / 500 mL (0.032 mg / mL) | Cont/Bolus or PCA
1158 | zz norepinephrine | --- mg / --- mL | Cont/Bolus or PCA
1159 | zz phenylephrine | 10 mg / 250 mL (0.04 mg / mL) | Cont/Bolus or PCA
1160 | zz phenylephrine | 20 mg / 500 mL (0.04 mg / mL) | Cont/Bolus or PCA
1161 | zz phenylephrine | 40 mg / 1,000 mL (0.04 mg / mL) | Cont/Bolus or PCA

## Master Fluids
Master Fluids | Master Fluids | Master Fluids | Master Fluids
Line Item | Fluid Name | Supports Secondary | Aliases
1 | ..IVF Maintenance | Yes | 1000000000F,1000000000I
2 | .Bolus IVF | Yes
3 | .IVF + KCL | Yes | 1000000100F
4 | .IVF + KCL 10 mEq/L | Yes | 1000000010F
5 | .IVF + KCL 20 mEq/L | Yes | 1000000020F
6 | .IVF + KCL 30 mEq/L | Yes
7 | .IVF + KCL 40 mEq/L | Yes | 1000000040F
8 | .IVF + Oxytocin | Yes
9 | .IVF + Vitamins | Yes
10 | .Maint IVF + KCL | Yes
11 | .Maint IVF+Additives | Yes
12 | 0.45%NS+KCL10 mEq/L | Yes
13 | 0.45%NS+KCL20 mEq/L | Yes
14 | 0.45%NS+KCL30 mEq/L | Yes
15 | 0.45%NS+KCL40 mEq/L | Yes
16 | 0.9% Normal Saline | Yes
17 | 0.9%NS+KCL 10mEq/L | Yes
18 | 0.9%NS+KCL 20mEq/L | Yes
19 | 0.9%NS+KCL 30mEq/L | Yes
20 | 0.9%NS+KCL 40mEq/L | Yes
21 | 1/2NS Hep 0.5unit/mL | Yes
22 | 1/4NS Hep 0.5unit/mL | Yes
23 | 3% Sodium Chloride | Yes | 2778727F,2778727I
24 | Bicarb Drip | Yes | 2768888F,2768888I
25 | Blood (RBCs) | No
26 | Blood (whole) | No
27 | Blood Reinfusion | No
28 | Cryoprecipitate | No
29 | D10W 1/4NS | Yes
30 | D10W 1/4NS Hep 0.5un | Yes
31 | D10W Hep 0.5units/mL | Yes
32 | D5 + Normal Saline | Yes
33 | D5 1/2Normal Saline | Yes
34 | D5 1/2NS+KCL10mEq/L | Yes
35 | D5 1/2NS+KCL20mEq/L | Yes
36 | D5 1/2NS+KCL30mEq/L | Yes
37 | D5 1/2NS+KCL40mEq/L | Yes
38 | D5 1/3Normal Saline | Yes
39 | D5 1/3NS+KCL10mEq/L | Yes
40 | D5 1/3NS+KCL20mEq/L | Yes
41 | D5 1/3NS+KCL30mEq/L | Yes
42 | D5 1/3NS+KCL40mEq/L | Yes
43 | D5 1/4Normal Saline | Yes
44 | D5 1/4NS Hep 0.5unit | Yes
45 | D5 1/4NS+KCL10mEq/L | Yes
46 | D5 1/4NS+KCL20mEq/L | Yes
47 | D5 1/4NS+KCL30mEq/L | Yes
48 | D5 1/4NS+KCL40mEq/L | Yes
49 | D5 Lactated Ringers | Yes
50 | D5NS+KCL 10 mEq/L | Yes
51 | D5NS+KCL 20 mEq/L | Yes
52 | D5NS+KCL 30 mEq/L | Yes
53 | D5NS+KCL 40 mEq/L | Yes
54 | D5W | Yes
55 | D5W Hep 0.5units/mL | Yes
56 | D5W+KCL 10 mEq/L | Yes
57 | D5W+KCL 20 mEq/L | Yes
58 | D5W+KCL 30 mEq/L | Yes
59 | D5W+KCL 40 mEq/L | Yes
60 | Dextrose 10% | Yes | 2778755F,2778755I
61 | Dextrose in Ringers | Yes
62 | FLUSH | Yes
63 | Fresh Frozen Plasma | No
64 | Hetastarch 6% NS | Yes
65 | Infliximab(Remicade) | Yes
66 | Insulin Maternity | Yes | 2761029F
67 | IVF +Hep 0.5units/mL | Yes
68 | Lactated Ringers | Yes
69 | Lipids 20% | No
70 | Mannitol | No
71 | NS w Hep 0.5units/mL | Yes
72 | Parenteral Nutrition | No | 1000000001F,1000000001I
73 | Platelets | No
74 | PPN | No
75 | PRBCs | No
76 | Replacement Fluid | Yes
77 | Rituximab (Rituxan) | Yes
78 | TPN | No
79 | UAC | No
80 | UVC | No

## Master Clinical Advisories
Master Clinical Advisories | Master Clinical Advisories | Master Clinical Advisories
Line Item | Advisory Name | Advisory Text
1 | 0.2 to 5 micron filt | Use a low-protein binding in-line or add-on filter between 0.2 and 5 microns.
2 | 0.22 Filter Needed | Use a 0.22 micron in-line filter when administering this product.
3 | 1 unit = 1 mL | 1 unit = 1 mL
4 | 12 Hour Tubing | Tubing must be changed every 12 hours
5 | 1unit=1million | 1 unit = 1 million units
6 | 2nd RN to verify | 2 RNs to verify and document concentration and programming prior to infusion.
7 | 48 Hour Tubing | Change tubing every 48 hours.
8 | Alaris Unit | 1 Unit = 1 million ex. 0.6units = 600,000 units
9 | Allergic Reaction | Infuse slowly for the first 10 minutes (25-50 mL/hr) to watch for allergic reaction, if none, rate may be increased to 250 mL/hr.
10 | Alteplase wt based | Dose and volume to be infused need to be set according to weight based protocol to avoid potential over infusion of entire contents of bag.
11 | AmBisome | Existing IV lines must be flushed with D5w prior to infusion. If unable, must start a new line. **An in-fine filter >= 1 micron may be used with infusion.
12 | Anaphylaxis Warning | Observe patient for 1 hour after starting infusion for signs and symptoms of anaphylaxis.
13 | Antivenin | 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
14 | Bleeding Precautions | Monitor for signs of bleeding.
15 | Calcium Precipitate | Do not infuse Calcium and Phosphate though the same line.
16 | CeftriSyringe | Adminster over 60 minutes to decrease risk of bilirubin encephalopathy.
17 | Central Line Only | Run through a CENTRAL line only.
18 | Cerebyx | Not to be infused faster than 150 mgPE / min, due to risk of hypotension.
19 | Check Concentration | *CHECK CONCENTRATION* 2nd RN to verify and document concentration and programming prior to starting infusion.
20 | Code only drugs | This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
21 | Dedicated IV Line | Do not infuse any other drugs through this line.
22 | DEHP | This medication must be infused using DEHP FREE tubing.
23 | Double Concentrated | ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
24 | EKG Required | EKG Montoring required for infusions greater than 10mEq per hour.
25 | Filter and DEHP Free | Requires Low protein binding 0.2 micron filter and DEHP free bag
26 | Fluid Bolus | Administer a fluid bolus prior to starting the cidofovir infusion.
27 | Flush After NS-LR | Following infusion, flush with 30 mL of sterile normal saline or lactated ringers.
28 | Flush Before/After | If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w or NS.
29 | Flush Before/AfterD5 | If a common IV line is being used to administer other drugs in addition to this drug, the line should be flushed before and after each infusion with D5w.
30 | Flush Before/AfterNS | Flush line before and after administration. Can be given on same line as sodium chloride 0.9%.
31 | Heparin | Monitor PTT. Do not administer other drugs through same IV line.
32 | High Risk | This is a HIGH RISK medication. 2nd RN to verify prior to initiating therapy.
33 | High Risk Central | This is a HIGH RISK medication that should be infused via a CENTRAL LINE. 2nd RN to verify prior to initiating therapy.
34 | Hypersensitivity | Risk of hypersensitivity. Check vital signs frequently during infusion.
35 | IFEX | Dose is based on Ifosfamide (IFEX) content
36 | In-line filter | Infusion requires the use of an in-line filter.
37 | INR > 6 | INR > 6 ** Max dose of Kcentra should not exceed 5000 units **
38 | INR 2-4 | INR 2 - 4 ** Max dose of Kcentra should not exceed 2500 units **
39 | INR 4-6 | INR 4 - 6 ** Max dose of Kcentra should not exceed 3500 units **
40 | Insulin Maternity | 75 ml/hr = 0.75 u/hr 100 ml/hr = 1 u/hr 125 ml/hr = 1.25 u/hr 150 ml/hr = 1.5 u/hr
41 | Levophed | *****Must have physician order for rates greater than 60 mcg/min*****
42 | Low Binding 0.2or1.2 | This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
43 | Low Binding Filter | This medication must be infused using a 0.2- micron, low protein binding filter.
44 | Low Binding Tubing | * Use Low Binding Tubing *
45 | MAX DOSE 180 mg | MAXIMUM dose of brentuximab is 180 mg.
46 | Midline Warning | Do NOT infuse through a midline catheter.
47 | Monitoring required | Monitoring module required.
48 | MVI | 1 unit = 1 dose
49 | Neuromuscular BA | *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
50 | No Dextrose | Do NOT administer through lines containing dextrose.
51 | NO FILTER | Do NOT use an in-line filter for administration.
52 | Non-PVC | This medication must be infused using Non-PVC tubing.
53 | NS only | Do not dilute with D5W Dilute with NS 0.9% only
54 | oritavacin | Contraindicated to run ****Heparin**** within 48 hours of Oritavacin administration. May increase risk of bleeding with oral anticoagulant therapy.
55 | Patient Lie Down | Advise patient to lie down during and 30 minutes after infusion to prevent hypotension.
56 | PenG-Syringe | 1 unit in pump = 1,000 units on label. Eg. 200 units in pump = 200,000 units on drug label.
57 | Phenylephrine | *****Must have physician order for rates greater than 300 mcg/min*****
58 | Premedicate | Premedicate before each infusion.
59 | Protect From Light | Protect this drug from light.
60 | SCr 2.0-4.0 | SCr 2.0 - 4.0
61 | SCr<2.0 | SCr < 2.0
62 | Siltuximab | Use PVC with DEHP or PU containing administration set containing a 0.2 micron filter.
63 | Stress Test | For STRESS TEST only.
64 | Vesicant | Drug is a vesicant. Check blood return.
65 | Vesicant No Filter | Drug is a vesicant. Check blood return. Do not use an in-line filter.
66 | Vincristine Lipo | Use dedicated line. Do not filter.

## Master Therapies
Master Therapies | Master Therapies
Line Item | Therapy
1 | ..Less than 3 kg
2 | ..Standard
3 | ..Std. Cardiac/PVD
4 | .Double Concentrated
5 | .Induction
6 | .INR 2 to < 4
7 | .INR 4 - 6
8 | .MG Dosing
9 | .mg, pt weight <40kg
10 | .mg, pt weight>=40kg
11 | .Standard & wt>121kg
12 | .Standard Dose
13 | .Standard NON STEMI
14 | .STANDARD STEMI PCI
15 | .Standard wt < 115kg
16 | .Standard wt >=50 kg
17 | .Std Cardiac or PVD
18 | .Test Dose
19 | .VTE (DVT, PE)
20 | .VTE/DVT/PE
21 | .Weight >= 30 kg
22 | .wt less than 100kg
23 | 0.2 mg / 1 mL
24 | 1 mg/1mL
25 | 1 mg/mL *HIGH CONC*
26 | 10 mEq/L
27 | 15 mEq/L
28 | 1st Dose <=100kg
29 | 1st Dose >100kg
30 | 20 mEq/L
31 | 2nd Dose <=100kg
32 | 2nd Dose >100kg
33 | 3 - 9.9 kg
34 | 30 mEq/L
35 | 3rd Dose <=100kg
36 | 3rd Dose >100kg
37 | 40 mEq/L
38 | 5 mg/1mL *HIGH CONC*
39 | Anavip
40 | Anesthesia
41 | Anti-Convulsant
42 | Anticonvulsant Thera
43 | AUC 2-3 q week
44 | AUC 3-7.5 q3weeks
45 | Based on AUC
46 | Based on BSA
47 | Bleeding
48 | Bleeding/Hemophilia
49 | BSA Based Dose
50 | Cardiac Arrest
51 | CENTRAL < 20kg
52 | CENTRAL >=20kg
53 | Central IV Occlusion
54 | Central Line
55 | Cerebrovascular
56 | CrCl <= 60
57 | CrCl > 60
58 | Crofab
59 | Cycle 1
60 | Cycle 2
61 | Cyclic
62 | Diabetes Insipidus
63 | DVT / Ischemic Limb
64 | Echo Stress Test
65 | ESRD
66 | Every 2 weeks
67 | Every 3 weeks
68 | First Dose
69 | GI Bleed
70 | GM dosing
71 | gram, pt wt < 40kg
72 | gram, pt wt >= 40kg
73 | Hepatic Dose
74 | High Dose
75 | High Dose - GRAMS
76 | Induction
77 | INR > 6
78 | Insulin Maternity
79 | Kimsyra gram
80 | Kimyrsa mg
81 | Load <67kg WtBasedGM
82 | Load <67kg WtBasedMG
83 | Load >67kg WtBasedGM
84 | Load >67kg WtBasedMG
85 | Load NonWtBased MG
86 | Load NonWtBasedGM
87 | Loading Dose
88 | LoDose & >= 153 kg
89 | LoDose & CrCl < 30
90 | LoDose & CrCl >= 30
91 | LoDose,Renal&>=153kg
92 | Maintenance Dose
93 | Maintenance Dose GM
94 | Maintenance Dose MG
95 | Non-Anesthesia
96 | Non-Based Based Dose
97 | OB/L&D
98 | Orbactiv gram
99 | Orbactiv mg
100 | PERIPH < 20kg
101 | PERIPH >=20kg
102 | Peripheral Line
103 | Postpartum
104 | pt age < = 12 years
105 | pt age > = 12 years
106 | pt BSA < 1 m2
107 | pt BSA >=1
108 | pt weight < 20 kg
109 | pt weight < 30 kg
110 | pt weight < 40 kg
111 | pt weight < 50 kg
112 | pt weight < 60 kg
113 | pt weight >= 20 kg
114 | pt weight >= 30 kg
115 | pt weight >= 40 kg
116 | pt weight >= 50 kg
117 | pt weight >= 60 kg
118 | Pulmonary Embolism
119 | Renal Dose
120 | Renal Dose CrCl < 30
121 | Renal Dose wt >121kg
122 | Renal NON STEMI
123 | Renal STEMI PCI
124 | Rheumatoid Arthritis
125 | Second Dose
126 | Sedation
127 | Shock
128 | Spinal bolus
129 | Stress Test
130 | Stroke < 101 kg
131 | Stroke 101kg or more
132 | Subsequent Doses
133 | Third Dose
134 | Undergoing PCI
135 | Weekly
136 | Weight < 30 kg
137 | Weight < 50 kg
138 | Weight >= 115 kg
139 | Weight above 110kg
140 | Weight Based
141 | Weight over 80kg
142 | wt < or = 1.5 kg
143 | WT = 10 - 19.9 kg
144 | WT = 20 - 39.9 kg
145 | WT = 40 kg and over
146 | wt 1.5 - 2.5 kg
147 | wt above 2.5 kg
148 | wt more than 100kg

## Channel Labels
Channel Labels | Channel Labels | Channel Labels | Channel Labels | Channel Labels
Line Item | Channel Labels are not used | Channel Labels are not used | Channel Labels are not used | Channel Labels are not used

## Master Syringe List
Master Syringe List | Master Syringe List | Master Syringe List
Line Item | Syringe Type | Included?
1 | BD Plastipak 1 mL | False
2 | BD Plastipak 3 mL | True
3 | BD Plastipak 5 mL | True
4 | BD Plastipak 10 mL | True
5 | BD Plastipak 20 mL | True
6 | BD Plastipak 30 mL | True
7 | BD Plastipak 50 and 60 mL | True
8 | Monoject 3 mL | True
9 | Monoject 6 mL | False
10 | Monoject 12 mL | False
11 | Monoject 20 mL | True
12 | Monoject 35 mL | False
13 | Monoject 60 mL | False
14 | Terumo 3 mL | True
15 | Terumo 5 mL | True
16 | Terumo 10 mL | True
17 | Terumo 20 mL | True
18 | Terumo 30 mL | True
19 | Terumo 60 mL | True

## Master PCA Syringe List
Master PCA Syringe List | Master PCA Syringe List | Master PCA Syringe List
Line Item | PCA Syringe Type | Included?
1 | BD Plastipak 20 mL | True
2 | BD Plastipak 30 mL | True
3 | BD Plastipak 50 and 60 mL | True
4 | Monoject 20 mL | True
5 | Monoject 35 mL | True
6 | Monoject 60 mL | True
7 | Terumo 20 mL | True
8 | Terumo 30 mL | True
9 | Terumo 60 mL | True
10 | IMS 30 mL Prefill | True