## Intermittent Drugs
Intermittent Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount (- - = custom concentration) | Drug Amount Units | Diluent Volume (- - = custom concentration) | Diluent Volume Units | Final Concentration | Final Concentration Units | Module Availability | Available As | Duration Limits (hh:mm) Hard Min | Duration Limits (hh:mm) Soft Min | Duration Limits (hh:mm) Soft Max | Duration (hh:mm) Initial Value | Dosing Units | Dose Limits Soft Min | Dose Limits Soft Max | Dose Limits Hard Max | Conc. Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max | Clinical Advisory
1 | Infusion Center | aprepitant |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:30 | mg | 99 | 131 |  | n/a |  |  |  | DEHP
2 | Infusion Center | aprepitant |  | 130 | mg | 100 | mL | 1.3 | mg/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:30 | mg | 99 | 131 |  | n/a |  |  |  | DEHP
3 | Infusion Center | atezolizumab |  | 800 | mg | 250 | mL | 3.2 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 799 | 1681 |  | n/a
4 | Infusion Center | atezolizumab |  | 1200 | mg | 250 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 799 | 1681 |  | n/a
5 | Infusion Center | atezolizumab |  | 1680 | mg | 250 | mL | 6.72 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 799 | 1681 |  | n/a
6 | Infusion Center | Bendamusti (Bendeka) |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:10 | 00:15 | 00:10 | mg/m2 | 25 | 120 |  | mg/mL |  | 1.85 | 5.6
7 | Infusion Center | Bendamusti (Treanda) |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:20 | 00:40 | 00:30 | mg/m2 | 25 | 120 |  | mg/mL |  | 0.2 | 0.6
8 | Infusion Center | bevacizumab | First Dose | - - | mg | - - | mL |  |  | Pump | Primary and Secondary | 01:15 | 01:30 | 03:00 | 01:30 | mg/kg | 5 | 15 |  | mg/mL |  | 1 | 25
9 | Infusion Center | bevacizumab | Second Dose | - - | mg | - - | mL |  |  | Pump | Primary and Secondary | 00:45 |  | 01:30 | 01:00 | mg/kg | 5 | 15 |  | mg/mL |  | 1 | 25
10 | Infusion Center | bevacizumab | Subsequent Doses | - - | mg | - - | mL |  |  | Pump | Primary and Secondary | 00:20 |  | 01:30 | 00:30 | mg/kg | 5 | 15 |  | mg/mL |  | 1 | 25
11 | Infusion Center | bleomycin |  | - - | unit | - - | mL |  |  | Pump | Primary and Secondary |  | 00:05 | 00:15 | 00:10 | unit/m2 | 2 | 30 |  | unit/mL |  | 0.1 | 3
12 | Infusion Center | brentuximab |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:20 | 00:40 | 00:30 | mg/kg | 12 | 180 | 180.1 | mg/mL |  | 0.4 | 1.8
13 | Infusion Center | cabazitaxel |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:50 | 01:10 |  | mg/m2 | 15 | 25 |  | mg/mL |  | 0.1 | 0.26
14 | Infusion Center | CARBOplatin | Based on AUC | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:15 | 24:00 | 00:30 | mg | 120 | 900 |  | mg/mL |  | 0.1 | 5
15 | Infusion Center | CARBOplatin | Based on BSA | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:15 | 03:00 | 00:30 | mg/m2 | 100 | 700 |  | mg/mL |  | 0.1 | 5
16 | Infusion Center | carfilzomib |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:05 | 00:35 | 00:30 | mg/m2 | 15 | 70 |  | mg/mL |  | 0.3 | 1.3
17 | Infusion Center | cemiplimab |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:25 | 00:35 | 00:30 | mg | 349 | 351 |  | mg/mL |  | 1 | 20 | In-line filter
18 | Infusion Center | cetuximab | Loading Dose | - - | mg | - - | mL |  |  | Pump | Primary and Secondary | 01:45 |  | 04:00 | 02:00 | mg/m2 | 400 | 500 |  | No
19 | Infusion Center | cetuximab | Maintenance Dose | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg/m2 | 150 | 250 |  | No
20 | Infusion Center | CISplatin |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 01:00 | 24:00 |  | mg/m2 | 10 | 100 |  | mg/mL |  | 0.04 | 2
21 | Infusion Center | cyclophosphamide |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:30 | 01:30 | 01:00 | mg/m2 | 100 | 2000 |  | mg/mL |  | 2 | 20
22 | Infusion Center | dacarbazine |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:15 | 01:00 |  | mg/m2 | 175 | 800 |  | mg/mL |  | 0.5 | 8
23 | Infusion Center | daratumumab | First Dose | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 06:30 | 20:00 |  | mg/kg | 15 | 16 | 16.1 | mg/mL |  | 0.5 | 3
24 | Infusion Center | daratumumab | Second Dose | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 04:00 | 10:00 |  | mg/kg | 15 | 16 | 16.1 | mg/mL |  | 1 | 5
25 | Infusion Center | daratumumab | Subsequent Doses | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 03:50 | 10:00 |  | mg/kg | 15 | 16 | 16.1 | mg/mL |  | 1 | 5
26 | Infusion Center | dexamethasone |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:05 | 00:30 | 00:15 | mg | 1 | 21 |  | No
27 | Infusion Center | DOCEtaxel |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg/m2 | 16 | 100 |  | mg/mL |  | 0.3 | 0.74
28 | Infusion Center | DOXOrubicin |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:10 | 24:00 |  | mg/m2 | 3 | 90 |  | mg/mL |  | 0.1 | 5
29 | Infusion Center | DOXOrubicin, liposo |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg/m2 | 2.5 | 75 |  | mg/mL |  | 0.01 | 0.36
30 | Infusion Center | erBULin |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:02 | 00:10 | 00:10 | mg/m2 | 0.7 | 1.4 |  | mg/mL |  | 0.01 | 0.5
31 | Infusion Center | etoposide |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:30 | 03:00 |  | mg/m2 | 35 | 750 |  | mg/mL |  | 0.1 | 0.4
32 | Infusion Center | fluorouracil |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:15 | 24:00 |  | mg/m2 | 200 | 1000 |  | mg/mL |  | 0.1 | 50
33 | Infusion Center | fosaprepitant |  | 150 | mg | 150 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 00:40 | 00:30 | mg | 149 | 151 |  | n/a
34 | Infusion Center | gemcitabine |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg/m2 | 300 | 1250 |  | mg/mL |  | 0.1 | 25
35 | Infusion Center | ifosfamide |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:30 | 25:00 |  | mg/m2 | 250 | 5000 |  | mg/mL |  | 0.6 | 20
36 | Infusion Center | ipilimumab |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 01:25 | 01:35 | 01:30 | mg/kg | 3 | 10 |  | mg/mL |  | 1 | 2
37 | Infusion Center | irinotecan |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary | 01:20 |  | 02:30 | 01:30 | mg/m2 | 20 | 350 |  | mg/mL |  | 0.12 | 2.8
38 | Infusion Center | irinotecan liposomal |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 01:20 | 02:00 | 01:30 | mg/m2 | 34 | 71 |  | mg/mL | 0.05 | 0.1 | 0.6 | NO FILTER
39 | Infusion Center | leucovorin |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mg/m2 | 10 | 1000 |  | mg/mL |  | 0.05 | 20
40 | Infusion Center | magnesium sulfate |  | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 02:30 | 01:00 | gram | 0.9 | 1.1 |  | n/a
41 | Infusion Center | mesna |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:15 | 25:00 |  | mg/m2 | 300 | 5000 |  | mg/mL |  | 4 | 20
42 | Infusion Center | mitoMYcin |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:15 | 00:45 | 00:30 | mg/m2 | 5 | 20 |  | mg/mL |  | 0.02 | 0.04
43 | Infusion Center | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a
44 | Infusion Center | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a
45 | Infusion Center | obinutuzumab | First Dose | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 04:00 | 04:01 | 04:00 | mg | 100 | 100.1 |  | mg/mL |  | 0.9 | 1.1
46 | Infusion Center | obinutuzumab | Subsequent Doses | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 03:00 | 08:00 |  | mg | 900 | 1000 |  | mg/mL |  | 3 | 3.5
47 | Infusion Center | ondansetron |  | 8 | mg | 50 | mL | 0.16 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 4 | 16 |  | n/a
48 | Infusion Center | oxaliplatin |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 01:50 | 06:30 | 02:00 | mg/m2 | 65 | 130 |  | mg/mL |  | 0.2 | 1.2
49 | Infusion Center | PACLitaxel |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:50 | 03:10 |  | mg/m2 | 40 | 225 |  | mg/mL |  | 0.3 | 1.2
50 | Infusion Center | PACLitaxel protein-b |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:25 | 00:35 | 00:30 | mg/m2 | 49 | 261 |  | mg/mL |  | 4.9 | 5.1
51 | Infusion Center | panitumumab |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg/kg | 3 | 6 |  | mg/mL |  | 2 | 10
52 | Infusion Center | pembrolizumab |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 00:35 | 00:30 | mg | 100 | 250 |  | n/a
53 | Infusion Center | PEMEtrexed |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:10 | mg/m2 | 250 | 900 |  | mg/mL |  | 3 | 13
54 | Infusion Center | pertuzumab | Loading Dose | 840 | mg | 250 | mL | 3.36 | mg/mL | Pump | Primary and Secondary | 01:00 |  | 02:00 | 01:00 | mg | 839 |  | 841 | n/a
55 | Infusion Center | pertuzumab | Maintenance Dose | 420 | mg | 250 | mL | 1.68 | mg/mL | Pump | Primary and Secondary | 00:30 |  | 02:00 | 00:30 | mg | 419 |  | 421 | n/a
56 | Infusion Center | ramucirumab |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 01:00 | 02:00 | 01:00 | mg/kg | 8 | 10 |  | mg/mL |  | 0.8 | 5
57 | Infusion Center | raNITIdine |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:20 | mg | 25 | 50 |  | n/a
58 | Infusion Center | riTUXimab | Based on BSA | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg/m2 | 250 | 500 |  | No
59 | Infusion Center | riTUXimab | Rheumatoid Arthritis | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg | 500 | 1200 |  | mg/mL |  | 0.5 | 5
60 | Infusion Center | temozolomide |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 01:20 | 02:00 | 01:30 | mg/m2 | 74 | 201 |  | mg/mL |  | 2.4 | 2.6 | Flush Before/After
61 | Infusion Center | topotecan |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg/m2 | 0.4 | 1.5 | 4 | mg/mL |  | 0.01 | 0.1
62 | Infusion Center | trastuzumab | Loading Dose | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 01:00 | 02:00 | 01:30 | mg/kg | 2 | 8 |  | mg/mL |  | 0.4 | 5
63 | Infusion Center | trastuzumab | Subsequent Doses | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg/kg | 2 | 8 |  | mg/mL |  | 0.4 | 5
64 | Infusion Center | vinBLAStine |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:05 | 00:20 | 00:15 | mg/m2 | 0.75 | 18.5 | 18.6 | mg/mL |  | 0.02 | 0.8
65 | Infusion Center | vinCRIStine |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:05 | 00:20 | 00:10 | mg | 0.4 | 2 | 2.1 | mg/mL |  | 0.01 | 0.1
66 | Infusion Center | vinorelbine |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:06 | 00:30 | 00:10 | mg/m2 | 5 | 35 |  | mg/mL |  | 0.5 | 2
67 | Infusion Center | ziv-aflibercept |  | - - | mg | - - | mL |  |  | Pump | Primary and Secondary |  | 00:55 | 01:05 | 01:00 | mg/kg | 2 | 4 |  | mg/mL |  | 0.6 | 8
68 | Infusion Center | zoledronic (ZOMETA) |  | 4 | mg | 100 | mL | 0.04 | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 02:00 | 00:30 | mg | 3 | 4 |  | n/a
Note on Y3: This medication must be infused using DEHP FREE tubing.
Note on Y4: This medication must be infused using DEHP FREE tubing.
Note on Y19: Infusion requires the use of an in-line filter.
Note on Y40: Do NOT use an in-line filter for administration.
Note on Y62: Flush line before and after administration. Can be given on same line as sodium chloride 0.9%.

## Fluids
Fluids
Line Item | Profile | Fluid Name | Therapy Name | Module Availability | Supports Secondary (Y/N) | Rate Units | Rate Limits (mL/h) Soft Min | Rate Limits (mL/h) Soft Max | Rate Limits (mL/h) Hard Max | Clinical Advisory
1 | Infusion Center | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 999
2 | Infusion Center | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 1 | 999
3 | Infusion Center | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 667
4 | Infusion Center | .IVF + KCL 30 mEq/L |  | Pump | Y | mL/h | 1 | 334 | 667
5 | Infusion Center | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 1 | 250 | 500
6 | Infusion Center | .Maint IVF+Additives |  | Pump | Y | mL/h | 1 | 250
7 | Infusion Center | Dextrose 10% |  | Pump | Y | mL/h | 15 | 200
8 | Infusion Center | Rituximab (Rituxan) |  | Pump | Y | mL/h | 0.1 | 250

## PC Unit Config
| PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit
Line Item | Profile | Alarm Audio Profile | Minimum Audio Volume | Default Audio Volume | Anesthesia Mode | Battery Meter Display | Color Display | Key Click Audio | Limit Checking | Max Patient BSA (m2) | Max Patient Weight (kg) | Patient ID Entry | Preventive Maintenance Reminder | Tamper Resistant | Authorized User Mode
1 | Infusion Center | Profile 1 | 1 | 5 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 500 | Enabled | Enabled | Enabled | Disabled

## Pump Config
| Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module
Line Item | Profile | Accumulated Air-In-Line | Air In Line (microliter) | Auto-Restart Attempts | KVO Rate Adjust (mL/h) | Max Occlusion Pressure (mmHg) | Default Pressure (mmHg) | Maximum Rate (mL/h) | Maximum VTBI (mL) | Pressure Mode Lock Status | Pressure Mode Selection | Secondary Infusion | SEC-PRI Audio Alert
1 | Infusion Center | Enabled | 75 | 0 | 1 | 525 | 525 | 999 | 9999 | Unlocked | Pump | Enabled | Enabled

## Pump & Syringe Shared Config
| Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe)
Line Item | Profile | Delay Start Options | Delay Start Callback | Drug Calculation | Drug Calculation Bolus Dose | Multidose | Multidose Callback | Pressure Dynamic Display | Volume/Duration
1 | Infusion Center | Disabled | n/a | Enabled | Disabled | Disabled | n/a | Disabled | Enabled