## PC Unit
BD Alaris™ Guardrails™ Editor v12.1.2 Configuration Worksheets
BD Alaris™ PC unit
BD Alaris™ Products
Fields highlighted the same color are related features AND grey highlights mean they are enabled with Drug Setups
By entering the profile name in the PCU worksheet, it will populate the other worksheets. To delete modules not being used highlight the corresponding tab, right click and delete. Enter account name, version date, if final submission, and consultant name. |  |  | Conway Medical Center |  |  |  |  | 2022-01-01 |  | Final Submission |  | YES |  | Cheryl Martineau
Profiles
BD Alaris™ PC unit |  |  | Adult | Infusion Center | Nursery | Pediatrics
Feature | Default Setting | Options | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
Alarm Audio Profile | Profile 1 | Profile 1, 2, 3, or 4 | Profile 1 | Profile 1 | Profile 1 | Profile 1
Minimum Audio Volume | 1 | 1, 2, 3, 4, or 5 | 3 | 3 | 2 | 3
Default Alarm Volume | 5 | 1, 2, 3, 4, or 5 | 4 | 4 | 3 | 4
Anesthesia Mode | Disabled | Enabled, Disabled | Enabled | Disabled | Disabled | Disabled
Battery Meter Display | Enabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
Color Display | Enabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
Key Click Audio | Enabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
Limit Checking | Always | Always, Smart | Always | Always | Always | Always
Max Patient BSA (m2) | 3 | 0.05 - 3.0 | 3 | 3 | 3 | 3
Max Patient Weight (Kg) | 500 kg | 0.1 - 500 | 500 | 500 | 7 | 250
Patient ID Entry | Enabled | Enabled, Disabled | Disabled | Disabled | Disabled | Disabled
Preventive Maintenance Reminder | Enabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
Tamper Resistant | Enabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
Authorized User Mode | Disabled | Enabled, Disabled, Enabled with Override | Disabled | Disabled | Disabled | Disabled
No-Guardrails™, Basic Infusion Clinical Advisory | Disabled | Enabled, Disabled - When enabled, text can be edited | Enabled | Enabled | Enabled | Enabled
BD Alaris™ PCU Settings
Setting | [10 empty] | Changed at Bedside |  | Consider for Default |  | NOTES
Alarm Audio Profile | Four different alarm patterns can be selected to help differentiate between audio alarms in certain hospital environments. To demonstrate the alarm Profiles, the BD Alaris™ PCU must be in the Configuration Mode. To access Configuration Mode, press Options and System On at the same time; continue holding both keys until you see the configurations screen appear. Once in Configuration Mode, press Options > Page Down > System Configurations > PC Unit > Alarm Audio > Test NOTE: Profile 4 audio settings support standards compliance. Setting the Alarm Audio Profile to Profile 4 could potentially result in the BD Alaris™ System alarms sounding similar to other devices, such as respirators and monitoring services that are compliant with the same safety standards. | [9 empty] | No |  | Profile 1
Minimum Audio Volume | Ability to set a minimum audio volume limit. The user will only be able to lower the device audio volume to this set level. When the minimum audio volume level has been reached the Softer soft key becomes unavailable and the PCU emits an illegal key press audio tone. | [9 empty] | No |  | Level 1
Default Alarm Volume | Five different volume settings for the selected alarm audio pattern. To demonstrate the volume, select Audio Adjust on the Main Screen. Type Sound Note: • Advisory/Message: One short beep every 2 seconds. Variable volume; can be silenced for 2 minutes. • Alarm: Choice of four alarm audio profiles. Variable volume; can be silenced for 2 minutes. • Error (Hardware Detected): Fixed maximum decibel volume; cannot be silenced. • Error (Software Detected): Fixed maximum decibel volume; can be silenced for 2 minutes. • Illegal Key Press: Two short beeps. Variable volume; cannot be silenced. • Key Click: One short beep. Fixed minimum volume; can be silenced/disabled in System Configuration. • Prompt: One short beep every 2 seconds. Variable volume; can be silenced. | [9 empty] | Yes |  | NICU 1-3 PICU 3-5 ICU 3-5 MedSurg 5 Clinics 2-4 Oncology 3-5
Anesthesia Mode | Optional setting for each Profile that sets the device to a mode that is user friendly for anesthesiology providers. It allows a subset of features and a unique, anesthesia-only drug list to be available. Users have access to this specific drug list and all drugs within the Profile. Some additional features of Anesthesia Mode: • A channel can be paused indefinitely without a call back alert. • The air-in-line associated with the Profile can be adjusted by the user up to 500 mcL. • All Guardrails™ Limits are set to Soft. • Dose Checking Mode is set to Smart. • Key click audio is disabled. • Tamper Resistant mode (panel locked) is not available. • All Guardrails™ Drug Library entries are available for selection. • Basic Bolus dose is available regardless of configuration settings. • Review of drug calculation setup page is omitted when restoring a stopped drug calculation. • Clinical Advisories are not displayed. • Alaris™ Auto ID module is not available. | [9 empty] | No |  | ____
Battery Meter Display | When enabled, the instrument will display the remaining battery run time calculated based on the current operating condition. | [9 empty] | No |  | Enabled
Color Display | When enabled, the system will display in color. The color display will help provide quick recognition of display notifications. If disabled, the system will have a monochrome display. To support standards compliance, the Color Display setting must be set to Enabled | [9 empty] | No |  | Enabled
Key Click Audio | Provides an audible sound (single beep) with each valid key press. The invalid key press audio (double beep) is louder than the valid key press. The audio alarm applicable to invalid key presses cannot be disabled and is not affected by the selection of a lower alarm volume (through Audio Adjust). | [9 empty] | No |  | Enabled
Limit Checking | The Limit Checking feature has two options: 1. Always option causes a Guardrails™ alert to occur each time a dose limit is exceeded 2. Smart option causes an initial Guardrails™ alert to occur when a dose limit is exceeded; subsequent programming beyond the dose limit will not receive additional alerts. With both options, the drug label on the Channel Message display provides an indicator (↑↑↑ or LLL) that the dose is beyond the current Guardrails™ Soft limit. | [9 empty] | No |  | Always
Max Patient BSA (m2) | Sets maximum patient body surface area. | [9 empty] | No |  | Adult - 3 Pediatric 20-50kg - 2.5 Pediatric 5-20kg - 1.5 Pediatric <5kg - 0.99
Max Patient Weight (kg) | Sets maximum patient weight. | [9 empty] | No |  | Adult - 250-400 NICU - <2kg NICU - 7kg Pediatric - 99 Pediatric - 250
Patient ID Entry | If enabled, the Patient ID entry screen appears at start-up, after selecting Yes to “New Patient?” question and selection of Profile. An alphanumeric identifier, of up to 16 characters, can be entered and will be displayed on the Main Screen. When disabled, the ID entry dialog box is only accessible from the System Options menu. | [9 empty] | No |  | Disabled
Preventive Maintenance Reminder | This feature displays a reminder when the preventive maintenance date for the PCU or attached modules has been reached. Enabled will display the Preventive Maintenance Reminder when designated. Disabled will not provide a Preventive Maintenance Reminder display. | [9 empty] | No |  | Disabled
Tamper Resistant | This feature provides a quick, one-touch lockout of the front panel keypad for all attached modules and the PCU. | [9 empty] | No |  | Enabled
Authorized User Mode | Authorized User Mode is a feature that combines the PCU Tamper Resist feature with the BD Alaris™ Auto-ID module. This feature is designed to ensure that only clinicians with a bar code on their ID badge can program the BD Alaris™ System. When Authorized User Mode is enabled, the PCU automatically turns on Tamper Resist mode upon power on and five minutes after programming is completed. When Enabled with Override is selected, the Authorized User Mode can be temporarily disabled (for five minutes) by pressing and holding the tamper resist switch. This may be desired for those situations when the clinician’s ID badge is not available. | [9 empty] | No |  | Disabled
No-Guardrails™, Basic Infusion Clinical Advisory | When enabled, a clinical advisory message appears on the PCU when Basic primary, Basic secondary, or Drug Calc is selected. This clinical advisory is used to reinforce that there is no Guardrails™ protection for the selection made. A default clinical advisory message is provided and can be edited to meet the hospital's need. The Advisory text will appear in all profiles that have this Advisory feature enabled. Because the PCU display screen uses a variable width font, the number of characters allowed on each line of the clinical advisory message will vary. The total number of characters for all eight lines will vary from 112-200 characters. Default text if enabled: This infusion selection DOES NOT have Guardrails Protection | [9 empty] | No |  | Disabled
BD-10751
BD, the BD Logo and Alaris™, Guardrails™ are trademarks of Becton, Dickinson and Company. © 2021 BD and its subsidiaries. All rights reserved
Note on A28: Default text if enabled: This infusion selection DOES NOT have Guardrails Protection

## PUMP
BD Alaris™ Guardrails™ Editor v12.1.2 Configuration Worksheets
BD Alaris™ Pump Module
BD Alaris™ Products
Fields highlighted the same color are related features AND grey highlights mean they are enabled with Drug Setups
Profile Names
BD Alaris™ Pump Module |  |  | Adult | Infusion Center | Nursery | Pediatrics
Feature | Default Setting | Options | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
Accumulated Air-in-line | Enabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
Air In Line (microliter) | 75 | 50, 75, 125, 175 or 250. | 250 | 250 | 50 | 75
Auto-Restart Attempts | 0 | 0 - 9 attempts | 9 | 9 | 1 | 9
KVO Rate Adjust (mL/h) | 1 | 0.1 - 20 | 20 | 20 | 2 | 5
Max Occlusion Pressure (mmHg) | 525 | 50 - 525 | 525 | 525 | 525 | 525
Default Pressure (mmHg) | 525 | 50 - 525 | 525 | 525 | 525 | 525
Maximum Rate (mL/h) | 999 | 0.1 - 999 | 999 | 999 | 25 | 999
Maximum VTBI (mL) | 9999 | 0.1 - 9999 | 3000 | 3000 | 1200 | 2200
Pressure Mode Lock Status | Unlocked | Locked, Unlocked | Unlocked | Unlocked | Unlocked | Unlocked
Pressure Mode Selection | Pump (P) | Pump (P), Selectable (S) | Pump | Pump | Pump | Pump
Secondary Infusion | Enabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
SEC-PRI Audio Alert | Enabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
BD Alaris™ Pump Module Settings
Setting | [10 empty] | Changed at Bedside |  | Consider for Default |  | NOTES
Accumulated Air-In-Line | Detects the presence of multiple air bubbles that are too small to be detected by the single bolus AIL detection limit. The volume of air that trips the accumulated air detection limit is a percentage of a specified averaging window volume that is variable based upon the current setting for single air bolus. • For the 50 microliters setting, the air volume that causes an alarm is 0.18mL (1.2mL window volume with 15% air in window volume). • For the 75 microliters setting, the air volume that causes an alarm is 1.1mL (4.4mL window volume with 25% air in window volume). • For the 125 microliters setting, the air volume that causes an alarm is 1.1mL (4.4mL window volume with 25% air in window volume). • For the 175 microliters setting, the air volume that causes an alarm is 1.1mL (4.4mL window volume with 25% air in window volume). • For the 250 microliters setting, the air volume that causes an alarm is 1.82mL (5.2mL window volume with 35% air in window volume). • When operating in Anesthesia Mode with a single bolus air-in-line limit of 500 microliters, the accumulated air detection limit is 2.73mL (7.8mL window volume with 35% air in window volume). NOTE: The 500 microliters limit cannot be set in the Guardrails™ Editor software and is only available when the PCU is operating in Anesthesia Mode. When Disabled is selected, the Accumulated AIL detection system is turned off for all pump module infusions. When Enabled is selected, the Accumulated AIL detection system is turned on for all pump module infusions. | [9 empty] | No |  | Enabled
Air-in-Line (microliter) | Sets the upper limit for a single bolus of air to pass without alarm. This is the amount of air allowed to pass through the detector before the “Air-in-Line” alarm will sound. Five different air-in-line detection settings can be selected: 50, 75, 125, 175 or 250 mcL. The 500 microliters setting is available in Anesthesia Mode only. Warning: Ensure that Air-in-line configuration settings are appropriate for the dsesignated care area proile. Failure to do so may increase the risk of clinically significant volumes of air infusing into the patient. Size of Air bubble: Based on commonly used, standard bore administration set #2420-0500; 50mcL is approximately equivalent to 0.34" (1/4 of an inch), 75mcL is approximately equivalent to 0.51" (1/2 of an inch), 125mcL is approximately equivalent to 0.7" (3/4 of an inch), 175mcL is approximately equivalent to 1.0" and 250mcL is approximately 1.7 (1 ½ inches)". | [9 empty] | No |  | NICU: 50 - 75 PEDS: 75 - 175 Adult: 75 - 250
Auto-Restart Attempts | The Auto-Restart feature is part of the BD Alaris™ System’s Downstream Occlusion Detection system designed to minimize nuisance, patient-side occlusion alarms. It allows the infusion device to automatically continue an infusion following detection of a patient-side occlusion if downstream pressure falls to an acceptable level within a 15-second “Checking Line” period. The system can be configured to allow from 0 to 9 restart attempts within a rolling 10-minute period. If the allowable number of restarts is exceeded or if the feature is set to 0, an Occluded-Patient Side alarm occurs when the system detects downstream pressure over the pressure limit. Auto-Restart for Anesthesia Mode is set to 9 and not configurable. | [9 empty] | No |  | NICU - 0-1 PICU - 1-3 Pediatrics - 4-7 OB/L&D - 9 Clinic - 9 MedSurg - 5-8 ICU - 3-5 Oncology - 1-3
KVO Rate Adjust (mL/hr) | The KVO rate range is from 0.1 to 20 mL/h. This will determine the rate of fluid flow after “Infusion Complete” has occurred, based on VTBI reaching zero mL. This prevents the infusion from stopping. Users will hear a continuous audible alert until the KVO is addressed. | [9 empty] | No |  | NICU - 10 Pediatrics - 10-20 Adult - 20
Max Occlusion Pressure (mmHg) | This is the patient-side occlusion alarm threshold for the Selectable Pressure Mode. The instrument defaults to this setting when New Patient is selected. This is the maximum allowable value the user can select when in Selectable Pressure Mode. 50–525 mmHg may be selected in 25 mmHg increments. | [9 empty] | No |  | 525
Default Pressure (mmHg) | The starting occlusion pressure limit for the Selectable Pressure Mode. The range is 50 to 525 mmHg and may be selected in 25 mmHg increments; the default pressure must be less than or equal to the maximum occlusion pressure. | [9 empty] | Yes |  | 525
Maximum Rate (mL/hr) | The maximum rate range is 0.1–999 mL/h. • Volumes of 0.1–99.9 mL/h may be selected in 0.1 mL increments. • Rates of 100–999 mL/h are selected in 1 mL increments. | [9 empty] | No |  | NICU - 20 mL/kg/hr Pediatrics <50kg - 20mL/kg/h Pediatrics >50kg - 999 mL/h Adult - 999 mL/h
Maximum VTBI (mL) | The volume-to-be-infused (VTBI) range is 0.1–9,999 mL. • Volumes of 0.1–9.99 mL may be selected in 0.01 mL increments. • Volumes of 10–999.9 mL may be selected in 0.1 mL increments. • VTBIs of 1,000–9,999 mL are selected in 1 mL increments. | [9 empty] | No |  | NICU - 20 mL/kg Pediatrics <50kg - 20 mL/kg/hr Pediatrics >50kg - TPN Adult - TPN
Pressure Mode Lock Status | Unlocked – Both pressure modes (Pump or Selectable) are available Locked – Only the default Pressure Mode Selection is available | [9 empty] | No |  | Unlocked
Pressure Mode Selection | There are two pressure modes available to determine the patient-side occlusion limit. The selected mode will be the default setting for the Profile. • Pump Mode - the downstream occlusion alarm threshold is 525 mmHg at flow rates of 30 mL/h or greater. For flow rates less than 30 mL/h, the occlusion pressure is rate-dependent to ensure rapid response to occlusions. • Selectable Pressure Mode - the downstream occlusion alarm threshold can be adjusted by the user in 25 mmHg increments from 50 mmHg up to the value set as the Profile’s Maximum Occlusion Pressure. | [9 empty] | Pump, Selectable |  | Pump
Secondary Infusion | This mode is designed to support automatic secondary infusions (piggybacking) in the same instrument channel. When the secondary VTBI reaches 0, an auditory tone is sounded, indicating completion of the secondary infusion. The primary infusion resumes automatically. When the instrument is programmed and delivering in the secondary mode, the primary infusion is temporarily stopped and fluid is drawn from the secondary container. Delivery from the primary container resumes when the fluid level in the secondary line is level with the fluid in the primary container. | [9 empty] | No |  | Dictated by the drug library build
SEC-PRI Audio Alert | When enabled, six audio tones will sound to signify the secondary infusion VTBI has completed and the primary infusion programming has resumed; no intervention is required from the clinician. When disabled, the primary infusion will resume when the secondary infusion has completed with no audio alert. | [9 empty] | No |  | NICU - Disable PICU - Disable Pediatric - Disable Quiet Zone - Disable
BD-10751
BD, the BD Logo and Alaris™, Guardrails™ are trademarks of Becton, Dickinson and Company. © 2021 BD and its subsidiaries. All rights reserved
Note on D2: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Can those HUGE "Page X" be removed?

## SYRINGE
BD Alaris™ Guardrails™ Editor v12.1.2 Configuration Worksheets
BD Alaris™ Syringe Module
BD Alaris™ Products
Fields highlighted the same color are related features AND grey highlights mean they are enabled with Drug Setups
Profile Names
BD Alaris™ Syringe Module |  |  | Adult | Infusion Center | Nursery | Pediatrics
Feature | Default Setting | Options | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
All Mode | Disabled | Enabled, Disabled | Disabled | Disabled | Enabled | Disabled
Auto Occlusion Pressure | Enabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
Back-Off After Occlusion | Enabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
Fast Start | Enabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
Occlusion Pressure Limit with Disc (mmHg) | 1000 | 25 - 1000 | 1000 | 1000 | 1000 | 1000
Occlusion Pressure Limit - No Disc | High | Low, Med, High | High | High | High | High
Priming | Disabled | Enabled, Disabled | Disabled | Disabled | Enabled | Disabled
KVO Option | Disabled | Enabled, Disabled | Disabled | Disabled | Enabled | Disabled
KVO Rate Adjust (mL/h) | 1 | 0.1 - 2.5 | n/a | n/a | 2 | n/a
KVO Volume Adjust (%) | 5 | 0.5 - 5 | n/a | n/a | 5 | n/a
Near End Of Infusion (Continuous) | Disabled | Enabled, Disabled | Disabled | Disabled | Enabled | Disabled
Alert Time Continuous (minutes) | 60 | 1 - 60 | n/a | n/a | 2 | n/a
Near End of Infusion (Intermittent) | Disabled | Enabled, Disabled | Disabled | Disabled | Enabled | Disabled
Alert Time Intermittent (minutes) | 15 | 1 - 60 | n/a | n/a | 2 | n/a
BD Alaris™ Syringe Module Settings
Setting | Definition | [9 empty] | Changed at Bedside |  | Consider for Default |  | NOTES
All Mode | When enabled, the default value displayed in the VTBI field is “ALL” with the intent that the entire contents of the syringe is to be delivered. When disabled, the default value displayed in the VTBI field is blank and the user will enter a VTBI. | [9 empty] | Yes |  | Enabled
Auto Occlusion Pressure | When enabled and a pressure sensing disc is in use, the Auto Pressure option is displayed in the Pressure Limit screen. Auto Pressure automatically sets the alarm limit for a shorter time to alarm, as follows: • If current pressure is 100 mmHg or less, system adds 30 mmHg to current pressure to create a new alarm limit. • If current pressure is greater than 100 mmHg, system adds 30% to current pressure to create a new alarm limit. | [9 empty] | No If using PSD it is applicable; if not using PSD, it is not applicable |  | Enabled
Back-Off After Occlusion | This feature is only available when the administration set in use has a pressure-sensing disc. When enabled, the motor reverses plunger movement during an occlusion until the pressure returns to preocclusion levels, utomatically reducing bolus after occlusion. | [9 empty] | No If using PSD it is applicable; if not using PSD, it is not applicable |  | Enabled
Fast-Start | When Fast-Start is enabled and an administration set having a pressure sensing disc is used, the instrument runs at an increased rate when an infusion is first started, taking up any slack in the drive mechanism. The net result is that infusions start delivering fluid to the patient within a few seconds, even at very low flow rates. | [9 empty] | No If using PSD it is applicable; if not using PSD, it is not applicable |  | Enabled
Occlusion Pressure Limit with Disc (mmHg) | When a pressure sensing disc is in use, this establishes a limit on the maximum pressure in the line before the occlusion warning alarms and the infusion stops. The downstream occlusion alarm threshold can be set from 25 to 1000 mmHg. NOTE: The occlusion pressure limit becomes the default pressure. The user can manually adjust the pressure lower but cannot adjust higher than the established occlusion pressure limit. | [9 empty] | Yes, when using PSD |  | 1000
Occlusion Pressure Limit No Disc | This establishes a limit on the maximum pressure in the line before the occlusion warning alarms and the infusion stops when a pressure-sensing disc set is NOT used. The downstream occlusion alarm threshold can be set to low (approximately 200 mmHg), medium (approximately 500 mmHg), or high (approximately 800 mmHg). NOTE: The occlusion pressure limit becomes the default pressure. The user can manually adjust the pressure lower but cannot adjust higher than the established occlusion pressure limit. | [9 empty] | Yes |  | High
Priming | The Priming option allows a limited volume of fluid to be delivered to prime the administration set prior to being connected to a patient or after changing a syringe. When priming, a single continuous press of the PRIME soft key on the PCU delivers up to 2 mL of priming fluid. | [9 empty] | No |  | Enabled
KVO Option | The Selectable KVO option allows some infusions to automatically switch into KVO mode upon completion. The KVO option setting cannot be changed after the instrument is powered on and a Profile selected. If the KVO feature is enabled a reserve of fluid identified by the KVO Rate Adjust is used to administer the KVO infusion. If the intent is to deliver the entire contents of the syringe, the KVO option should be disabled. | [9 empty] | No |  | Disabled
KVO Rate Adjust (mL/h) | This option is used to select the KVO rate (0.1–2.5 mL/h). This determines the rate of fluid flow after Infusion Complete has occurred. The KVO Rate shall be less than or equal to the Maximum Rate. NOTE: If flow rate is less than 0.1 mL/h with 1 mL or 3 mL syringe, the KVO rate will stay the same rate down to 0.01 mL/h. | [9 empty] | No |  | ___
KVO Volume Adjust (%) | The KVO volume trigger range is 0.5 to 5% of remaining fluid volume. It is the established amount remaining in the syringe, which will trigger the device to go into KVO mode. This amount is calculated from syringe size and, once this remaining amount is reached, the infusion will continue at the established KVO rate while providing an “Infusion Complete” alert. | [9 empty] | No |  | ___
Near End Of Infusion (Continuous) | The near end of infusion (NEOI) option is a configurable alert that indicates that the syringe infusion is nearing a completion of the volume to be infused. The NEOI can be configured separately between continuous type infusions and intermittent infusion. The NEOI option for Continuous, Fluids and Basic infusions allows an alert to be configured to sound before the infusion is completed. | [9 empty] | No |  | Enabled
Alert Time Continuous (minutes) | The NEOI alert occurs at the configured time (1–60 minutes) A separate alert time can be set between continuous and intermittent infusions. | [9 empty] | No |  | ___
Near End Of Infusion (Intermittent) | The near end of infusion (NEOI) option is a configurable alert that indicates that the syringe infusion is nearing a completion of the volume to be infused. The NEOI can be configured separately between continuous type infusions and intermittent infusion. The NEOI option for Intermittent infusions allows an alert to be configured to sound before the infusion is completed. When NEOI Intermittent is enabled, NEOI for Continuous is automatically enabled. | [9 empty] | No |  | Disabled
Alert Time Intermittent (minutes) | The NEOI alert occurs at the configured time (1–60 minutes) A separate alert time can be set between continuous and intermittent infusions. | [9 empty] | No |  | n/a
BD-10751
BD, the BD Logo and Alaris™, Guardrails™ are trademarks of Becton, Dickinson and Company. © 2021 BD and its subsidiaries. All rights reserved

## SHARED INFUSION SETTINGS
BD Alaris™ Guardrails™ Editor v12.x Configuration Worksheets
Shared Infusion Settings
BD Alaris™ Products
Fields highlighted the same color are related features AND grey highlights mean they are enabled with Drug Setups
Profile Names
BD Alaris™ Syringe module |  |  | Adult | Infusion Center | Nursery | Pediatrics
Feature | Default Setting | Options | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
Shared Infusion Settings (BD Alaris™ Pump Module and Syringe Module)
Delay Start Options | Disabled | Enabled, Disabled | Disabled | Disabled | Disabled | Disabled
Delay Start Callback | After | None, Before, Before & After, After | n/a | n/a | n/a | n/a
Drug Calculation | Enabled | Enabled, Disabled | Enabled | Disabled | Enabled | Enabled
Drug Calculation Bolus Dose | Disabled | Enabled, Disabled | Enabled | n/a | Enabled | Enabled
Pressure Dynamic Display | Disabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
Volume/Duration | Enabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
Shared Infusion Settings (BD Alaris™ PCA Module and Syringe Module) |  |  | Adult | Infusion Center | Nursery | Pediatrics
Maximum Rate (mL/h) | 999 | 0.1 - 999 | 999 | 999 | 25 | 999
Near End Of Infusion (NEOI) Snooze | Disabled | Enabled, Disabled | Disabled | Disabled | Enabled | Disabled
Snooze Time (minutes) | 15 | 5, 10 or 15 | n/a | n/a | 5 | n/a
Shared Infusion Settings (BD Alaris™ Pump and Syringe Modules)
Setting | Definition | [9 empty] | Changed at Bedside |  | Consider for Default |  | NOTES
Delay Start Options | This feature allows the system to be programmed to delay the start of an infusion for 1 minute or up to 11 hours and 59 minutes. | [9 empty] | No |  | Enabled
Delay Start Callback | When programming a Delay for infusion, a Callback can be scheduled for that infusion. There are three types of Callback: • Before gives an alert when delay is completed and infusion needs to be initiated. • After gives an alert when delayed infusion has completed. • Before and After gives an alert when delay is completed and infusion needs to be initiated and when delayed infusion has completed. The default Callback (After), or the Callback for the current Profile appear in the main display. Clinician can schedule a different callback on the Alaris System Delay Options page. NOTE: With Guardrails sofware version 12.1.0 and higher, all infusion programmed with a Delay will revert to KVO when the infusion is complete regardless of the callback chosen. | [9 empty] | Yes |  | Before and After
Drug Calculation | The Drug Calculation mode allows one of the following: • Entry of drug dose (BD Alaris™ System calculated correct flow rate to achieve desired dose). • Entry of flow rate (BD Alaris™ System calculated corresponding drug dose). NOTE: When enabled, the No-Guardrails™, Basic Infusion Clinical Advisory will be displayed when a Drug Calculation is programmed | [9 empty] | No |  | Enabled
Drug Calculation Bolus Dose | The Drug Calculation Bolus Dose feature allows a bolus infusion to be programmed using either the Guardrails™ Drug Library or the Drug Calculation feature. The bolus infusion can be programmed with or without a continuous infusion following the bolus. Medications programmed using the Drug Calculator feature are not protected by Guardrails™. | [9 empty] | No |  | Enabled
Pressure Dynamic Display | The Dynamic Pressure Display is located in the Main Disply of the PCU, just below the channel status information. If enabled, it provides a graphical display of the patient-side (downstream) pressure and the current patient-side operating pressure for that channel. This feature is available when using the Pump module and when using a pressure sensing disc set with the Syringe module. | [9 empty] | No |  | Enabled
Volume/Duration | The Volume/Duration infusion feature allows a volume-to-be-infused (VTBI) and duration (infusion time) to be programmed; the flow rate is then automatically calculated. | [9 empty] | No |  | Enabled
Shared Infusion Settings (BD Alaris™ PCA and Syringe Modules)
Maximum Rate | The maximum rate that can be infused. Range 0.1–999 mL/h. The Maximum Rate is also dependent on the size of the compatible syringe installed within the Alaris System. • 0.1–9.9 in 0.01 mL/h increments. • 10–99.9 in 0.1 mL/h increments. • 100–999 in 1 mL/h increments. | [9 empty] | No |  | 999
NEOI Snooze | An optional alert that provides an audio-only tone when a Near End of Infusion alert has been previously silenced. This alert can be configured to give an audio-only tone every five, 10 or 15 minutes. | [9 empty] | No |  | Disabled
NEOI Snooze Time (min) | The NEOI reminder signal can be configured in 5, 10, or 15 minute intervals. | [9 empty] | No |  | n/a
BD-10751
BD, the BD Logo and Alaris™, Guardrails™ are trademarks of Becton, Dickinson and Company. © 2020 BD and its subsidiaries. All rights reserved

## PCA
BD Alaris™ Guardrails™ Editor v12.1.2 Configuration Worksheets
BD Alaris™ PCA Module
BD Alaris™ Products
Fields highlighted the same color are related features AND grey highlights mean they are enabled with Drug Setups
Profile Names
BD Alaris™ PCA module |  |  | Adult | Infusion Center | Nursery | Pediatrics
Feature | Default Setting | Options | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
Bolus Delivery Rate (mL/h) | 150 | 75 - 500 | 150 | 150 | 150 | 150
Dose Request Cord (Profile) | 2 | 1, 2, or 3 | 2 | 2 | 2 | 2
Max Dose Limit | Enabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
Time Window (h) | 4 | 1, 2, or 4 | 1 | 4 | 4 | 1
Module Location Enforcement | Enabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
Near End of Infusion (NEOI) | Disabled | Enabled, Disabled | Enabled | Disabled | Disabled | Enabled
NEOI Alert Time % | 5 | 5 - 25 | 5 | n/a | n/a | 5
Pressure Limit | High | High, Med, Low | High | High | High | High
Priming | Enabled | Enabled, Disabled | Enabled | Enabled | Enabled | Enabled
Security Access Level | 1 | 1, 2, or 3 | 1 | 1 | 1 | 1
Security Code | None | 0000 - 9999 | n/a | n/a | n/a | n/a
BD Alaris™ PCA Module Settings
Setting | Definition | [9 empty] | Changed at Bedside |  | Consider for Default |  | NOTES
Bolus Delivery Rate (mL/h) | The rate at which PCA, Bolus, and Loading Doses is infused. The range is 75 to 500 mL/hr. The rate is limited by syringe size. Warning: Avoid delivering extrememly small volume bolus or loading doses (less than 0.5mL/h). Under infusion can occur under these conditions. | [9 empty] | No |  | Adult -150 Pediatrics - 75 Opioid Naïve - 75 Opioid Tolerant - 150 PCEA - <150
Dose Request Cord (Profiles) | The dose request cord can be configured to provide both audio and visual (LED indicator) notification to the patient in response to the requesting a dose. One of three profiles can be selected: | [9 empty] | Yes |  | 1
Max Dose Limit (Max Accumulated Dose Limit) | Optional configuration that limits total amount of drug allowed to be delivered to the patient in a defined time period (1, 2, or 4 hours). NOTE: This setting should be configured in the data set before the drug library is developed. After drugs are in the Profile PCA Drug Library, the Max Accumulated Dose Limit cannot be changed. NOTE: This optional setting applies to all drug setups within the Profile PCA Drug Library. | [9 empty] | If enabled, the user can still choose not to use this parameter when programming the PCA module. |  | ___
Time Window (h) | In conjunction with the Max Dose Limit, the Time Window (1, 2 or 4 hours) establishes a limit on the total amount of drug delivered to the patient for the defined time window. | [9 empty] | No |  | ___
Module Location Enforcement | Tamper resistant security feature that ensures the PCA Module is in a tamper evident position. When enabled, PCA Module must be located to the direct right of the PCU before programming an infusion. | [9 empty] | No |  | Enabled
Near End of Infusion (NEOI) | The NEOI option allows an audio alert to be configured based on a percentage of the syringe size. | [9 empty] | No |  | Enabled
NEOI Alert Time (%) | Alert Time is the trigger for the Near End of Infusion alert. Alerts the clinician when 5 to 25% of the remaining volume in the PCA syringe remains, based on syringe size. | [9 empty] | No |  | 5%
Pressure Limit | This establishes a limit on the maximum pressure in the line before the occlusion warning alarms and the infusion stops. The downstream occlusion alarm threshold can be set to low (approximately 200 mmHg), medium (approximately 500 mmHg), or high (approximately 800 mmHg). NOTE: Syringe variability may impact occlusion pressure sensing. The variability may reduce the device's time to alarm and/or may require that a higher alarm pressure limit be programmed. | [9 empty] | No |  | Adult/Pediatric - High Epidural - Medium
Priming | The Priming option allows a limited volume of fluid to be delivered to prime the administration set prior to being connected to a patient or after changing a syringe. When priming, a single continuous press of the PRIME soft key delivers up to 2 mL of priming fluid. | [9 empty] | No |  | Enabled
Security Access Level | The security access level can be configured to provide varying levels of access to the device. Security access is accomplished through either the use of a key or a four-digit authorization code. | [9 empty] | No |  | 1
Security Code | If Security Access Level 2 or 3 was selected above, a four-digit, numeric code will need to be provided to clinicians to access the PCU for administering bolus doses and making subsequent programming changes. | [9 empty] | No |  | Customizable
BD-10751
BD, the BD Logo and Alaris™, Guardrails™ are trademarks of Becton, Dickinson and Company. © 2021 BD and its subsidiaries. All rights reserved

## PCA PAUSE PROTOCOL
BD Alaris™ Guardrails™ Editor v12.1.2 Configuration Worksheets
BD Alaris™ PCA Pause Protocol
BD Alaris™ Products
Fields highlighted the same color are related features AND grey highlights mean they are enabled with Drug Setups
Profile Names
BD Alaris™ PCA Pause Protocol |  |  | Adult | Infusion Center | Nursery | Pediatrics
Feature | Default Setting | Options | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
PCA Pause Protocol Enabled | Disabled | Enabled, Disabled | Enabled | Disabled | Disabled | Disabled
Monitoring Module Attach Enforcement | Disabled | Enabled, Disabled | Disabled | n/a | n/a | n/a
PCA Pause Protocol (Alarm Message) | “PCA Infusion has paused due to a decline in respiratory status. Check Patient” | Select Edit Text to change | PCA Infusion has paused due to a decline in respiratory status. Check Patient! | n/a | n/a | n/a
BD PCA Pause Protocol
Setting | Definition | [9 empty] | Changed at Bedside |  | Consider for Default |  | NOTES
PCA Pause Protocol | PCA pause protocol can be enabled for the EtCO2 Module. When PCA pause protocol is enabled, the PCA pauses upon exceeding and sustaining the PCA pause protocol alarm limits established for the EtCO2 Module. Once PCA Pause Protocol is enabled, you are allowed to enter PCA/Monitoring Alarm Limits for Low Respiratory Rate and the default or initial alarm value. See EtCO2 Module settings to establish the PCA pause protocol alarm limits. | [9 empty] | Yes |  | n/a
Monitoring Module Attach Enforcement | This feature allows the facility to provide a notice to the user to attach the monitoring module EtCO2 for the PCA pause protocol features to be enabled. When enabled, the user must attach the applicable monitoring module, or modules, prior to moving forward with PCA programming. | [9 empty] | No |  | Disabled
PCA Pause Protocol (Alarm Message) | This text alarm notice alerts the user that PCA infusion has paused. This default text alarm message is recommended but editable to enforce hospital protocol. The default alarm text reads: PCA infusion has paused due to decline in respiratory status. Check patient. The PCA Pause Protocol message on PCU displays one line for the header and has remaining 7 lines for the PCA Pause Protocol message. | [9 empty] | No |  | System Default
BD-10751
BD, the BD Logo and Alaris™, Guardrails™ are trademarks of Becton, Dickinson and Company. © 2021 BD and its subsidiaries. All rights reserved

## EtCO2
BD Alaris™ Guardrails™ Editor v12.1.2 Configuration Worksheets
BD Alaris™ EtCO2 Module
BD Alaris™ Products
Fields highlighted the same color are related features AND grey highlights mean they are enabled with Drug Setups
Profile Names
BD Alaris™ EtCO2 module |  |  | Adult | Infusion Center | Nursery | Pediatrics
Feature | Default Setting | Options | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
Alarm Limit Type | Adult | Adult/Neonatal | Adult | Adult | Adult | Adult
EtCO2 High (mmHg) | Adult = 60 | 5 - 99 | 60 | 60 | 60 | 60
Neonatal = 60
EtCO2 Low (mmHg) | Adult = 10 | 0 - 98 | 10 | 10 | 10 | 10
Neonatal = 10
Respiratory Rate High (bpm) | Adult = 35 | 1 - 150 | 35 | 35 | 35 | 35
Neonatal = 80
Respiratory Rate Low (bpm) | Adult = 6 | 0 - 149 | 7 | 6 | 6 | 6
Neonatal = 12
No Breath Alarm (Seconds) | Adult = 30 | 10 - 60 | 30 | 30 | 30 | 30
Neonatal = 20
FiCO2 High (mmHg) | 8 | 2 - 99 | 8 | 8 | 8 | 8
Waveform Time Scale (Seconds) | Adult = 5 | 5 or 10 | 10 | 10 | 10 | 10
Neonatal = 5
PCA Pause Protocol |  |  | Adult | Infusion Center | Nursery | Pediatrics
Resp. Rate Lower Limit (bpm) | Off | 0 – 149 | 5 | n/a | n/a | n/a
Initial Value (bpm) | Off | 0 – 149 | 6 | n/a | n/a | n/a
BD Alaris™ EtCO2 Module Settings
Setting | Definition | [9 empty] | Changed at Bedside |  | Consider for Default |  | NOTES
Alarm Limit Type | Enable/Disable for Adult/Neonatal | [9 empty] | No |  | Adult
EtCO2 High (mmHg) | High default alarm limit for EtCO2. | [9 empty] | Yes |  | 60
EtCO2 Low (mmHg) | Low default alarm limit for EtCO2. | [9 empty] | Yes |  | 10
Respiratory Rate High | High default alarm limit for Resp. Rate. | [9 empty] | Yes |  | 35
Respiratory Rate Low | Low default alarm limit for Resp. Rate. | [9 empty] | Yes |  | 6 - 8
No Breath Alarm (Seconds) | The period of time no breath can be detected before an alarm. | [9 empty] | Yes |  | 20 - 30
FiCO2 High (mmHg) | High default alarm limit for Fraction of Inspired CO2. | [9 empty] | Yes |  | 8
Waveform Time Scale (Seconds) | The amount of waveform time (in seconds) displayed on the EtCO2 Main display. | [9 empty] | Yes |  | ___
PCA Pause Protocol
Setting | Definition | [9 empty] | Changed at Bedside |  | Consider for Default
PCA Pause Protocol | The following parameters can be set if PCA Pause Protocol is enabled for the EtCO2 module:
Resp. Rate Lower Limit (bpm) | This value represents the lowest alarm value established by hospital protocol. This value is not editable by the user. | [9 empty] | No |  | 5
Initial Value (bpm) | This is the default alarm value established by hospital protocol. If this value is exceeded and sustained, the PCA infusion will pause and alarm. This default value must be higher than the Resp. Rate Lower Limit and is editable by the user. | [9 empty] | Yes |  | 6
BD-10751
BD, the BD Logo and Alaris™, Guardrails™ are trademarks of Becton, Dickinson and Company. © 2021 BD and its subsidiaries. All rights reserved

## Appendix
BD Alaris™ Guardrails™ Editor v12.x Configuration Worksheets
BD Alaris™ Products
Default Clinical Advisories List
Advisory Name (Up to 20 characters) | Clinical Advisory Text (Up to approximately 145 characters) |  |  |  | Approval Signatures
2nd RN to verify | 2 RNs to verify and document concentration and programming prior to infusion.
Assess and Document | Assess and document pain level, sedation level, pulse, O2 sat and respiratory rate Q2-4H PRN.
Check with MD | Check with MD before exceeding limits
Cisatracurium | *Train of Four Required* BEFORE beginning infusion obtain baseline TOF. Mechanical ventilation and concurrent sedation mandatory.
Code only drugs | This selection is reserved for use during Code Pink, ACLS events. DO NOT USE for routine infusions!
Crash Cart | Have crash cart or emergency equipment available.
Extravasation | Watch IV site carefully for signs of extravasation.
Heparin | Monitor PTT. Do not administer other drugs through same IV line.
Hypersensitivity | Risk of hypersensitivity. Check vital signs frequently during infusion.
In-line filter | Infusion requires the use of an in-line filter.
Lipids | Infuse with 1.2 micron filter. Do not infuse other drugs through this line.
Monitor BP | Monitor blood pressure continuously during infusion.
Monitor O2 & Cardiac | Continuously monitor respiratory and cardiac function during infusion.
Monitor Sedation | Monitor respiratory rate and sedation level. Titrate to Ramsay scale per orders
Monitor Vitals | Closely monitor vital signs during infusion.
Monitor Vitals & O2 | During infusion, monitor vital signs and oxygen saturation continuously.
Monitoring required | Monitoring module required.
O2 Depression | Drug may cause respiratory depression. Monitor respiratory function closely during infusion.
Preferred Route | Please verify preferred route of infusion for this drug.
Premeds | Check for dexamethasone premeds. Use nitro/taxol tubing. Use in-line filter.
TPN | Infuse with 0.22 micron filter. Do not infuse other drugs through same IV line.
Unit policy | Please refer to unit policy prior to infusion.
Vesicant | Drug is a vesicant. Check blood return.
Vesicant-Cent. Line | Drug is a vesicant. Central line required for infusion.
New or Modified Advisory Name | Clinical Advisory Text (Up to approximately 145 characters) |  |  |  | Approval Signatures
Default Therapy List
Therapy Name (Up to 20 characters) | Matches Known Patient Care Standard/Protocol/Practice |  |  |  | Approval Signatures
24 hour infusion
BMT
Central
CODE only
DVT
ECMO
Epidural Route
Every 3 weeks
Every 4 hours
Every 6 hours
Every 8 hours
First Dose
High Dose
Induction
Intravenous Route
Loading Dose
Maintenance Dose
MI
PE
Peripheral
Postpartum
Renal
Solid Tumor
Stroke
Weekly
New or Modified Therapy Name | Matches Known Patient Care Standard/Protocol/Practice |  |  |  | Approval Signatures
Master Syringe List
Syringe Module Syringes |  |  | PCA Module Syringes
Included? | Syringe Name |  | Included? | Syringe Name
BD Plastipak |  |  | BD Plastipak
| BD Plastipak 1 mL |  |  | BD Plastipak 20 mL
| BD Plastipak 3 mL |  |  | BD Plastipak 30 mL
| BD Plastipak 5 mL |  |  | BD Plastipak 50/60 mL
| BD Plastipak 10 mL |  | Monoject
| BD Plastipak 20 mL |  |  | Monoject 20 mL
| BD Plastipak 30 mL |  |  | Monoject 35 mL
| BD Plastipak 50/60 mL |  |  | Monoject 60 mL
Monoject |  |  | Terumo
| Monoject 3 mL |  |  | Terumo 20 mL
| Monoject 6 mL |  |  | Terumo 30 mL
| Monoject 12 mL |  |  | Terumo 60 mL
| Monoject 20 mL |  | IMS
| Monoject 35 mL |  |  | IMS 30 mL Prefill
Monoject 60 mL
Terumo |  |  | NOTES:
Terumo 3 mL
Terumo 5 mL
Terumo 10 mL
Terumo 20 mL
Terumo 30 mL
Terumo 60 mL
Default Channel Labels List
AMNIOTIC INFUSION | MEDIAL LUMEN | NOTES:
ARTERIAL LINE | MEDIAL PORT
BLADDER IRRIGATION | MIDLINE CATHETER
BLOOD PRODUCTS | PA LINE
BLUE PORT | PERCUTANEOUS LINE
CENTRAL LINE | PERIPHERAL ART LINE
CRRT | PERIPHERAL IV
CSF DRAINAGE | PERIPHERAL LINE
CVVHDF DIALYSATE | PICC LINE
DIALYSATE | PICC LINE #1
DIALYSIS CATHETER | PICC LINE #2
DISTAL LUMEN | PROXIMAL LUMEN
EPIDURAL | RA LINE
FEMORAL LINE | RED PORT
FLUSH | REPLACEMENT FLUID
GROSHONG | RESEARCH AGENT
HICKMAN | SUBCLAVIAN
IMPLANTED PORT | SUBCUTANEOUS
INTRAJUGULAR | SWAN
INTRAOSSEOUS | TRANSHEPATIC LINE
INTRAPERITONEAL | UAC
INVESTIGATIONAL DRUG | UVC
IRRIGATION *NOT IV* | WHITE PORT
IV FLUIDS + VITAMINS | WOUND IRRIGATION
LA LINE
New or Modified Channel Label List
Channel Label â Profile à | Adult | Infusion Center | Nursery | Pediatrics
1
2
3
4
5
6
7
8
9
10
Channel Labels should not include drug or fluid names. They are intended to provide a channel identifier for route of administration or type of IV line, such as epidural or distal lumen. There is no BD Alaris™ Guardrails™ Suite MX protection with the use of Channel Labels.
BD-10751
BD, the BD Logo and Alaris™, Guardrails™ are trademarks of Becton, Dickinson and Company. © 2021 BD and its subsidiaries. All rights reserved
CAUTION