## Data Set Properties
Data Set Properties
Data Set Name: | Conway2022v35is
Hospital Name: | Conway Medical Center
Status: | Released
Data Set ID: | 02a0bf1fb-R

## ContBolus Non-Anesth Drugs
Continuous/Bolus Non-Anesthesia Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount (- - = custom concentration) | Drug Amount Units | Diluent Volume (- - = custom concentration) | Diluent Volume Units | Final Concentration | Final Concentration Units | Module Availability | Dosing Units | Dose Limits Soft Min | Dose Limits Soft Max | Dose Limits Hard Max | Dose Initial Value | Bolus Dose Units | Bolus Dose Limits Soft Min | Bolus Dose Limits Soft Max | Bolus Dose Limits Hard Max | Bolus Dose Limits Initial Value | Bolus Dose Admin. Rate Units | Bolus Dose Admin. Rate Limits Soft Min | Bolus Dose Admin. Rate Limits Soft Max | Bolus Dose Admin. Rate Limits Hard Max | Bolus Dose Admin. Rate Initial Value | Conc. Limit Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max | Clinical Advisory
1 | Adult | adenosine (CATH LAB) |  | 90 | mg | 90 | mL | 1 | mg/mL | Pump | mcg/kg/min | 135 | 145 |  | 140 | [10 empty] | n/a
2 | Adult | alteplase -stroke | Stroke < 101 kg | - - | mg | - - | mL |  | mg/mL | Pump | mg/kg/h | 0.7 | 0.9 |  | 0.81 | [10 empty] | mg/mL | 0.85 | 0.99 | 1.01 | Alteplase wt based
3 | Adult | alteplase -stroke | Stroke 101kg or more | 81 | mg | 81 | mL | 1 | mg/mL | Pump | mg/h | 71 | 90 |  | 81 | [10 empty] | n/a
4 | Adult | amiodarone Drip |  | 360 | mg | 200 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.51 | [11 empty] | n/a |  |  |  | In-line filter
5 | Adult | amiodarone Drip |  | 450 | mg | 250 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.51 | [11 empty] | n/a |  |  |  | In-line filter
6 | Adult | amiodarone Drip |  | 900 | mg | 500 | mL | 1.8 | mg/mL | Pump | mg/min | 0.49 | 1 | 1.51 | [11 empty] | n/a |  |  |  | In-line filter
7 | Adult | argatroban | .Standard Dose | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 10 | 11.5 | 2 | [10 empty] | n/a
8 | Adult | argatroban | .Standard Dose | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 10 | 11.5 | 2 | [10 empty] | n/a
9 | Adult | argatroban | Hepatic Dose | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.1 | 5 | 10.1 | 0.5 | [10 empty] | n/a
10 | Adult | argatroban | Hepatic Dose | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.1 | 5 | 10.1 | 0.5 | [10 empty] | n/a
11 | Adult | argatroban | Undergoing PCI | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 15 | 30 | 34 | 25 | mcg/kg | 149 | 351 |  |  | mcg/kg/min | 29 | 117 |  |  | n/a
12 | Adult | argatroban | Undergoing PCI | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 15 | 30 | 34 | 25 | mcg/kg | 149 | 351 |  |  | mcg/kg/min | 29 | 117 |  |  | n/a
13 | Adult | atracurium |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 0.5 | 29.5 | 39 |  | mg/kg | 0.07 | 0.5 | 0.58 |  | mg/kg/min | 0.07 | 0.5 | 0.58 |  | mcg/mL | 200 | 5000 | 5100 | Neuromuscular BA
14 | Adult | bivalirudin | .Standard Dose | 250 | mg | 50 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.1 | 1.8 | 2 | 1.75 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.76 |  | 0.75 | n/a
15 | Adult | bivalirudin | .Standard Dose | 500 | mg | 100 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.1 | 1.8 | 2 | 1.75 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.76 |  | 0.75 | n/a
16 | Adult | bivalirudin | .Standard Dose | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.1 | 1.8 | 2 | 1.75 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.76 |  | 0.75 | n/a
17 | Adult | bivalirudin | ESRD | 250 | mg | 50 | mL | 5 | mg/mL | Pump | mcg/kg/h | 0.1 | 1.2 | 1.4 | 0.25 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
18 | Adult | bivalirudin | ESRD | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/h | 0.1 | 1.2 | 1.4 | 0.25 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
19 | Adult | bivalirudin | Renal Dose CrCl < 30 | 250 | mg | 50 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.24 | 1.2 | 1.4 | 1 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
20 | Adult | bivalirudin | Renal Dose CrCl < 30 | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.24 | 1.2 | 1.4 | 1 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
21 | Adult | bivalirudin | Renal Dose CrCl < 30 | 500 | mg | 100 | mL | 5 | mg/mL | Pump | mg/kg/h | 0.24 | 1.2 | 1.4 | 1 | mg/kg | 0.3 | 0.8 | 0.81 | 0.75 | mg/kg/min | 0.3 | 0.8 |  | 0.75 | n/a
22 | Adult | bumetanide continuos |  | - - | mg | - - | mL |  | mg/mL | Pump | mg/h | 0.1 | 2 |  | 0.5 | mg | 0.25 | 4 |  |  | mg/min | 0.25 | 4 |  |  | mg/mL | 0.009 | 0.01 | 0.26 | Protect From Light
23 | Adult | cangrelor |  | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 2 | 4.2 | 4.6 | 4 | mcg/kg | 26 |  | 35 | 30 | mcg/kg/min | 26 |  | 35 |  | n/a
24 | Adult | cisatracurium |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | mcg/kg/min | 0.5 | 5 | 7 | 3 | [10 empty] | n/a |  |  |  | Neuromuscular BA
25 | Adult | conivaptan |  | 20 | mg | 100 | mL | 0.2 | mg/mL | Pump | mg/h | 0.6 | 1.7 | 2 | 0.8 | mg | 10 | 20 |  | 20 | mg/min | 0.333 | 0.7 |  | 0.667 | n/a
26 | Adult | conivaptan |  | 20 | mg | 250 | mL | 0.08 | mg/mL | Pump | mg/h | 0.6 | 1.7 | 2 | 0.8 | mg | 10 | 20 |  | 20 | mg/min | 0.333 | 0.7 |  | 0.667 | n/a
27 | Adult | dexmedetomidine |  | 200 | mcg | 50 | mL | 4 | mcg/mL | Pump | mcg/kg/h | 0.05 | 0.7 | 1.5 | 0.2 | mcg/kg | 0.9 | 1.1 | 2 | 1 | mcg/kg/min | 0.01 | 0.11 | 0.2 | 0.1 | n/a
28 | Adult | dexmedetomidine |  | 400 | mcg | 100 | mL | 4 | mcg/mL | Pump | mcg/kg/h | 0.05 | 0.7 | 1.5 | 0.2 | mcg/kg | 0.9 | 1.1 | 2 | 1 | mcg/kg/min | 0.01 | 0.11 | 0.2 | 0.1 | n/a
29 | Adult | dilTIAZem |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | mg/h | 0.5 | 15 | 25 | 5 | [10 empty] | n/a
30 | Adult | dilTIAZem |  | 125 | mg | 125 | mL | 1 | mg/mL | Pump | mg/h | 0.5 | 15 | 25 | 5 | [10 empty] | n/a
31 | Adult | DOBUTamine | ..Standard | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | 2 | [10 empty] | n/a
32 | Adult | DOBUTamine | ..Standard | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | 2 | [10 empty] | n/a
33 | Adult | DOBUTamine | .Double Concentrated | 500 | mg | 250 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | 2 | [10 empty] | n/a |  |  |  | Double Concentrated
34 | Adult | DOBUTamine | .Double Concentrated | 1000 | mg | 500 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | 2 | [10 empty] | n/a |  |  |  | Double Concentrated
35 | Adult | DOBUTamine | Echo Stress Test | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 4.9 | 50 |  | 5 | [10 empty] | mcg/mL | 900 |  | 5000
36 | Adult | DOBUTamine | Stress Test | 500 | mg | 250 | mL | 2 | mg/mL | Pump | mcg/kg/min | 5 | 50 |  | 10 | [10 empty] | n/a
37 | Adult | DOPamine | ..Standard | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 0.9 | 20 | 30 | 2 | [10 empty] | n/a
38 | Adult | DOPamine | .Double Concentrated | 800 | mg | 250 | mL | 3.2 | mg/mL | Pump | mcg/kg/min | 0.9 | 20 | 30 | 2 | [10 empty] | n/a |  |  |  | Double Concentrated
39 | Adult | DOPamine | Renal Dose | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 0.9 | 5 | 6 | 1 | [10 empty] | n/a
40 | Adult | doxapram |  | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mg/min | 0.1 | 6 | [12 empty] | n/a
41 | Adult | doxapram |  | - - | mg | - - | mL |  | mg/mL | Pump | mg/min | 0.1 | 6 | [12 empty] | mg/mL | 0.89 | 0.9 | 2.1
42 | Adult | epidural w/ fentanyl |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 0.01 | 12 | 15 |  | unit | 2 | 7 | 8 | 5 | unit/min | 4 | 14 | 16 | 10 | n/a |  |  |  | 1 unit = 1 mL
43 | Adult | epidural w/o fentany |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 0.01 | 12 | 15 |  | unit | 2 | 7 | 8 | 5 | unit/min | 4 | 14 | 16 | 10 | n/a |  |  |  | 1 unit = 1 mL
44 | Adult | EPINEPHrine | ..Standard | 5 | mg | 250 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 20 | 30 | [11 empty] | n/a |  |  |  | Dedicated IV Line
45 | Adult | EPINEPHrine | ..Standard | 10 | mg | 500 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 20 | 30 | [11 empty] | n/a |  |  |  | Dedicated IV Line
46 | Adult | EPINEPHrine | Cardiac Arrest | 30 | mg | 250 | mL | 0.12 | mg/mL | Pump | mcg/min | 0.33 | 40 |  | 33.33 | [10 empty] | n/a |  |  |  | Dedicated IV Line
47 | Adult | epoprostenol |  | 500 | mcg | 100 | mL | 5 | mcg/mL | Pump | nanogram/kg/min | 1 | 40 |  | 2 | [10 empty] | n/a
48 | Adult | epoprostenol |  | 1000 | mcg | 100 | mL | 10 | mcg/mL | Pump | nanogram/kg/min | 1 | 40 |  | 2 | [10 empty] | n/a
49 | Adult | epoprostenol |  | 3000 | mcg | 100 | mL | 30 | mcg/mL | Pump | nanogram/kg/min | 1 | 40 |  | 2 | [10 empty] | n/a
50 | Adult | epoprostenol |  | - - | mcg | - - | mL |  | mcg/mL | Pump | nanogram/kg/min | 1 | 40 |  | 2 | [10 empty] | nanogram/mL | 2000 | 3000 | 50000
51 | Adult | eptifibatide | .Standard & wt>121kg | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mg/h | 13 | 17 |  | 15 | [10 empty] | n/a |  |  |  | SCr<2.0
52 | Adult | eptifibatide | .Standard & wt>121kg | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mg/h | 13 | 17 |  | 15 | [10 empty] | n/a |  |  |  | SCr<2.0
53 | Adult | eptifibatide | .Standard Dose | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mcg/kg/min | 0.5 | 2.2 | 2.5 | 2 | [10 empty] | n/a |  |  |  | SCr<2.0
54 | Adult | eptifibatide | .Standard Dose | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 2.2 | 2.5 | 2 | [10 empty] | n/a |  |  |  | SCr<2.0
55 | Adult | eptifibatide | Renal Dose | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mcg/kg/min | 0.25 | 1.2 | 1.4 | 1 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
56 | Adult | eptifibatide | Renal Dose | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.25 | 1.2 | 1.4 | 1 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
57 | Adult | eptifibatide | Renal Dose wt >121kg | 75 | mg | 100 | mL | 0.75 | mg/mL | Pump | mg/h | 6 | 8 | 9 | 7.5 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
58 | Adult | eptifibatide | Renal Dose wt >121kg | 200 | mg | 100 | mL | 2 | mg/mL | Pump | mg/h | 6 | 8 | 9 | 7.5 | [10 empty] | n/a |  |  |  | SCr 2.0-4.0
59 | Adult | esmolol |  | 2500 | mg | 250 | mL | 10 | mg/mL | Pump | mcg/kg/min | 25 | 300 |  | 100 | mcg/kg | 100 | 500 |  | 250 | mcg/kg/min | 100 | 500 |  | 200 | n/a
60 | Adult | fenoldopam |  | 10 | mg | 250 | mL | 0.04 | mg/mL | Pump | mcg/kg/min | 0.03 | 1.6 | 1.8 | 0.1 | [10 empty] | n/a
61 | Adult | fentaNYL |  | 2500 | mcg | 100 | mL | 25 | mcg/mL | Pump | mcg/h | 0.0001 | 250 | 750 |  | mcg | 0.0001 | 250 | 500 |  | mcg/min | 0.0001 | 250 | 500 |  | n/a
62 | Adult | furosemide |  | 300 | mg | 30 | mL | 10 | mg/mL | Pump | mg/h | 1 | 10 | [12 empty] | n/a |  |  |  | Protect From Light
63 | Adult | furosemide |  | - - | mg | - - | mL |  | mg/mL | Pump | mg/h | 1 | 10 | [12 empty] | mg/mL | 0.5 | 1 | 10 | Protect From Light
64 | Adult | glucagon |  | 10 | mg | 50 | mL | 0.2 | mg/mL | Pump | mg/h | 0.1 | 10 | [12 empty] | n/a
65 | Adult | glucagon |  | 10 | mg | 100 | mL | 0.1 | mg/mL | Pump | mg/h | 0.1 | 10 | [12 empty] | n/a
66 | Adult | heparin | ..Std. Cardiac/PVD | 25000 | unit | 500 | mL | 50 | unit/mL | Pump | unit/kg/h | 1 | 100 |  | 12 | unit/kg | 30 |  | 82 |  | unit/kg/min | 30 |  | 82 |  | n/a |  |  |  | Heparin
67 | Adult | heparin | .VTE (DVT, PE) | 25000 | unit | 500 | mL | 50 | unit/mL | Pump | unit/kg/h | 1 | 100 |  | 18 | unit/kg | 30 |  | 82 |  | unit/kg/min | 30 |  | 82 |  | n/a |  |  |  | Heparin
68 | Adult | heparin | Cerebrovascular | 25000 | unit | 500 | mL | 50 | unit/mL | Pump | unit/kg/h | 1 | 100 |  | 15 | [10 empty] | n/a |  |  |  | Heparin
69 | Adult | insulin |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 1 | 22.1 | 30 |  | unit | 1 | 12 |  |  | unit/min | 1 |  | 10 | 10 | n/a |  |  |  | 2nd RN to verify
70 | Adult | isoproterenol |  | 1 | mg | 250 | mL | 0.004 | mg/mL | Pump | mcg/min | 0.15 | 4 | 30 | 1 | [10 empty] | n/a
71 | Adult | isoproterenol |  | 2 | mg | 500 | mL | 0.004 | mg/mL | Pump | mcg/min | 0.15 | 4 | 30 | 1 | [10 empty] | n/a
72 | Adult | ketamine ANES |  | - - | mg | - - | mL |  | mg/mL | Pump | mg/kg/h | 0.05 | 0.5 | 1 | 0.1 | [10 empty] | mg/mL | 0.1 | 0.5 | 1.5
73 | Adult | labetalol |  | 200 | mg | 200 | mL | 1 | mg/mL | Pump | mg/min | 0.05 | 2 |  | 2 | [10 empty] | n/a
74 | Adult | lidocaine |  | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | mg/min | 0.5 | 4 | 4.6 | 2 | [10 empty] | n/a
75 | Adult | mag OB maintenance |  | 20 | gram | 500 | mL | 0.04 | gram/mL | Pump | gram/h | 0.5 | 3 | 5 | [11 empty] | n/a
76 | Adult | midazolam |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | mg/h | 0.0001 | 5 | 10 |  | mg | 0.0001 | 1 | 5 |  | mg/min | 0.0001 | 1 | 5 |  | n/a
77 | Adult | milrinone |  | 40 | mg | 200 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 0.19 | 0.75 |  | 0.5 | mcg/kg | 25 | 50.1 |  | 50 | mcg/kg/min | 1 | 5.01 |  | 5 | n/a
78 | Adult | naloxone |  | 2 | mg | 500 | mL | 0.004 | mg/mL | Pump | mg/h | 0.25 | 2 | [12 empty] | n/a
79 | Adult | nesiritide |  | 1.5 | mg | 250 | mL | 0.006 | mg/mL | Pump | mcg/kg/min | 0.005 | 0.03 | 0.035 | 0.01 | [10 empty] | n/a
80 | Adult | niCARdipine |  | 20 | mg | 200 | mL | 0.1 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
81 | Adult | niCARdipine |  | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
82 | Adult | niCARdipine |  | 40 | mg | 200 | mL | 0.2 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
83 | Adult | niCARdipine |  | 50 | mg | 500 | mL | 0.1 | mg/mL | Pump | mg/h | 1 | 15 |  | 5 | [10 empty] | n/a
84 | Adult | nitroglycerin |  | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mcg/min | 5 | 150 | 400 | 5 | [10 empty] | n/a
85 | Adult | nitroglycerin |  | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mcg/min | 5 | 150 | 400 | 5 | [10 empty] | n/a
86 | Adult | nitroprusside |  | 50 | mg | 500 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.1 | 3 | 10 | 0.3 | [10 empty] | n/a |  |  |  | Protect From Light
87 | Adult | norepinephrine | ..Standard | 4 | mg | 250 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 45 | [12 empty] | n/a |  |  |  | Levophed
88 | Adult | norepinephrine | ..Standard | 8 | mg | 500 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 45 | [12 empty] | n/a |  |  |  | Levophed
89 | Adult | norepinephrine | .Double Concentrated | 16 | mg | 500 | mL | 0.032 | mg/mL | Pump | mcg/min | 0.3 | 45 | [12 empty] | n/a |  |  |  | Double Concentrated
90 | Adult | octreotide |  | 600 | mcg | 240 | mL | 2.5 | mcg/mL | Pump | mcg/h | 25 | 50 |  |  | mcg | 25 | 150 |  | 50 | mcg/min | 1 | 50 |  | 10 | n/a
91 | Adult | oxytocin | Induction | 30 | unit | 500 | mL | 0.06 | unit/mL | Pump | milliunit/min | 1 | 41 |  |  | unit | 0.9 | 30.1 |  | 10 | unit/min | 0.1 | 1.01 |  | 0.999 | n/a
92 | Adult | oxytocin | Postpartum | 10 | unit | 1000 | mL | 0.01 | unit/mL | Pump | milliunit/min | 1 | 75 | [12 empty] | n/a
93 | Adult | oxytocin | Postpartum | 20 | unit | 1000 | mL | 0.02 | unit/mL | Pump | milliunit/min | 1 | 75 | [12 empty] | n/a
94 | Adult | oxytocin | Postpartum | 30 | unit | 1000 | mL | 0.03 | unit/mL | Pump | milliunit/min | 1 | 75 | [12 empty] | n/a
95 | Adult | oxytocin | Postpartum | 40 | unit | 1000 | mL | 0.04 | unit/mL | Pump | milliunit/min | 1 | 75 | [12 empty] | n/a
96 | Adult | pancuronium |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 0.01 | 2 | 3 |  | mcg/kg | 0.06 | 0.1 |  |  | mcg/kg/min | 0.06 | 1 |  |  | mcg/mL | 99 | 100 | 2001 | Neuromuscular BA
97 | Adult | pantoprazole gtt |  | 40 | mg | 100 | mL | 0.4 | mg/mL | Pump | mg/h | 7 | 9 |  | 8 | [10 empty] | n/a
98 | Adult | pantoprazole gtt |  | 96 | mg | 240 | mL | 0.4 | mg/mL | Pump | mg/h | 7 | 9 |  | 8 | [10 empty] | n/a
99 | Adult | phenylephrine |  | 10 | mg | 250 | mL | 0.04 | mg/mL | Pump | mcg/min | 10 | 300 |  | 100 | [10 empty] | n/a |  |  |  | Phenylephrine
100 | Adult | phenylephrine |  | 20 | mg | 500 | mL | 0.04 | mg/mL | Pump | mcg/min | 10 | 300 |  | 100 | [10 empty] | n/a |  |  |  | Phenylephrine
101 | Adult | phenylephrine |  | 40 | mg | 1000 | mL | 0.04 | mg/mL | Pump | mcg/min | 10 | 300 |  | 100 | [10 empty] | n/a |  |  |  | Phenylephrine
102 | Adult | phenylephrine ANES |  | 10 | mg | 101 | mL | 0.099 | mg/mL | Pump | mcg/min | 40 | 120 | 180 | [11 empty] | n/a
103 | Adult | phenylephrine ANES |  | 10 | mg | 251 | mL | 0.0398 | mg/mL | Pump | mcg/min | 40 | 120 | 180 | [11 empty] | n/a
104 | Adult | procainamide | Loading Dose | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | mg/min | 15 | 50 |  | 33.33 | [10 empty] | n/a
105 | Adult | procainamide | Maintenance Dose | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | mg/kg/min | 0.01 | 0.06 |  | 0.05 | [10 empty] | n/a
106 | Adult | propofol |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | mcg/kg/min | 1 | 55 | 100 | 5 | [10 empty] | n/a
107 | Adult | rocuronium |  | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 4 | 12 | 16 | 8 | [10 empty] | n/a |  |  |  | Neuromuscular BA
108 | Adult | tirofiban | CrCl <= 60 | 12.5 | mg | 250 | mL | 0.05 | mg/mL | Pump | mcg/kg/min | 0.065 |  | 0.085 | 0.075 | mcg/kg | 22 |  | 28 | 25 | mcg/kg/min | 4 |  | 6 | 5 | n/a |  |  |  | Bleeding Precautions
109 | Adult | tirofiban | CrCl > 60 | 12.5 | mg | 250 | mL | 0.05 | mg/mL | Pump | mcg/kg/min | 0.13 |  | 0.17 | 0.15 | mcg/kg | 22 |  | 28 | 25 | mcg/kg/min | 4 |  | 6 | 5 | n/a |  |  |  | Bleeding Precautions
110 | Adult | vasopressin | ..Standard | 40 | unit | 100 | mL | 0.4 | unit/mL | Pump | unit/min | 0.01 | 0.05 | 0.06 | 0.04 | [10 empty] | n/a
111 | Adult | vasopressin | .Double Concentrated | 80 | unit | 100 | mL | 0.8 | unit/mL | Pump | unit/min | 0.01 | 0.05 | 0.06 | 0.04 | [10 empty] | n/a |  |  |  | Double Concentrated
112 | Adult | vecuronium |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | mcg/kg/min | 0.01 | 1.3 | 1.67 | 1 | mg | 5 | 10 |  |  | mg/min | 5 | 100 |  |  | n/a |  |  |  | Neuromuscular BA
113 | Adult | zz amiodarone |  | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | mg/min | 0.5 | 1 | [12 empty] | n/a |  |  |  | Code only drugs
114 | Adult | zz amiodarone |  | 900 | mg | 500 | mL | 1.8 | mg/mL | Pump | mg/min | 0.5 | 1 | [12 empty] | n/a |  |  |  | Code only drugs
115 | Adult | zz DOBUTamine | ..Standard | 500 | mg | 250 | mL | 2 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 |  | 2 | [10 empty] | n/a |  |  |  | Code only drugs
116 | Adult | zz DOBUTamine | .Double Concentrated | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 |  | 2 | [10 empty] | n/a |  |  |  | Code only drugs
117 | Adult | zz DOPamine | ..Standard | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 2 | 20 |  | 2 | [10 empty] | n/a |  |  |  | Code only drugs
118 | Adult | zz DOPamine | .Double Concentrated | 800 | mg | 250 | mL | 3.2 | mg/mL | Pump | mcg/kg/min | 2 | 20 |  | 2 | [10 empty] | n/a |  |  |  | Code only drugs
119 | Adult | zz EPINEPHrine |  | 5 | mg | 250 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 20 | [12 empty] | n/a |  |  |  | Code only drugs
120 | Adult | zz EPINEPHrine |  | 10 | mg | 500 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 20 | [12 empty] | n/a |  |  |  | Code only drugs
121 | Adult | zz lidocaine |  | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | mg/min | 0.5 | 4 |  | 2 | [10 empty] | n/a |  |  |  | Code only drugs
122 | Adult | zz nitroglycerin |  | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mcg/min | 1 | 150 |  | 5 | [10 empty] | n/a |  |  |  | Code only drugs
123 | Adult | zz nitroprusside |  | 50 | mg | 250 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 0.1 | 3 |  | 0.3 | [10 empty] | n/a |  |  |  | Code only drugs
124 | Adult | zz norepinephrine | ..Standard | 4 | mg | 250 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 45 | [12 empty] | n/a |  |  |  | Code only drugs
125 | Adult | zz norepinephrine | ..Standard | 8 | mg | 500 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 45 | [12 empty] | n/a |  |  |  | Code only drugs
126 | Adult | zz norepinephrine | .Double Concentrated | 16 | mg | 500 | mL | 0.032 | mg/mL | Pump | mcg/min | 0.3 | 45 | [12 empty] | n/a |  |  |  | Code only drugs
127 | Adult | zz phenylephrine |  | 10 | mg | 250 | mL | 0.04 | mg/mL | Pump | mcg/min | 10 | 120 |  | 100 | [10 empty] | n/a |  |  |  | Code only drugs
128 | Nursery | No drugs exist in this section
129 | ONC Infusion Center | No drugs exist in this section
130 | OutPt InfusionCenter | No drugs exist in this section
131 | Pediatrics | alprostadil |  | 500 | mcg | 250 | mL | 2 | mcg/mL | Pump | mcg/kg/min | 0.05 | 0.1 | 0.4 | 0.05 | [10 empty] | n/a
132 | Pediatrics | alprostadil |  | - - | mcg | - - | mL |  | mcg/mL | Pump | mcg/kg/min | 0.05 | 0.1 | 0.4 | 0.05 | [10 empty] | mcg/mL | 0.5 | 1 | 20
133 | Pediatrics | aminophylline |  | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.2 | 1.1 |  |  | mg/kg | 0.001 |  | 6 |  | mg/kg/min | 0.19 |  | 0.36 | 0.2 | n/a
134 | Pediatrics | DOBUTamine |  | 250 | mg | 250 | mL | 1 | mg/mL | Pump | mcg/kg/min | 0.5 | 20 | 30 | 2 | [10 empty] | n/a
135 | Pediatrics | DOPamine |  | 400 | mg | 250 | mL | 1.6 | mg/mL | Pump | mcg/kg/min | 2 |  | 20 | 2 | [10 empty] | n/a
136 | Pediatrics | EPINEPHrine |  | 5 | mg | 250 | mL | 0.02 | mg/mL | Pump | mcg/min | 0.33 | 20 | 30 | [11 empty] | n/a
137 | Pediatrics | insulin |  | 100 | unit | 100 | mL | 1 | unit/mL | Pump | unit/h | 1 | 15 | 30 |  | unit | 1 | 12 |  |  | unit/min | 1 |  | 10 | 10 | n/a
138 | Pediatrics | lidocaine |  | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | mg/min | 0.5 | 4 | 4.6 | 2 | [10 empty] | n/a
139 | Pediatrics | milrinone |  | 20 | mg | 100 | mL | 0.2 | mg/mL | Pump | mcg/kg/min | 0.125 | 0.75 |  | 0.5 | mcg/kg | 25 | 50 |  |  | mcg/kg/min | 1 | 5 |  |  | n/a
140 | Pediatrics | naloxone |  | 2 | mg | 500 | mL | 0.004 | mg/mL | Pump | mg/h | 0.25 | 2 | [12 empty] | n/a
141 | Pediatrics | nitroglycerin |  | 25 | mg | 250 | mL | 0.1 | mg/mL | Pump | mcg/min | 5 | 150 | 400 | 5 | [10 empty] | n/a
142 | Pediatrics | nitroprusside |  | 50 | mg | 500 | mL | 0.1 | mg/mL | Pump | mcg/kg/min | 0.1 | 8 | 8.1 | 1 | [10 empty] | n/a |  |  |  | Protect From Light
143 | Pediatrics | norepinephrine |  | 8 | mg | 500 | mL | 0.016 | mg/mL | Pump | mcg/min | 0.3 | 45 | [12 empty] | n/a
Note on AE4: Dose and volume to be infused need to be set according to weight based protocol to avoid potential over infusion of entire contents of bag.
Note on AE6: Infusion requires the use of an in-line filter.
Note on AE7: Infusion requires the use of an in-line filter.
Note on AE8: Infusion requires the use of an in-line filter.
Note on J15: Units Only concentration and Dosing Units have different units of measure
Note on AA15: Units Only concentration and Dosing Units have different units of measure
Note on AE15: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE24: Protect this drug from light.
Note on AE26: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE35: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on AE36: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on J37: Units Only concentration and Dosing Units have different units of measure
Note on AA37: Units Only concentration and Dosing Units have different units of measure
Note on AE40: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on AE44: 1 unit = 1 mL
Note on AE45: 1 unit = 1 mL
Note on AE46: Do not infuse any other drugs through this line.
Note on AE47: Do not infuse any other drugs through this line.
Note on AE48: Do not infuse any other drugs through this line.
Note on J52: Units Only concentration and Dosing Units have different units of measure
Note on AA52: Units Only concentration and Dosing Units have different units of measure
Note on AE53: SCr < 2.0
Note on AE54: SCr < 2.0
Note on AE55: SCr < 2.0
Note on AE56: SCr < 2.0
Note on AE57: SCr 2.0 - 4.0
Note on AE58: SCr 2.0 - 4.0
Note on AE59: SCr 2.0 - 4.0
Note on AE60: SCr 2.0 - 4.0
Note on AE64: Protect this drug from light.
Note on AE65: Protect this drug from light.
Note on AE68: Monitor PTT. Do not administer other drugs through same IV line.
Note on AE69: Monitor PTT. Do not administer other drugs through same IV line.
Note on AE70: Monitor PTT. Do not administer other drugs through same IV line.
Note on AE71: 2 RNs to verify and document concentration and programming prior to infusion.
Note on AE88: Protect this drug from light.
Note on AE89: *****Must have physician order for rates greater than 60 mcg/min*****
Note on AE90: *****Must have physician order for rates greater than 60 mcg/min*****
Note on AE91: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on J98: Units Only concentration and Dosing Units have different units of measure
Note on AA98: Units Only concentration and Dosing Units have different units of measure
Note on AE98: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE101: *****Must have physician order for rates greater than 300 mcg/min*****
Note on AE102: *****Must have physician order for rates greater than 300 mcg/min*****
Note on AE103: *****Must have physician order for rates greater than 300 mcg/min*****
Note on AE109: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE110: Monitor for signs of bleeding.
Note on AE111: Monitor for signs of bleeding.
Note on AE113: ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
Note on AE114: *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
Note on AE115: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE116: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE117: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE118: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE119: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE120: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE121: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE122: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE123: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE124: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE125: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE126: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE127: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE128: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE129: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on AE144: Protect this drug from light.

## ContBolus Anesth Drugs
Continuous/Bolus Anesthesia Only Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount (- - = custom concentration) | Drug Amount Units | Diluent Volume (- - = custom concentration) | Diluent Volume Units | Final Concentration | Final Concentration Units | Module Availability | Dosing Units | Dose Limits Soft Min | Dose Limits Soft Max | Dose Limits Hard Max | Dose Initial Value | Bolus Dose Units | Bolus Dose Limits Soft Min | Bolus Dose Limits Soft Max | Bolus Dose Limits Hard Max | Bolus Dose Limits Initial Value | Bolus Dose Admin. Rate Units | Bolus Dose Admin. Rate Limits Soft Min | Bolus Dose Admin. Rate Limits Soft Max | Bolus Dose Admin. Rate Limits Hard Max | Bolus Dose Admin. Rate Initial Value | Conc. Limit Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max
1 | Adult | ketamine ANES |  | 50 | mg | 50 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.05 | 0.5 | 1 | 0.1 | [10 empty] | n/a
2 | Adult | ketamine ANES |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | mg/kg/h | 0.05 | 0.5 | 1 | 0.1 | [10 empty] | n/a
3 | Adult | phenylephrine ANES |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/min | 40 | 120 | 180 | [11 empty] | mcg/mL | 10 | 38 | 110
4 | Adult | propofol ANES |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | mcg/kg/min | 1 |  | 300 | 5 | mg | 19 | 71 |  | 30 | mg/min | 100 | 140 |  | 120 | n/a
5 | Adult | propofol ANES |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | mcg/kg/min | 1 |  | 300 | 5 | mg | 19 | 71 |  | 30 | mg/min | 100 | 140 |  | 120 | n/a
6 | Adult | remifentanil ANES |  | 1000 | mcg | 40 | mL | 25 | mcg/mL | Pump | mcg/kg/min | 0.025 | 2 | 4 | 0.05 | mcg/kg | 0.1 |  | 2 | 1 | mcg/min | 1 | 2000 | 4000 | 1000 | n/a
7 | Adult | remifentanil ANES |  | 2000 | mcg | 80 | mL | 25 | mcg/mL | Pump | mcg/kg/min | 0.025 | 2 | 4 | 0.05 | mcg/kg | 0.1 |  | 2 | 1 | mcg/min | 1 | 2000 | 4000 | 1000 | n/a
8 | Adult | remifentanil ANES |  | - - | mcg | - - | mL |  | mcg/mL | Pump | mcg/kg/min | 0.025 | 2 | 4 | 0.05 | mcg/kg | 0.1 |  | 2 | 1 | mcg/min | 1 | 2000 | 4000 | 1000 | mcg/mL | 8.9 | 9 | 250
9 | Adult | rocuronium ANES |  | 500 | mg | 500 | mL | 1 | mg/mL | Pump | mcg/kg/min | 4 | 12 | 16 | 8 | [10 empty] | n/a
10 | Adult | rocuronium ANES |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 4 | 12 | 16 | 8 | [10 empty] | mcg/mL | 890 | 900 | 1100
11 | Adult | SUFentanil ANES |  | 100 | mcg | 100 | mL | 1 | mcg/mL | Pump | mcg/kg/h | 0.1 | 1 | [12 empty] | n/a
12 | Adult | SUFentanil ANES |  | - - | mcg | - - | mL |  | mcg/mL | Pump | mcg/kg/h | 0.1 | 1 | [12 empty] | mcg/mL | 0.1 | 0.2 | 4
13 | Adult | vecuronium |  | - - | mg | - - | mL |  | mg/mL | Pump | mcg/kg/min | 0.01 | 1.3 | 1.67 | 1 | mg | 5 | 10 |  |  | mg/min | 5 | 100 |  |  | mcg/mL | 440 | 450 | 550
14 | Nursery | No drugs exist in this section
15 | ONC Infusion Center | No drugs exist in this section
16 | OutPt InfusionCenter | No drugs exist in this section
17 | Pediatrics | No drugs exist in this section
Note on J5: Units Only concentration and Dosing Units have different units of measure
Note on AA5: Units Only concentration and Dosing Units have different units of measure
Note on J12: Units Only concentration and Dosing Units have different units of measure
Note on AA12: Units Only concentration and Dosing Units have different units of measure
Note on J15: Units Only concentration and Dosing Units have different units of measure
Note on AA15: Units Only concentration and Dosing Units have different units of measure

## Intermittent Drugs
Intermittent Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount (- - = custom concentration) | Drug Amount Units | Diluent Volume (- - = custom concentration) | Diluent Volume Units | Final Concentration | Final Concentration Units | Module Availability | Available As | Duration Limits (hh:mm) Hard Min | Duration Limits (hh:mm) Soft Min | Duration Limits (hh:mm) Soft Max | Duration (hh:mm) Initial Value | Dosing Units | Dose Limits Soft Min | Dose Limits Soft Max | Dose Limits Hard Max | Conc. Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max | Clinical Advisory
1 | Adult | acetaminophen | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 500 | 1000 |  | n/a
2 | Adult | acetaminophen | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | gram | 0.5 | 1 |  | n/a
3 | Adult | acetaZOLAMIDE |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 00:40 | 00:30 | mg | 100 | 1000 |  | n/a
4 | Adult | acetylcysteine | 1st Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg/kg | 142.5 |  | 157.5 | mg/mL | 1 |  | 75
5 | Adult | acetylcysteine | 1st Dose >100kg | 15000 | mg | 275 | mL | 54.5455 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 13500 | 16500 | 18000 | n/a
6 | Adult | acetylcysteine | 2nd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg/kg | 47.5 |  | 52.5 | mg/mL | 1 |  | 10
7 | Adult | acetylcysteine | 2nd Dose >100kg | 5000 | mg | 525 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg | 4000 | 5500 | 6500 | n/a
8 | Adult | acetylcysteine | 3rd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg/kg | 95 |  | 105 | mg/mL | 1 |  | 10
9 | Adult | acetylcysteine | 3rd Dose >100kg | 10000 | mg | 1050 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg | 9000 | 11000 | 12500 | n/a
10 | Adult | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 100 | 2000 |  | mg/mL | 0.1 | 1 | 7
11 | Adult | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg/kg | 1 | 10.1 |  | mg/mL | 0.1 | 1 | 7
12 | Adult | albumin 25% |  | 12.5 | gram | 50 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:16 | 01:00 | 00:50 | gram | 12.5 | 25 |  | n/a
13 | Adult | albumin 25% |  | 25 | gram | 100 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:32 | 02:00 | 01:40 | gram | 12.5 | 25 |  | n/a
14 | Adult | alteplase-non-stroke | Central IV Occlusion | 2 | mg | 50 | mL | 0.04 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:00 | 00:15 | mg | 2 | 4 |  | n/a
15 | Adult | alteplase-non-stroke | Central IV Occlusion | 2 | mg | 100 | mL | 0.02 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:00 | 00:15 | mg | 2 | 4 |  | n/a
16 | Adult | alteplase-non-stroke | DVT / Ischemic Limb | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:00 |  | 24:00 |  | mg | 0.9 |  | 24 | mg/mL | 0.004 | 0.008 | 0.51
17 | Adult | alteplase-non-stroke | Pulmonary Embolism | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary Only |  | 01:45 | 02:15 | 02:00 | mg | 49 | 115 | 130 | n/a
18 | Adult | amikacin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 |  | mg/kg | 1 | 10 |  | mg/mL | 2 | 2.5 | 11
19 | Adult | aminocaproic acid | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 500 | 5000 |  | mg/mL | 1 | 5 | 50
20 | Adult | aminocaproic acid | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | gram | 0.5 | 5.7 |  | gram/mL | 0.001 | 0.005 | 0.05
21 | Adult | amiodarone *Loading* |  | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:08 | 00:15 | 00:10 | mg | 149 | 151 |  | n/a
22 | Adult | amiodarone *Loading* |  | 150 | mg | 103 | mL | 1.4563 | mg/mL | Pump | Primary and Secondary |  | 00:08 | 00:15 | 00:10 | mg | 149 | 151 |  | n/a
23 | Adult | ampho B (liposomal) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only |  | 01:00 | 04:00 | 02:00 | mg/kg | 2 | 7 |  | mg/mL | 0.1 | 0.2 | 2 | AmBisome
24 | Adult | amphotericin B |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 02:00 | 03:00 | 10:00 | 06:00 | mg/kg | 0.05 | 1 | 1.5 | mg/mL | 0.01 |  | 0.1
25 | Adult | ampicillin | .MG Dosing | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2500 |  | n/a
26 | Adult | ampicillin | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2500 |  | n/a
27 | Adult | ampicillin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2500 |  | n/a
28 | Adult | ampicillin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2500 |  | n/a
29 | Adult | ampicillin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2500 |  | n/a
30 | Adult | ampicillin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2500 |  | n/a
31 | Adult | ampicillin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | mg | 200 | 2500 |  | mg/mL | 0.9 | 1 | 30
32 | Adult | ampicillin | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.5 |  | n/a
33 | Adult | ampicillin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.5 |  | n/a
34 | Adult | ampicillin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.5 |  | n/a
35 | Adult | ampicillin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.5 |  | n/a
36 | Adult | ampicillin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.5 |  | n/a
37 | Adult | ampicillin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:13 | 01:00 | 00:30 | gram | 0.2 | 2.5 |  | gram/mL | 0.001 | 0.002 | 0.03
38 | Adult | ampicillin/sulbact | .MG Dosing | 1500 | mg | 50 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | mg | 1200 | 3400 |  | n/a
39 | Adult | ampicillin/sulbact | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | mg | 1200 | 3400 |  | n/a
40 | Adult | ampicillin/sulbact | .MG Dosing | 3000 | mg | 100 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:13 | 04:00 | 01:00 | mg | 1200 | 3400 |  | n/a
41 | Adult | ampicillin/sulbact | GM dosing | 1.5 | gram | 50 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | gram | 1.2 | 3.4 |  | n/a
42 | Adult | ampicillin/sulbact | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 02:00 | 00:30 | gram | 1.2 | 3.4 |  | n/a
43 | Adult | ampicillin/sulbact | GM dosing | 3 | gram | 100 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:13 | 04:00 | 01:00 | gram | 1.2 | 3.4 |  | n/a
44 | Adult | andulafungin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 04:00 | 02:00 | mg | 50 | 200 |  | mg/mL | 0.1 | 0.35 | 0.5
45 | Adult | antivenin crotalidae | Anavip | 4 | unit | 250 | mL | 0.016 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 10.1 | 11 | n/a |  |  |  | Antivenin
46 | Adult | antivenin crotalidae | Anavip | 10 | unit | 250 | mL | 0.04 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 10.1 | 11 | n/a |  |  |  | Antivenin
47 | Adult | antivenin crotalidae | Anavip | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 10.1 | 11 | unit/mL | 0.003 |  | 0.045 | Antivenin
48 | Adult | antivenin crotalidae | Crofab | 4 | unit | 250 | mL | 0.016 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 6.1 | 13 | n/a |  |  |  | Antivenin
49 | Adult | antivenin crotalidae | Crofab | 6 | unit | 250 | mL | 0.024 | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 6.1 | 13 | n/a |  |  |  | Antivenin
50 | Adult | antivenin crotalidae | Crofab | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary | 00:52 | 00:53 | 11:30 | 10:00 | unit | 1 | 6.1 | 13 | unit/mL | 0.004 |  | 0.025 | Antivenin
51 | Adult | azithromycin |  | 250 | mg | 250 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:10 | 01:00 | mg | 210 | 575 |  | n/a
52 | Adult | azithromycin |  | 500 | mg | 250 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:10 | 01:00 | mg | 210 | 575 |  | n/a
53 | Adult | aztreonam | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
54 | Adult | aztreonam | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
55 | Adult | aztreonam | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg | 250 | 2000 |  | n/a
56 | Adult | aztreonam | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg | 250 | 2000 |  | n/a
57 | Adult | aztreonam | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.25 | 2 |  | n/a
58 | Adult | aztreonam | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.25 | 2 |  | n/a
59 | Adult | aztreonam | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | gram | 0.25 | 2 |  | n/a
60 | Adult | aztreonam | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 02:00 | gram | 0.25 | 2 |  | n/a
61 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 110 | mL | 6.3636 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:20 | 06:00 | 00:22 | mg | 699 | 701 |  | n/a
62 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 160 | mL | 4.375 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 06:00 | 00:31 | mg | 699 | 701 |  | n/a
63 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 210 | mL | 3.3333 | mg/mL | Pump | Primary and Secondary | 00:36 | 00:37 | 06:00 | 00:41 | mg | 699 | 701 |  | n/a
64 | Adult | bamlaniv/estesivimab | .Standard wt >=50 kg | 700 | mg | 310 | mL | 2.2581 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 06:00 | 01:00 | mg | 699 | 701 |  | n/a
65 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 110 | mL | 6.3636 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:20 | 06:00 | 00:22 | mg | 699 | 701 |  | n/a
66 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 160 | mL | 4.375 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 06:00 | 00:31 | mg | 699 | 701 |  | n/a
67 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 210 | mL | 3.3333 | mg/mL | Pump | Primary and Secondary | 00:36 | 00:37 | 06:00 | 00:41 | mg | 699 | 701 |  | n/a
68 | Adult | bamlaniv/estesivimab | Weight < 50 kg | 700 | mg | 310 | mL | 2.2581 | mg/mL | Pump | Primary and Secondary | 00:59 | 01:00 | 06:00 | 01:10 | mg | 699 | 701 |  | n/a
69 | Adult | bamlanivimab |  | 700 | mg | 270 | mL | 2.5926 | mg/mL | Pump | Primary and Secondary | 01:00 |  | 10:00 | 01:00 | mg | 650 | 750 |  | n/a |  |  |  | Low Binding Filter
70 | Adult | caffeine citrate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 06:00 |  | mg | 1 | 1200 |  | mg/mL |  | 0.01 | 10
71 | Adult | caffeine/sod benzoa |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 04:00 | 01:00 | mg | 50 | 501 | 1001 | mg/mL |  | 0.24 | 6 | 0.22 Filter Needed
72 | Adult | calcium chloride | .MG Dosing | 1000 | mg | 60 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | mg | 100 | 2000 |  | n/a
73 | Adult | calcium chloride | .MG Dosing | 1000 | mg | 110 | mL | 9.0909 | mg/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | mg | 100 | 2000 |  | n/a
74 | Adult | calcium chloride | .MG Dosing | 2000 | mg | 120 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 | 02:00 | mg | 100 | 2000 |  | n/a
75 | Adult | calcium chloride | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 |  | mg | 100 | 2000 |  | mg/mL |  | 0.1 | 20
76 | Adult | calcium chloride | GM dosing | 1 | gram | 60 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | gram | 0.1 | 10 |  | n/a
77 | Adult | calcium chloride | GM dosing | 1 | gram | 110 | mL | 0.0091 | gram/mL | Pump | Primary and Secondary | 00:10 | 00:45 | 24:00 | 01:00 | gram | 0.1 | 10 |  | n/a
78 | Adult | calcium chloride | GM dosing | 2 | gram | 120 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 | 02:00 | gram | 0.1 | 10 |  | n/a
79 | Adult | calcium chloride | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:20 | 00:45 | 24:00 |  | gram | 0.1 | 10 |  | gram/mL |  | 0.001 | 0.04
80 | Adult | calcium gluconate | .MG Dosing | 1000 | mg | 60 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | mg | 500 | 2400 |  | n/a
81 | Adult | calcium gluconate | .MG Dosing | 1000 | mg | 110 | mL | 9.0909 | mg/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | mg | 500 | 2400 |  | n/a
82 | Adult | calcium gluconate | .MG Dosing | 2000 | mg | 70 | mL | 28.5714 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | mg | 500 | 2400 |  | n/a
83 | Adult | calcium gluconate | .MG Dosing | 2000 | mg | 120 | mL | 16.6667 | mg/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | mg | 500 | 2400 |  | n/a
84 | Adult | calcium gluconate | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:25 | 06:00 |  | mg | 500 | 2400 |  | mg/mL | 1 |  | 40
85 | Adult | calcium gluconate | GM dosing | 1 | gram | 60 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
86 | Adult | calcium gluconate | GM dosing | 1 | gram | 110 | mL | 0.0091 | gram/mL | Pump | Primary and Secondary | 00:09 | 00:15 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
87 | Adult | calcium gluconate | GM dosing | 2 | gram | 70 | mL | 0.0286 | gram/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
88 | Adult | calcium gluconate | GM dosing | 2 | gram | 120 | mL | 0.0167 | gram/mL | Pump | Primary and Secondary | 00:19 | 00:30 | 06:00 |  | gram | 0.5 | 2.4 |  | n/a
89 | Adult | calcium gluconate | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 | 00:25 | 06:00 |  | gram | 0.5 | 2.4 |  | gram/mL | 0.001 |  | 0.04
90 | Adult | casirivimab/imdevima |  | 1200 | mg | 260 | mL | 4.6154 | mg/mL | Pump | Primary and Secondary | 00:48 | 00:49 | 04:00 | 01:00 | mg | 1190 | 1201 | 1210 | n/a |  |  |  | Low Binding Filter
91 | Adult | casirivimab/imdevima |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:18 | 00:19 | 04:00 |  | mg | 1190 | 1201 | 1210 | No |  |  |  | Low Binding Filter
92 | Adult | caspofungin |  | 35 | mg | 100 | mL | 0.35 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg | 35 | 90 |  | n/a |  |  |  | Dedicated IV Line
93 | Adult | caspofungin |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg | 35 | 90 |  | n/a |  |  |  | Dedicated IV Line
94 | Adult | caspofungin |  | 70 | mg | 250 | mL | 0.28 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg | 35 | 90 |  | n/a |  |  |  | Dedicated IV Line
95 | Adult | ceFAZolin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 |  | mg | 480 | 2300 |  | n/a
96 | Adult | ceFAZolin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
97 | Adult | ceFAZolin | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
98 | Adult | ceFAZolin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
99 | Adult | ceFAZolin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | mg/mL | 5 | 10 | 40
100 | Adult | ceFAZolin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
101 | Adult | ceFAZolin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
102 | Adult | ceFAZolin | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
103 | Adult | ceFAZolin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
104 | Adult | ceFAZolin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | gram/mL | 0.005 | 0.01 | 0.04
105 | Adult | cefepime | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
106 | Adult | cefepime | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
107 | Adult | cefepime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
108 | Adult | cefepime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
109 | Adult | cefepime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
110 | Adult | cefepime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
111 | Adult | cefepime | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 4 | 5 | 40
112 | Adult | cefepime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
113 | Adult | cefepime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
114 | Adult | cefepime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
115 | Adult | cefepime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
116 | Adult | cefepime | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.04
117 | Adult | cefiderocol | .MG Dosing | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
118 | Adult | cefiderocol | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
119 | Adult | cefiderocol | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
120 | Adult | cefiderocol | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | mg | 740 | 2050 | 2051 | n/a
121 | Adult | cefiderocol | GM dosing | 0.75 | gram | 100 | mL | 0.0075 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
122 | Adult | cefiderocol | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
123 | Adult | cefiderocol | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
124 | Adult | cefiderocol | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary | 02:39 | 02:40 | 04:00 | 03:00 | gram | 0.74 | 2.05 | 2.051 | n/a
125 | Adult | cefotaxime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
126 | Adult | cefotaxime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
127 | Adult | cefotaxime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
128 | Adult | cefotaxime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
129 | Adult | cefotaxime | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 40
130 | Adult | cefotaxime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
131 | Adult | cefotaxime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
132 | Adult | cefotaxime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
133 | Adult | cefotaxime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
134 | Adult | cefotaxime | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.04
135 | Adult | cefoTEtan | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
136 | Adult | cefoTEtan | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
137 | Adult | cefoTEtan | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
138 | Adult | cefoTEtan | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
139 | Adult | cefoTEtan | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
140 | Adult | cefoTEtan | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
141 | Adult | cefoTEtan | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
142 | Adult | cefoTEtan | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
143 | Adult | cefOXitin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
144 | Adult | cefOXitin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | n/a
145 | Adult | cefOXitin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 20
146 | Adult | cefOXitin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
147 | Adult | cefOXitin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | n/a
148 | Adult | cefOXitin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.02
149 | Adult | ceftaroline |  | 200 | mg | 50 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
150 | Adult | ceftaroline |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
151 | Adult | ceftaroline |  | 300 | mg | 50 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
152 | Adult | ceftaroline |  | 300 | mg | 100 | mL | 3 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
153 | Adult | ceftaroline |  | 400 | mg | 50 | mL | 8 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
154 | Adult | ceftaroline |  | 400 | mg | 100 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
155 | Adult | ceftaroline |  | 600 | mg | 50 | mL | 12 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
156 | Adult | ceftaroline |  | 600 | mg | 100 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:07 | 01:15 | 01:00 | mg | 200 | 600 |  | n/a
157 | Adult | ceftaroline |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:07 | 04:00 | 01:00 | mg | 200 | 600 |  | mg/mL |  | 0.1 | 14
158 | Adult | cefTAZidime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
159 | Adult | cefTAZidime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
160 | Adult | cefTAZidime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
161 | Adult | cefTAZidime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | n/a
162 | Adult | cefTAZidime | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 250 | 2000 |  | mg/mL | 1 | 2.5 | 40
163 | Adult | cefTAZidime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
164 | Adult | cefTAZidime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
165 | Adult | cefTAZidime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
166 | Adult | cefTAZidime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.249 | 2 |  | n/a
167 | Adult | ceftazidime/avibacta |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 01:42 | 02:18 | 02:00 | gram | 0.93 |  | 2.8 | gram/mL | 0.009 |  | 0.051
168 | Adult | ceftolozane/tazo | .MG Dosing | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 01:30 | 01:00 | mg | 149 | 3010 |  | n/a
169 | Adult | ceftolozane/tazo | .MG Dosing | 375 | mg | 100 | mL | 3.75 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 01:30 | 01:00 | mg | 149 | 3010 |  | n/a
170 | Adult | ceftolozane/tazo | .MG Dosing | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 01:30 | 01:00 | mg | 149 | 3010 |  | n/a
171 | Adult | ceftolozane/tazo | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 01:30 | 01:00 | mg | 149 | 3010 |  | n/a
172 | Adult | ceftolozane/tazo | .MG Dosing | 3000 | mg | 100 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 01:30 | 01:00 | mg | 149 | 3010 |  | n/a
173 | Adult | ceftolozane/tazo | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 01:30 | 01:00 | mg | 149 | 3010 |  | mg/mL | 1 | 1.4 | 30.1
174 | Adult | ceftolozane/tazo | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 01:30 | 01:00 | gram | 0.149 | 3.1 |  | n/a
175 | Adult | ceftolozane/tazo | GM dosing | 3 | gram | 100 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 01:30 | 01:00 | gram | 0.149 | 3.1 |  | n/a
176 | Adult | ceftolozane/tazo | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:45 | 01:30 | 01:00 | gram | 0.149 | 3.1 |  | gram/mL | 0.001 |  | 0.031
177 | Adult | cefTRIAXone | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
178 | Adult | cefTRIAXone | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
179 | Adult | cefTRIAXone | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
180 | Adult | cefTRIAXone | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
181 | Adult | cefTRIAXone | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
182 | Adult | cefTRIAXone | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.5 | 2 |  | n/a
183 | Adult | cefTRIAXone | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.5 | 2 |  | n/a
184 | Adult | cefTRIAXone | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.5 | 2 |  | n/a
185 | Adult | cefTRIAXone | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.5 | 2 |  | n/a
186 | Adult | cefTRIAXone | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
187 | Adult | cefuroxime | .MG Dosing | 750 | mg | 50 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
188 | Adult | cefuroxime | .MG Dosing | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
189 | Adult | cefuroxime | .MG Dosing | 1500 | mg | 50 | mL | 30 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
190 | Adult | cefuroxime | .MG Dosing | 1500 | mg | 100 | mL | 15 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 750 | 1501 | 2000 | n/a
191 | Adult | cefuroxime | GM dosing | 0.75 | gram | 50 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
192 | Adult | cefuroxime | GM dosing | 0.75 | gram | 100 | mL | 0.0075 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
193 | Adult | cefuroxime | GM dosing | 1.5 | gram | 50 | mL | 0.03 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
194 | Adult | cefuroxime | GM dosing | 1.5 | gram | 100 | mL | 0.015 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.75 | 1.51 | 2 | n/a
195 | Adult | chloramphenicol |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg/kg | 3 | 25 | 30 | mg/mL | 5 | 10 | 110
196 | Adult | chlorothiazide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 1000 | 2300 | mg/mL | 1 | 2 | 20
197 | Adult | chlorproMAZINE |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 02:00 | 00:15 | mg | 1 | 51 |  | mg/mL |  | 0.001 | 1.1 | Patient Lie Down
198 | Adult | ciprofloxacin |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
199 | Adult | ciprofloxacin |  | 400 | mg | 200 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
200 | Adult | clindamycin |  | 300 | mg | 50 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
201 | Adult | clindamycin |  | 300 | mg | 100 | mL | 3 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
202 | Adult | clindamycin |  | 600 | mg | 50 | mL | 12 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
203 | Adult | clindamycin |  | 600 | mg | 100 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
204 | Adult | clindamycin |  | 900 | mg | 50 | mL | 18 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
205 | Adult | clindamycin |  | 900 | mg | 100 | mL | 9 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
206 | Adult | colistimethate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 10 | 300 |  | mg/mL | 0.1 | 0.5 | 5
207 | Adult | colistimethate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 0.3 |  | 5 | mg/mL | 0.1 | 0.5 | 5
208 | Adult | conivaptan Loading |  | 20 | mg | 100 | mL | 0.2 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 15 | 25 |  | n/a
209 | Adult | dantrolene |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg/kg | 1 | 2.5 |  | mg/mL | 0.05 | 0.1 | 100
210 | Adult | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg | 100 | 2000 |  | mg/mL | 1 | 5 | 20
211 | Adult | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg/kg | 4 | 10.5 |  | mg/mL | 1 | 5 | 20
212 | Adult | deferoxamine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 04:00 | 08:00 | 24:00 |  | mg | 500 | 6000 |  | mg/mL | 1 |  | 95
213 | Adult | deferoxamine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 04:00 | 08:00 | 24:00 |  | mg/kg | 30 | 60 |  | mg/mL | 1 |  | 95
214 | Adult | desmopressin |  | - - | mcg | - - | mL |  | mcg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mcg | 0.1 | 22 |  | mcg/mL | 0.005 | 0.01 | 4
215 | Adult | desmopressin |  | - - | mcg | - - | mL |  | mcg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mcg/kg | 0.25 | 0.41 |  | mcg/mL | 0.005 | 0.01 | 4
216 | Adult | dexamethasone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:07 | 02:00 | 00:30 | mg | 1 | 20 |  | mg/mL | 0.01 |  | 10
217 | Adult | digoxin immune fab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 10 | 800 |  | mg/mL | 0.05 | 0.1 | 1
218 | Adult | doripenem |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 100 | 500 |  | n/a
219 | Adult | doripenem |  | 250 | mg | 100 | mL | 2.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 100 | 500 |  | n/a
220 | Adult | doripenem |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 100 | 500 |  | n/a
221 | Adult | doxycycline |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
222 | Adult | doxycycline |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
223 | Adult | ertapenem | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
224 | Adult | ertapenem | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
225 | Adult | ertapenem | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
226 | Adult | ertapenem | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
227 | Adult | ertapenem | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
228 | Adult | ertapenem | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
229 | Adult | ertapenem | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
230 | Adult | ertapenem | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
231 | Adult | erythromycin |  | 250 | mg | 100 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
232 | Adult | erythromycin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
233 | Adult | erythromycin |  | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:30 | 01:00 | mg | 100 | 1000 |  | n/a
234 | Adult | erythromycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | mg/mL | 0.05 | 1 | 5
235 | Adult | famotidine |  | 20 | mg | 50 | mL | 0.4 | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg | 8 | 44 | 50 | n/a
236 | Adult | famotidine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 02:00 |  | mg | 8 | 44 | 50 | mg/mL | 0.09 | 0.1 | 3.5
237 | Adult | fat emulsion 20% |  | 25 | gram | 125 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 11:00 | 24:00 |  | gram | 24 | 51 |  | n/a
238 | Adult | fat emulsion 20% |  | 50 | gram | 250 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 11:00 | 24:00 |  | gram | 24 | 51 |  | n/a
239 | Adult | ferric carboxymaltos |  | 750 | mg | 265 | mL | 2.8302 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg | 250 | 750 | 1501 | n/a |  |  |  | Hypersensitivity
240 | Adult | ferric carboxymaltos |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg | 250 | 750 | 1501 | mg/mL | 1.9 | 2 | 4 | Hypersensitivity
241 | Adult | ferric carboxymaltos |  | 750 | mg | 265 | mL | 2.8302 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg/kg | 5 | 16 |  | n/a |  |  |  | Hypersensitivity
242 | Adult | ferric carboxymaltos |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg/kg | 5 | 16 |  | mg/mL | 1.9 | 2 | 4 | Hypersensitivity
243 | Adult | fluconazole |  | 100 | mg | 50 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 04:00 | 00:30 | mg | 50 | 801 |  | n/a
244 | Adult | fluconazole |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 04:00 | 01:00 | mg | 50 | 801 |  | n/a
245 | Adult | fluconazole |  | 400 | mg | 200 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 01:00 | 04:00 | 02:00 | mg | 50 | 801 |  | n/a
246 | Adult | fluconazole |  | 800 | mg | 400 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 02:00 | 08:00 | 04:00 | mg | 50 | 801 |  | n/a
247 | Adult | folic acid |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 0.1 | 5 |  | mg/mL | 0.001 | 0.009 | 0.06
248 | Adult | fomepizole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 9 | 16 |  | mg/mL | 0.5 | 1 | 20
249 | Adult | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:11 | 01:00 | 00:20 | mg | 50 | 1500 |  | mg/mL | 0.5 | 1 | 30
250 | Adult | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:11 | 01:00 | 00:20 | mg/kg | 1 | 20 |  | mg/mL | 0.5 | 1 | 30
251 | Adult | ganciclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg/kg | 0.625 | 8 |  | mg/mL | 0.5 | 1 | 10 | In-line filter
252 | Adult | gentamicin |  | 80 | mg | 100 | mL | 0.8 | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 | 00:30 | mg | 10 | 1000 |  | n/a
253 | Adult | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 | 01:00 | mg | 10 | 1000 |  | mg/mL | 0.05 | 0.1 | 10
254 | Adult | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 |  | 01:20 | 01:00 | mg/kg | 0.5 | 10 |  | mg/mL | 0.05 | 0.1 | 10
255 | Adult | hetastarch 6% |  | 30 | gram | 500 | mL | 0.06 | gram/mL | Pump | Primary and Secondary |  | 00:31 | 24:00 |  | gram | 29 | 31 |  | n/a
256 | Adult | hydrocortisone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:30 | mg | 100 | 1000 |  | mg/mL | 0.5 | 1 | 20
257 | Adult | ibutilide | pt weight < 60 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:07 |  | 00:30 | 00:10 | mg/kg | 0.009 | 0.011 |  | mg/mL | 0.001 |  | 0.02
258 | Adult | ibutilide | pt weight >= 60 kg | 1 | mg | 60 | mL | 0.0167 | mg/mL | Pump | Primary and Secondary | 00:07 |  | 00:30 | 00:10 | mg | 0.4 | 1.25 |  | n/a
259 | Adult | imipenem/cilastatin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 199 | 1000 |  | n/a
260 | Adult | imipenem/cilastatin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 199 | 1000 |  | n/a
261 | Adult | imipenem/cilastatin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 01:00 | mg | 199 | 1000 |  | mg/mL | 1 | 2 | 5
262 | Adult | imipenem/cilastatin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 0.5 | 1 |  | n/a
263 | Adult | imipenem/cilastatin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | gram | 0.5 | 1 |  | n/a
264 | Adult | imipenem/cilastatin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 01:00 | gram | 0.5 | 1 |  | gram/mL | 0.001 | 0.003 | 0.006
265 | Adult | immune glob (IVIG) |  | 5 | gram | 50 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:00 |  | gram | 1 | 300 |  | n/a |  |  |  | Hypersensitivity
266 | Adult | immune glob (IVIG) |  | 5 | gram | 100 | mL | 0.05 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:00 |  | gram | 1 | 300 |  | n/a |  |  |  | Hypersensitivity
267 | Adult | immune glob (IVIG) |  | 10 | gram | 100 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:40 | 08:00 |  | gram | 1 | 300 |  | n/a |  |  |  | Hypersensitivity
268 | Adult | immune glob (IVIG) |  | 10 | gram | 200 | mL | 0.05 | gram/mL | Pump | Primary and Secondary |  | 00:40 | 08:00 |  | gram | 1 | 300 |  | n/a |  |  |  | Hypersensitivity
269 | Adult | immune glob (IVIG) |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 02:00 | 25:00 |  | gram | 1 | 300 |  | gram/mL | 0.01 | 0.05 | 0.1 | Hypersensitivity
270 | Adult | immune glob (IVIG) |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 02:00 | 36:00 |  | gram/kg | 1 | 2.4 |  | gram/mL | 0.01 | 0.05 | 0.1 | Hypersensitivity
271 | Adult | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg | 50 | 2000 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
272 | Adult | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg/kg | 2 | 10 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
273 | Adult | iron sucrose |  | 200 | mg | 110 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 08:00 | 00:30 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
274 | Adult | iron sucrose |  | 500 | mg | 275 | mL | 1.8182 | mg/mL | Pump | Primary and Secondary |  | 03:30 | 08:00 | 04:00 | mg | 25 | 500 |  | n/a |  |  |  | Hypersensitivity
275 | Adult | iron sucrose |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:00 | 08:00 | 04:00 | mg | 25 | 500 |  | mg/mL | 0.1 | 0.2 | 2.1 | Hypersensitivity
276 | Adult | itraconazole |  | 200 | mg | 60 | mL | 3.3333 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 100 | 400 |  | n/a |  |  |  | In-line filter
277 | Adult | itraconazole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 100 | 400 |  | mg/mL | 3.29 | 3.3 | 3.4 | In-line filter
278 | Adult | Kcentra (Factor IX) | .INR 2 to < 4 | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | unit/kg | 24 | 26 |  | unit/mL | 19 | 20 | 30 | INR 2-4
279 | Adult | Kcentra (Factor IX) | .INR 4 - 6 | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | unit/kg | 34 | 36 |  | unit/mL | 19 | 20 | 30 | INR 4-6
280 | Adult | Kcentra (Factor IX) | INR > 6 | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | unit/kg | 49 | 53 |  | unit/mL | 19 | 20 | 30 | INR > 6
281 | Adult | Kinevac (sincalide) |  | - - | mcg | - - | mL |  | mcg/mL | Pump | Primary and Secondary |  | 00:12 | 02:00 | 00:16 | mcg/kg | 0.018 | 0.13 |  | mcg/mL | 0.01 |  | 0.5
282 | Adult | lacosamide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 25 | 200 |  | mg/mL | 0.25 | 0.5 | 4.1
283 | Adult | levETIRAcetam |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:06 | 02:00 | 00:15 | mg | 250 | 1500 |  | mg/mL | 2 | 2.5 | 15
284 | Adult | levETIRAcetam |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:06 | 02:00 | 00:15 | mg/kg | 1 | 60 |  | mg/mL |  | 0.5 | 15
285 | Adult | levoFLOXacin |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
286 | Adult | levoFLOXacin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
287 | Adult | levoFLOXacin |  | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 01:15 | 02:00 | 01:30 | mg | 250 | 750 |  | n/a
288 | Adult | linezolid |  | 600 | mg | 300 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 04:00 | 01:00 | mg | 400 | 800 |  | n/a
289 | Adult | mag sulfate |  | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:30 | 05:00 | 04:00 | gram | 0.5 | 7 | 8 | n/a
290 | Adult | mag sulfate |  | 4 | gram | 100 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 09:00 |  | gram | 0.5 | 7 | 8 | n/a
291 | Adult | mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:30 | 12:00 |  | gram | 0.5 | 7 | 8 | gram/mL | 0.005 | 0.01 | 0.12
292 | Adult | mannitol 20% |  | 100 | gram | 500 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 00:31 | 04:00 |  | gram | 10 | 100 |  | n/a |  |  |  | In-line filter
293 | Adult | mannitol 20% |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:31 | 08:00 |  | gram | 10 | 100 |  | gram/mL | 0.15 | 0.18 | 0.21 | In-line filter
294 | Adult | mannitol 20% |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:31 | 08:00 |  | gram/kg | 0.1 | 2.1 |  | gram/mL | 0.15 | 0.18 | 0.21 | In-line filter
295 | Adult | meropenem | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
296 | Adult | meropenem | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
297 | Adult | meropenem | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
298 | Adult | meropenem | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
299 | Adult | meropenem | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
300 | Adult | meropenem | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.5 | 2 |  | n/a
301 | Adult | meropenem | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.5 | 2 |  | n/a
302 | Adult | meropenem | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.5 | 2 |  | n/a
303 | Adult | meropenem | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.5 | 2 |  | n/a
304 | Adult | meropenem | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.5 | 2 |  | n/a
305 | Adult | meropenem/vaborbacta |  | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
306 | Adult | meropenem/vaborbacta |  | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
307 | Adult | meropenem/vaborbacta |  | 2 | gram | 250 | mL | 0.008 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
308 | Adult | meropenem/vaborbacta |  | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
309 | Adult | meropenem/vaborbacta |  | 4 | gram | 250 | mL | 0.016 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
310 | Adult | meropenem/vaborbacta |  | 4 | gram | 500 | mL | 0.008 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
311 | Adult | meropenem/vaborbacta |  | 4 | gram | 1000 | mL | 0.004 | gram/mL | Pump | Primary and Secondary |  | 02:45 | 06:00 | 03:00 | gram | 0.9 | 4.1 |  | n/a
312 | Adult | methocarbamol |  | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:10 |  | 04:00 |  | mg | 249 |  | 1001 | n/a
313 | Adult | methocarbamol |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 04:00 |  | mg | 249 |  | 1001 | mg/mL | 0.5 | 1 | 15
314 | Adult | methylPREDNISolone |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | n/a
315 | Adult | methylPREDNISolone |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 125 | 2000 |  | n/a
316 | Adult | methylPREDNISolone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | mg/mL | 0.5 | 1 | 15
317 | Adult | metroNIDAZOLE |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 250 | 1000 |  | n/a
318 | Adult | metroNIDAZOLE |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 250 | 1000 |  | n/a
319 | Adult | micafungin |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg | 50 | 150 |  | n/a
320 | Adult | micafungin |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg | 50 | 150 |  | n/a
321 | Adult | micafungin |  | 150 | mg | 100 | mL | 1.5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg | 50 | 150 |  | n/a
322 | Adult | nafcillin | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
323 | Adult | nafcillin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
324 | Adult | nafcillin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
325 | Adult | nafcillin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 200 | 2000 |  | n/a
326 | Adult | nafcillin | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | mg | 200 | 2000 |  | n/a
327 | Adult | nafcillin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | mg | 200 | 2000 |  | n/a
328 | Adult | nafcillin | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
329 | Adult | nafcillin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
330 | Adult | nafcillin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
331 | Adult | nafcillin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.2 | 2 |  | n/a
332 | Adult | nafcillin | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | gram | 0.2 | 2 |  | n/a
333 | Adult | nafcillin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:30 | 01:15 | 01:00 | gram | 0.2 | 2 |  | n/a
334 | Adult | ondansetron |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:30 | mg | 2 | 32 |  | mg/mL | 0.02 | 0.04 | 0.64
335 | Adult | pamidronate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 02:00 | 20:00 | 04:00 | mg | 29 | 99 | 110 | mg/mL | 0.001 |  | 0.4 | Dedicated IV Line
336 | Adult | pantoprazole ivpb |  | 80 | mg | 100 | mL | 0.8 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:15 | mg | 40 | 80 |  | n/a
337 | Adult | penG 2.5 Mill Units |  | 2.5 | unit | 50 | mL | 0.05 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | unit | 2 | 3 |  | n/a |  |  |  | 1unit=1million
338 | Adult | penG 2.5 Mill Units |  | 2.5 | unit | 100 | mL | 0.025 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | unit | 2 | 3 |  | n/a |  |  |  | 1unit=1million
339 | Adult | penG 3 Million Units |  | 3 | unit | 50 | mL | 0.06 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | unit | 2.8 | 3.2 | 3.3 | n/a |  |  |  | 1unit=1million
340 | Adult | penG 5 Million Units |  | 5 | unit | 50 | mL | 0.1 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | unit | 4.5 | 5.5 |  | n/a |  |  |  | 1unit=1million
341 | Adult | penG 5 Million Units |  | 5 | unit | 100 | mL | 0.05 | unit/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | unit | 4.5 | 5.5 |  | n/a |  |  |  | 1unit=1million
342 | Adult | penG x Million Units |  | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | unit | 1 | 6 |  | unit/mL | 0.01 | 0.02 | 0.12 | 1unit=1million
343 | Adult | PHENobarbital |  | 130 | mg | 52 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:04 |  | 04:00 |  | mg | 5 | 270 |  | n/a
344 | Adult | PHENobarbital |  | 130 | mg | 102 | mL | 1.2745 | mg/mL | Pump | Primary and Secondary | 00:07 |  | 04:00 |  | mg | 5 | 270 |  | n/a
345 | Adult | PHENobarbital |  | 260 | mg | 54 | mL | 4.8148 | mg/mL | Pump | Primary and Secondary | 00:04 |  | 04:00 |  | mg | 5 | 270 |  | n/a
346 | Adult | PHENobarbital |  | 260 | mg | 104 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:07 |  | 04:00 |  | mg | 5 | 270 |  | n/a
347 | Adult | PHENobarbital |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:04 |  | 04:00 |  | mg | 5 | 270 |  | No
348 | Adult | phenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 | 00:30 | mg/kg | 0.5 | 20 |  | mg/mL | 0.5 | 1 | 10 | In-line filter
349 | Adult | phytonadione |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 01:00 | 00:30 | mg | 1 | 10 |  | mg/mL | 0.005 | 0.01 | 0.2
350 | Adult | piperacillin/tazo |  | 2.25 | gram | 50 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:30 | 00:30 | gram | 0.75 | 4.5 |  | n/a
351 | Adult | piperacillin/tazo |  | 2.25 | gram | 100 | mL | 0.0225 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:30 | 00:30 | gram | 0.75 | 4.5 |  | n/a
352 | Adult | piperacillin/tazo |  | 3.375 | gram | 50 | mL | 0.0675 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 | 04:00 | gram | 0.75 | 4.5 |  | n/a
353 | Adult | piperacillin/tazo |  | 3.375 | gram | 100 | mL | 0.0338 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 | 04:00 | gram | 0.75 | 4.5 |  | n/a
354 | Adult | piperacillin/tazo |  | 4.5 | gram | 50 | mL | 0.09 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 0.75 | 4.5 |  | n/a
355 | Adult | piperacillin/tazo |  | 4.5 | gram | 100 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:15 |  | gram | 0.75 | 4.5 |  | n/a
356 | Adult | potassium acetate |  | 40 | mEq | 520 | mL | 0.0769 | mEq/mL | Pump | Primary and Secondary | 02:00 | 03:30 | 12:00 | 04:00 | mEq | 1 | 46 | 48 | n/a
357 | Adult | potassium acetate |  | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:45 |  | 12:00 |  | mEq | 1 | 46 | 48 | mEq/mL |  | 0.01 | 0.46
358 | Adult | potassium chloride | Central Line | 10 | mEq | 100 | mL | 0.1 | mEq/mL | Pump | Primary and Secondary | 00:20 | 00:30 | 06:00 | 01:00 | mEq | 5 |  | 46 | n/a
359 | Adult | potassium chloride | Central Line | 20 | mEq | 100 | mL | 0.2 | mEq/mL | Pump | Primary and Secondary | 00:40 | 01:00 | 08:00 | 02:00 | mEq | 5 |  | 46 | n/a
360 | Adult | potassium chloride | Central Line | 40 | mEq | 520 | mL | 0.0769 | mEq/mL | Pump | Primary and Secondary | 01:40 | 02:00 | 12:00 | 04:00 | mEq | 5 |  | 46 | n/a
361 | Adult | potassium chloride | Central Line | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:50 | 01:00 | 12:00 |  | mEq | 5 |  | 46 | mEq/mL | 0.01 | 0.02 | 0.21
362 | Adult | potassium chloride | Peripheral Line | 10 | mEq | 100 | mL | 0.1 | mEq/mL | Pump | Primary and Secondary | 00:45 |  | 06:00 | 01:00 | mEq | 5 |  | 46 | n/a
363 | Adult | potassium chloride | Peripheral Line | 40 | mEq | 520 | mL | 0.0769 | mEq/mL | Pump | Primary and Secondary | 03:20 |  | 12:00 | 04:00 | mEq | 5 |  | 46 | n/a
364 | Adult | potassium chloride | Peripheral Line | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:50 | 01:00 | 12:00 |  | mEq | 5 |  | 46 | mEq/mL | 0.01 | 0.02 | 0.11
365 | Adult | potassium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 01:00 | 02:00 | 09:00 |  | mmol | 15 | 40 |  | mmol/mL | 0.005 | 0.01 | 0.24
366 | Adult | pralidoxime | Load <67kg WtBasedGM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:15 | gram/kg | 0.001 |  | 0.034 | gram/mL |  | 0.009 | 0.051
367 | Adult | pralidoxime | Load <67kg WtBasedMG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:15 | mg/kg | 1 |  | 34 | mg/mL |  | 9 | 51
368 | Adult | pralidoxime | Load >67kg WtBasedGM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:20 | gram | 1 | 21 |  | gram/mL |  | 0.009 | 0.051
369 | Adult | pralidoxime | Load >67kg WtBasedMG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:20 | mg | 500 |  | 2300 | mg/mL |  | 9 | 51
370 | Adult | pralidoxime | Load NonWtBased MG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:20 | mg | 900 |  | 2300 | mg/mL |  | 9 | 51
371 | Adult | pralidoxime | Load NonWtBasedGM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 |  | 01:00 | 00:15 | gram | 0.9 |  | 2.3 | gram/mL |  | 0.009 | 0.051
372 | Adult | pralidoxime | Maintenance Dose GM | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:10 |  | 24:00 |  | gram | 0.9 |  | 2.3 | gram/mL |  | 0.009 | 0.051
373 | Adult | pralidoxime | Maintenance Dose MG | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 |  | 24:00 |  | mg | 900 |  | 2300 | mg/mL | 5 | 9 | 51
374 | Adult | promethazine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | mg | 6.25 | 25 |  | mg/mL | 1 | 1.25 | 25
375 | Adult | raNITIdine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg | 25 | 51 |  | mg/mL | 0.1 | 0.5 | 2.5
376 | Adult | remdesivir |  | 100 | mg | 230 | mL | 0.4348 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
377 | Adult | remdesivir |  | 100 | mg | 250 | mL | 0.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
378 | Adult | remdesivir |  | 200 | mg | 210 | mL | 0.9524 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
379 | Adult | remdesivir |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:15 | 01:00 | mg | 99 | 201 |  | n/a
380 | Adult | rifAMPin |  | 300 | mg | 250 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
381 | Adult | rifAMPin |  | 600 | mg | 500 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
382 | Adult | rifAMPin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | mg/mL | 0.1 | 0.5 | 6
383 | Adult | riTUXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:15 | 02:45 | 25:00 |  | mg | 250 | 1100 |  | mg/mL | 0.1 | 0.2 | 4 | Premedicate
384 | Adult | SMX/TMP |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:00 | 02:00 | 01:30 | mg/kg | 1.5 | 7 |  | mg/mL | 0.05 | 0.1 | 1.2 | 48 Hour Tubing
385 | Adult | sodium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 00:45 |  | 12:00 |  | mmol | 15 | 60 |  | mmol/mL | 0.005 | 0.01 | 0.25
386 | Adult | sotrovimab |  | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:22 | 00:23 | 04:00 | 00:30 | mg | 490 | 510 | 511 | n/a |  |  |  | Low Binding Filter
387 | Adult | sotrovimab |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:22 | 00:23 | 04:00 | 00:30 | mg | 490 | 510 | 511 | n/a |  |  |  | Low Binding Filter
388 | Adult | tedizolid phosphate |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 01:15 | 01:00 | mg | 175 | 225 | 250 | n/a
389 | Adult | thiamine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:15 | 00:30 | mg | 90 | 550 |  | mg/mL | 0.001 | 0.7 | 5
390 | Adult | ticarcillin/clav |  | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
391 | Adult | ticarcillin/clav |  | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
392 | Adult | ticarcillin/clav |  | 3.1 | gram | 50 | mL | 0.062 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
393 | Adult | ticarcillin/clav |  | 3.1 | gram | 100 | mL | 0.031 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 2 | 3.1 |  | n/a
394 | Adult | tigecycline |  | 50 | mg | 50 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:15 | 00:30 | mg | 25 | 100 |  | n/a
395 | Adult | tigecycline |  | 50 | mg | 100 | mL | 0.5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:15 | 00:30 | mg | 25 | 100 |  | n/a
396 | Adult | tigecycline |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:15 | 00:30 | mg | 25 | 100 |  | n/a
397 | Adult | tobramycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 |  | 01:00 | 00:30 | mg/kg | 0.5 | 9 |  | mg/mL | 0.1 | 0.5 | 10
398 | Adult | tocilizumab |  | 800 | mg | 100 | mL | 8 | mg/mL | Pump | Primary Only | 00:54 | 00:55 | 04:00 | 01:00 | mg | 200 |  | 801 | n/a
399 | Adult | tocilizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 00:54 | 00:55 | 04:00 | 01:00 | mg | 200 |  | 801 | mg/mL | 0.9 | 1 | 8.1
400 | Adult | tocilizumab |  | 800 | mg | 100 | mL | 8 | mg/mL | Pump | Primary Only | 00:54 | 00:55 | 04:00 | 01:00 | mg/kg | 3.9 | 12.1 | 12.2 | n/a
401 | Adult | tocilizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 00:54 | 00:55 | 04:00 | 01:00 | mg/kg | 3.9 | 12.1 | 12.2 | mg/mL | 0.9 | 1 | 8.1
402 | Adult | tranexamic BOLUS | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:09 |  | 00:40 | 00:15 | mg | 500 | 2000 |  | n/a
403 | Adult | tranexamic BOLUS | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:07 | 00:20 | 00:40 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
404 | Adult | tranexamic BOLUS | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary | 00:10 |  | 00:40 | 00:15 | gram | 0.5 | 2 |  | n/a
405 | Adult | tranexamic BOLUS | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:07 | 00:20 | 00:40 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
406 | Adult | valproate sodium |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg | 250 | 2000 |  | mg/mL | 0.1 | 1 | 11
407 | Adult | valproate sodium |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg/kg | 1 | 40 |  | mg/mL | 0.1 | 1 | 11
408 | Adult | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg | 500 | 2000 |  | n/a
409 | Adult | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
410 | Adult | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
411 | Adult | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
412 | Adult | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg | 500 | 2000 |  | n/a
413 | Adult | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | mg | 500 | 2000 |  | n/a
414 | Adult | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg | 500 | 2000 |  | n/a
415 | Adult | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg | 500 | 2000 |  | n/a
416 | Adult | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
417 | Adult | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
418 | Adult | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
419 | Adult | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
420 | Adult | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 |  | mg | 500 | 2000 |  | mg/mL | 1 | 3 | 5
421 | Adult | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 04:00 | 00:30 | mg/kg | 5 |  | 30 | n/a
422 | Adult | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
423 | Adult | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
424 | Adult | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
425 | Adult | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
426 | Adult | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
427 | Adult | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
428 | Adult | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 05:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
429 | Adult | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
430 | Adult | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
431 | Adult | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
432 | Adult | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
433 | Adult | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 |  | 06:00 |  | mg/kg | 5 |  | 30 | mg/mL | 1 | 3 | 5
434 | Adult | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:20 |  | 04:00 | 00:30 | gram | 0.5 | 2 |  | n/a
435 | Adult | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
436 | Adult | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
437 | Adult | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
438 | Adult | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 04:00 | 01:00 | gram | 0.5 | 2 |  | n/a
439 | Adult | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 04:00 | 01:30 | gram | 0.5 | 2 |  | n/a
440 | Adult | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:15 |  | 05:00 | 01:30 | gram | 0.5 | 2 |  | n/a
441 | Adult | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:15 |  | 05:00 | 01:30 | gram | 0.5 | 2 |  | n/a
442 | Adult | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:35 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
443 | Adult | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:35 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
444 | Adult | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:50 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
445 | Adult | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:50 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
446 | Adult | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:50 | 04:00 |  | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.003 | 0.006
447 | Adult | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:20 |  | 06:00 | 00:30 | gram/kg | 0.005 |  | 0.03 | n/a
448 | Adult | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 06:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
449 | Adult | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 06:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
450 | Adult | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
451 | Adult | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
452 | Adult | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 06:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
453 | Adult | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 06:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
454 | Adult | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 06:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
455 | Adult | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:35 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
456 | Adult | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:35 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
457 | Adult | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
458 | Adult | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
459 | Adult | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:25 |  | 06:00 |  | gram/kg | 0.005 |  | 0.03 | gram/mL | 0.001 | 0.003 | 0.006
460 | Adult | voriconazole |  | 200 | mg | 50 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 02:00 | 01:00 | mg | 200 | 400 |  | n/a
461 | Adult | voriconazole |  | 400 | mg | 100 | mL | 4 | mg/mL | Pump | Primary and Secondary |  | 01:40 | 04:00 | 02:00 | mg | 200 | 400 |  | n/a
462 | Adult | voriconazole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 | 01:00 | 04:00 | 02:00 | mg/kg | 2 | 7 |  | mg/mL |  | 0.5 | 5
463 | Adult | zidovudine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 24:00 |  | mg/kg | 1 | 2.2 | 3 | mg/mL | 0.001 | 0.01 | 4.2
464 | Adult | zoledronic acid |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg | 2 | 5 |  | mg/mL | 0.01 | 0.03 | 0.05
465 | Adult | zz mag sulfate |  | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary | 00:07 | 00:14 | 02:00 | 00:20 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
466 | Adult | zz mag sulfate |  | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary | 00:07 | 00:14 | 02:00 | 00:20 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
467 | Adult | zz mag sulfate |  | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary | 00:14 | 00:28 | 04:00 | 00:40 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
468 | Adult | zz mag sulfate |  | 4 | gram | 100 | mL | 0.04 | gram/mL | Pump | Primary and Secondary | 00:28 | 00:56 | 08:00 | 01:20 | gram | 1 | 6 |  | n/a |  |  |  | Code only drugs
469 | Adult | zz mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:14 | 00:28 | 08:00 |  | gram | 1 | 6 |  | gram/mL | 0.005 | 0.01 | 0.08 | Code only drugs
470 | Nursery | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:50 | 00:55 | 02:00 | 01:00 | mg/kg | 4 | 21 | 22 | mg/mL | 1 | 4 | 7
471 | Nursery | ampicillin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:07 | 00:08 | 01:10 | 00:30 | mg/kg | 24 | 101 | 105 | mg/mL |  | 10 | 105
472 | Nursery | ceFAZolin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:07 | 00:08 | 01:05 | 00:30 | mg/kg | 8 | 55 | 60 | mg/mL |  | 45 | 55
473 | Nursery | cefotaxime |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:12 | 00:13 | 01:05 | 00:30 | mg/kg | 12.5 | 80 | 81 | mg/mL |  | 45 | 55
474 | Nursery | cefTRIAXone |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:45 | 01:10 | 01:00 | mg/kg | 20 | 135 | 140 | mg/mL |  | 35 | 45 | CeftriSyringe
475 | Nursery | clindamycin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:08 | 00:09 | 01:05 | 00:30 | mg/kg | 3 | 7.1 | 7.2 | mg/mL |  | 13 | 17
476 | Nursery | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:25 | 02:15 | 00:30 | mg/kg | 0.5 | 7 | 7.1 | mg/mL |  | 3 | 7
477 | Nursery | nafcillin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:25 | 01:10 | 00:30 | mg/kg | 10 | 53 | 55 | mg/mL |  | 35 | 45
478 | Nursery | piperacillin/tazo |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:22 | 00:25 | 04:15 | 00:30 | mg/kg | 49 | 125 | 130 | mg/mL |  | 15 | 85
479 | Nursery | vancomycin |  | - - | mg | - - | mL |  | mg/mL | Syringe | Primary Only | 00:52 | 00:55 | 03:10 | 01:00 | mg/kg | 4 | 22 | 23 | mg/mL |  | 4.5 | 6.5
480 | ONC Infusion Center | abatacept |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
481 | ONC Infusion Center | abatacept |  | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
482 | ONC Infusion Center | abatacept |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
483 | ONC Infusion Center | aprepitant |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:30 | mg | 99 | 131 |  | n/a |  |  |  | DEHP
484 | ONC Infusion Center | aprepitant |  | 130 | mg | 100 | mL | 1.3 | mg/mL | Pump | Primary and Secondary | 00:11 |  | 01:00 | 00:30 | mg | 99 | 131 |  | n/a |  |  |  | DEHP
485 | ONC Infusion Center | atezolizumab |  | 840 | mg | 250 | mL | 3.36 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 799 | 1681 |  | n/a
486 | ONC Infusion Center | atezolizumab |  | 1200 | mg | 250 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 799 | 1681 |  | n/a
487 | ONC Infusion Center | atezolizumab |  | 1680 | mg | 250 | mL | 6.72 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 799 | 1681 |  | n/a
488 | ONC Infusion Center | Bendamusti (Bendeka) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:08 | 00:15 | 00:10 | mg/m2 | 25 | 120 |  | mg/mL |  | 1.85 | 5.6
489 | ONC Infusion Center | Bendamusti (Treanda) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:20 | 00:30 | mg/m2 | 25 | 120 |  | mg/mL |  | 0.2 | 0.7
490 | ONC Infusion Center | bevacizumab | First Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:15 | 01:30 | 03:00 | 01:30 | mg/kg | 5 | 15 |  | mg/mL |  | 1 | 25
491 | ONC Infusion Center | bevacizumab | Second Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 01:30 | 01:00 | mg/kg | 5 | 15 |  | mg/mL |  | 1 | 25
492 | ONC Infusion Center | bevacizumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 01:30 | 00:30 | mg/kg | 5 | 15 |  | mg/mL |  | 1 | 25
493 | ONC Infusion Center | bleomycin | BSA Based Dose | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:05 | 00:15 | 00:10 | unit/m2 | 2 | 30 |  | unit/mL |  | 0.1 | 3
494 | ONC Infusion Center | bleomycin | Non-Based Based Dose | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:08 | 00:20 | 00:10 | unit | 1 | 60 | 61 | unit/mL |  | 0.1 | 3
495 | ONC Infusion Center | brentuximab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 00:40 | 00:30 | mg/kg | 1 | 1.8 | 1.81 | mg/mL |  | 0.4 | 1.8 | MAX DOSE 180 mg
496 | ONC Infusion Center | cabazitaxel |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:10 | 01:00 | mg/m2 | 15 | 25 |  | mg/mL |  | 0.1 | 0.26 | 0.22 Filter Needed
497 | ONC Infusion Center | CARBOplatin | Based on AUC | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 24:00 | 00:30 | mg | 100 | 800 |  | mg/mL |  | 0.4 | 5
498 | ONC Infusion Center | CARBOplatin | Based on BSA | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 | 00:30 | mg/m2 | 60 | 700 |  | mg/mL |  | 0.5 | 5
499 | ONC Infusion Center | carfilzomib |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:35 | 00:30 | mg/m2 | 15 | 70 |  | mg/mL |  | 0.3 | 1.6
500 | ONC Infusion Center | cemiplimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 00:35 | 00:30 | mg | 349 | 351 |  | mg/mL |  | 1 | 20 | In-line filter
501 | ONC Infusion Center | cetuximab | Loading Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:45 |  | 04:00 | 02:00 | mg/m2 | 400 | 500 |  | No |  |  |  | Low Binding Filter
502 | ONC Infusion Center | cetuximab | Maintenance Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 01:30 | 01:00 | mg/m2 | 150 | 250 |  | No |  |  |  | Low Binding Filter
503 | ONC Infusion Center | CISplatin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 24:00 |  | mg/m2 | 10 | 100 |  | mg/mL |  | 0.04 | 0.5
504 | ONC Infusion Center | cyclophosphamide | BSA Based Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:30 | 01:00 | mg/m2 | 100 | 2000 |  | mg/mL | 1.9 | 2 | 20
505 | ONC Infusion Center | cyclophosphamide | Weight Based | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 01:30 |  | mg/kg | 1 | 50 | 51 | mg/mL | 1.9 | 2 | 20
506 | ONC Infusion Center | dacarbazine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:40 |  | mg/m2 | 175 | 850 |  | mg/mL |  | 0.19 | 20.1
507 | ONC Infusion Center | daratumumab | First Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 02:20 | 02:30 | 20:00 |  | mg/kg | 7 | 16 | 16.1 | mg/mL |  | 0.4 | 4 | 0.22 Filter Needed
508 | ONC Infusion Center | daratumumab | Second Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 02:30 | 04:00 | 10:00 |  | mg/kg | 15 | 16 | 16.1 | mg/mL |  | 0.9 | 5 | 0.22 Filter Needed
509 | ONC Infusion Center | daratumumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 02:30 | 03:50 | 10:00 |  | mg/kg | 15 | 16 | 16.1 | mg/mL |  | 0.9 | 5 | 0.22 Filter Needed
510 | ONC Infusion Center | dexamethasone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:30 | 00:15 | mg | 1 | 21 |  | No
511 | ONC Infusion Center | DOCEtaxel |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg/m2 | 16 | 100 |  | mg/mL |  | 0.3 | 0.74 | DEHP
512 | ONC Infusion Center | DOXOrubicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 24:00 |  | mg/m2 | 3 | 90 |  | mg/mL |  | 0.01 | 5
513 | ONC Infusion Center | DOXOrubicin, liposo |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:45 | 03:00 | 01:00 | mg/m2 | 2.5 | 55 |  | mg/mL |  | 0.1 | 0.36 | NO FILTER
514 | ONC Infusion Center | eriBULin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:02 | 00:20 | 00:10 | mg/m2 | 0.7 | 1.4 |  | mg/mL |  | 0.001 | 0.5 | No Dextrose
515 | ONC Infusion Center | etoposide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 03:00 |  | mg/m2 | 35 | 750 |  | mg/mL |  | 0.2 | 0.4 | Non-PVC
516 | ONC Infusion Center | fluorouracil |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 24:00 |  | mg/m2 | 200 | 3000 |  | mg/mL |  | 0.1 | 50
517 | ONC Infusion Center | fosaprepitant |  | 150 | mg | 150 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 00:40 | 00:30 | mg | 149 | 151 |  | n/a
518 | ONC Infusion Center | gemcitabine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg/m2 | 300 | 1250 |  | mg/mL |  | 0.1 | 38
519 | ONC Infusion Center | golimumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:00 | 00:30 | mg/kg | 1.8 |  | 2.2 | mg/mL | 0.49 | 0.5 | 4 | Low Binding Filter
520 | ONC Infusion Center | ifosfamide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 25:00 |  | mg/m2 | 250 | 5000 |  | mg/mL |  | 0.6 | 80
521 | ONC Infusion Center | ipilimumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:25 | 01:35 | 01:30 | mg/kg | 1 | 10 |  | mg/mL |  | 1 | 2
522 | ONC Infusion Center | irinotecan |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 01:20 |  | 02:30 | 01:30 | mg/m2 | 20 | 600 |  | mg/mL |  | 0.12 | 2.8
523 | ONC Infusion Center | irinotecan liposomal |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:20 | 02:00 | 01:30 | mg/m2 | 34 | 71 |  | mg/mL | 0.05 | 0.1 | 0.6 | NO FILTER
524 | ONC Infusion Center | leucovorin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mg/m2 | 10 | 1000 |  | mg/mL |  | 0.05 | 20
525 | ONC Infusion Center | leucovorin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 03:00 |  | mg | 1 | 100 |  | mg/mL |  | 0.05 | 20
526 | ONC Infusion Center | mag sulfate |  | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 02:30 | 01:00 | gram | 0.9 | 1.1 |  | n/a
527 | ONC Infusion Center | mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:45 | 08:00 |  | gram | 0.9 | 1.1 |  | No
528 | ONC Infusion Center | mesna |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 25:00 |  | mg/m2 | 300 | 5000 |  | mg/mL |  | 0.01 | 20
529 | ONC Infusion Center | mesna |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 25:00 |  | mg | 1 | 1000 |  | mg/mL |  | 0.01 | 20
530 | ONC Infusion Center | mitoMYcin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg/m2 | 1 | 20 |  | mg/mL |  | 0.02 | 0.04
531 | ONC Infusion Center | mitoMYcin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 02:00 |  | mg | 1 | 20 |  | mg/mL |  | 0.02 | 0.04
532 | ONC Infusion Center | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
533 | ONC Infusion Center | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
534 | ONC Infusion Center | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
535 | ONC Infusion Center | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
536 | ONC Infusion Center | nivolumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | mg/mL |  | 1 | 10 | In-line filter
537 | ONC Infusion Center | obinutuzumab | First Dose | 100 | mg | 104 | mL | 0.9615 | mg/mL | Pump | Primary and Secondary |  | 04:00 | 08:00 | 04:00 | mg | 99 | 100.1 |  | n/a
538 | ONC Infusion Center | obinutuzumab | Subsequent Doses | 900 | mg | 286 | mL | 3.1469 | mg/mL | Pump | Primary and Secondary |  | 02:00 | 60:00 |  | mg | 900 | 1000 |  | n/a
539 | ONC Infusion Center | obinutuzumab | Subsequent Doses | 1000 | mg | 290 | mL | 3.4483 | mg/mL | Pump | Primary and Secondary |  | 02:00 | 60:00 |  | mg | 900 | 1000 |  | n/a
540 | ONC Infusion Center | obinutuzumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 02:00 | 60:00 |  | mg | 900 | 1000 |  | mg/mL |  | 0.9 | 3.5
541 | ONC Infusion Center | ondansetron |  | 8 | mg | 54 | mL | 0.1481 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 4 | 16 |  | n/a
542 | ONC Infusion Center | ondansetron |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 |  | mg | 4 | 16 |  | mg/mL |  | 0.01 | 0.64
543 | ONC Infusion Center | oxaliplatin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:00 | 06:30 | 02:00 | mg/m2 | 20 | 130 |  | mg/mL |  | 0.2 | 1.2
544 | ONC Infusion Center | PACLitaxel |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 24:00 |  | mg/m2 | 40 | 225 |  | mg/mL |  | 0.3 | 1.2 | Filter and DEHP Free
545 | ONC Infusion Center | PACLitaxel protein-b |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/m2 | 49 | 261 |  | mg/mL |  | 4.9 | 5.1 | NO FILTER
546 | ONC Infusion Center | panitumumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:00 | 01:00 | mg/kg | 3 | 6 |  | mg/mL |  | 2 | 10 | Low Binding Filter
547 | ONC Infusion Center | pembrolizumab |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 00:35 | 00:30 | mg | 100 | 250 |  | n/a |  |  |  | In-line filter
548 | ONC Infusion Center | pembrolizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 00:35 | 00:30 | mg | 100 | 250 |  | mg/mL |  | 1 | 10 | In-line filter
549 | ONC Infusion Center | PEMEtrexed |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:10 | mg/m2 | 250 | 900 |  | mg/mL |  | 3 | 13
550 | ONC Infusion Center | pertuzumab | Loading Dose | 840 | mg | 278 | mL | 3.0216 | mg/mL | Pump | Primary and Secondary | 01:00 |  | 02:00 | 01:00 | mg | 839 |  | 841 | n/a
551 | ONC Infusion Center | pertuzumab | Maintenance Dose | 420 | mg | 264 | mL | 1.5909 | mg/mL | Pump | Primary and Secondary | 00:30 |  | 02:00 | 00:30 | mg | 419 |  | 421 | n/a
552 | ONC Infusion Center | ramucirumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 |  | mg/kg | 5 | 10 |  | mg/mL |  | 0.8 | 12 | Low Binding Filter
553 | ONC Infusion Center | raNITIdine |  | 50 | mg | 102 | mL | 0.4902 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:20 | mg | 25 | 50 |  | n/a
554 | ONC Infusion Center | riTUXimab | Based on BSA | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg/m2 | 250 | 500 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
555 | ONC Infusion Center | riTUXimab | Rheumatoid Arthritis | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg | 90 | 1200 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
556 | ONC Infusion Center | temozolomide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:20 | 02:00 | 01:30 | mg/m2 | 74 | 201 |  | mg/mL |  | 2.4 | 2.6 | Flush Before/After
557 | ONC Infusion Center | topotecan |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg/m2 | 0.2 | 1.5 | 4 | mg/mL |  | 0.01 | 0.5
558 | ONC Infusion Center | trastuzumab | Loading Dose | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:00 | 03:00 | 01:30 | mg/kg | 4 | 8 |  | mg/mL |  | 0.4 | 8
559 | ONC Infusion Center | trastuzumab | Subsequent Doses | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 2 | 6 |  | mg/mL |  | 0.4 | 8
560 | ONC Infusion Center | vinBLAStine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 01:00 | 00:15 | mg/m2 | 0.75 | 18.5 | 18.6 | mg/mL |  | 0.001 | 0.6 | Vesicant
561 | ONC Infusion Center | vinCRIStine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:20 | 00:10 | mg | 0.4 | 2 | 2.1 | mg/mL |  | 0.001 | 0.1 | Vesicant
562 | ONC Infusion Center | vinorelbine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:06 | 00:30 | 00:10 | mg/m2 | 5 | 35 |  | mg/mL |  | 0.5 | 2
563 | ONC Infusion Center | ziv-aflibercept |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:55 | 01:05 | 01:00 | mg/kg | 2 | 4 |  | mg/mL |  | 0.6 | 8 | 0.22 Filter Needed
564 | ONC Infusion Center | zoledronic acid |  | 4 | mg | 100 | mL | 0.04 | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 02:00 | 00:30 | mg | 3 | 4 |  | n/a
565 | OutPt InfusionCenter | abatacept |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
566 | OutPt InfusionCenter | abatacept |  | 750 | mg | 100 | mL | 7.5 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
567 | OutPt InfusionCenter | abatacept |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:10 |  | mg | 499 |  | 1001 | n/a |  |  |  | Low Binding 0.2or1.2
568 | OutPt InfusionCenter | albumin 25% |  | 12.5 | gram | 50 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:16 | 01:30 |  | gram | 12.5 | 25 |  | n/a
569 | OutPt InfusionCenter | albumin 25% |  | 25 | gram | 100 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:32 | 03:00 |  | gram | 12.5 | 25 |  | n/a
570 | OutPt InfusionCenter | ampho B (liposomal) |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 00:55 | 01:00 | 06:30 | 02:00 | mg/kg | 2 | 7 |  | mg/mL | 0.1 | 0.2 | 2 | AmBisome
571 | OutPt InfusionCenter | bezlotoxumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 | 00:51 | 03:00 | 01:00 | mg/kg | 9 | 11 | 11.1 | mg/mL | 0.8 | 0.9 | 10.1 | 0.2 to 5 micron filt
572 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
573 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
574 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
575 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | n/a
576 | OutPt InfusionCenter | ceFAZolin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 480 | 2300 |  | mg/mL | 5 | 10 | 40
577 | OutPt InfusionCenter | ceFAZolin | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
578 | OutPt InfusionCenter | ceFAZolin | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
579 | OutPt InfusionCenter | ceFAZolin | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
580 | OutPt InfusionCenter | ceFAZolin | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | n/a
581 | OutPt InfusionCenter | ceFAZolin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.48 | 2.3 |  | gram/mL | 0.005 | 0.01 | 0.04
582 | OutPt InfusionCenter | cefepime | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
583 | OutPt InfusionCenter | cefepime | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
584 | OutPt InfusionCenter | cefepime | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
585 | OutPt InfusionCenter | cefepime | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
586 | OutPt InfusionCenter | cefepime | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
587 | OutPt InfusionCenter | cefepime | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | n/a
588 | OutPt InfusionCenter | cefepime | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 4 | 4.9 | 40
589 | OutPt InfusionCenter | cefepime | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
590 | OutPt InfusionCenter | cefepime | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
591 | OutPt InfusionCenter | cefepime | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
592 | OutPt InfusionCenter | cefepime | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | n/a
593 | OutPt InfusionCenter | cefepime | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.004 | 0.009 | 0.04
594 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
595 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
596 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 2000 | mg | 50 | mL | 40 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
597 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | 2000 | mg | 100 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 500 | 2000 |  | n/a
598 | OutPt InfusionCenter | cefTRIAXone | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
599 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.5 | 2 |  | n/a
600 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.5 | 2 |  | n/a
601 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 2 | gram | 50 | mL | 0.04 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.5 | 2 |  | n/a
602 | OutPt InfusionCenter | cefTRIAXone | GM dosing | 2 | gram | 100 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.5 | 2 |  | n/a
603 | OutPt InfusionCenter | cefTRIAXone | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:15 | 01:00 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
604 | OutPt InfusionCenter | ciprofloxacin |  | 200 | mg | 100 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
605 | OutPt InfusionCenter | ciprofloxacin |  | 400 | mg | 200 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 03:05 |  | mg | 180 | 440 | 500 | n/a
606 | OutPt InfusionCenter | clindamycin |  | 300 | mg | 50 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
607 | OutPt InfusionCenter | clindamycin |  | 300 | mg | 100 | mL | 3 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
608 | OutPt InfusionCenter | clindamycin |  | 600 | mg | 50 | mL | 12 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
609 | OutPt InfusionCenter | clindamycin |  | 600 | mg | 100 | mL | 6 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
610 | OutPt InfusionCenter | clindamycin |  | 900 | mg | 50 | mL | 18 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
611 | OutPt InfusionCenter | clindamycin |  | 900 | mg | 100 | mL | 9 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 50 | 901 |  | n/a
612 | OutPt InfusionCenter | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg | 100 | 2000 |  | mg/mL | 1 | 5 | 20
613 | OutPt InfusionCenter | DAPTOmycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg/kg | 4 | 10.5 |  | mg/mL | 1 | 5 | 20
614 | OutPt InfusionCenter | dexamethasone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:05 | 00:30 | 00:15 | mg | 1 | 21 |  | No
615 | OutPt InfusionCenter | doxycycline |  | 100 | mg | 100 | mL | 1 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
616 | OutPt InfusionCenter | doxycycline |  | 200 | mg | 250 | mL | 0.8 | mg/mL | Pump | Primary and Secondary | 00:49 | 00:50 | 04:00 | 01:00 | mg | 25 | 200 | 300 | n/a |  |  |  | Protect From Light
617 | OutPt InfusionCenter | ertapenem | .MG Dosing | 500 | mg | 50 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
618 | OutPt InfusionCenter | ertapenem | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
619 | OutPt InfusionCenter | ertapenem | .MG Dosing | 1000 | mg | 50 | mL | 20 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
620 | OutPt InfusionCenter | ertapenem | .MG Dosing | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | mg | 450 | 1250 |  | n/a
621 | OutPt InfusionCenter | ertapenem | GM dosing | 0.5 | gram | 50 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
622 | OutPt InfusionCenter | ertapenem | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
623 | OutPt InfusionCenter | ertapenem | GM dosing | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
624 | OutPt InfusionCenter | ertapenem | GM dosing | 1 | gram | 100 | mL | 0.01 | gram/mL | Pump | Primary and Secondary |  | 00:25 | 02:00 | 00:30 | gram | 0.45 | 1.25 |  | n/a
625 | OutPt InfusionCenter | erythromycin |  | 250 | mg | 100 | mL | 2.5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
626 | OutPt InfusionCenter | erythromycin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | n/a
627 | OutPt InfusionCenter | erythromycin |  | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:30 | 01:00 | mg | 100 | 1000 |  | n/a
628 | OutPt InfusionCenter | erythromycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:29 | 00:30 | 02:00 | 01:00 | mg | 100 | 1000 |  | mg/mL | 0.05 | 1 | 5
629 | OutPt InfusionCenter | famotidine |  | 20 | mg | 50 | mL | 0.4 | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg | 8 | 44 | 50 | n/a
630 | OutPt InfusionCenter | famotidine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 |  | 02:00 |  | mg | 8 | 44 | 50 | mg/mL | 0.09 | 0.1 | 3.5
631 | OutPt InfusionCenter | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 |  | mg | 10 | 1000 |  | mg/mL | 0.05 | 0.1 | 10
632 | OutPt InfusionCenter | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 01:20 |  | mg/kg | 0.5 | 10 |  | mg/mL | 0.05 | 0.1 | 10
633 | OutPt InfusionCenter | golimumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:27 | 00:28 | 01:00 | 00:30 | mg/kg | 1.8 |  | 2.2 | mg/mL | 0.49 | 0.5 | 4 | Low Binding Filter
634 | OutPt InfusionCenter | immune glob (IVIG) |  | 5 | gram | 50 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:00 |  | gram | 1 | 300 |  | n/a |  |  |  | Hypersensitivity
635 | OutPt InfusionCenter | immune glob (IVIG) |  | 5 | gram | 100 | mL | 0.05 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 04:00 |  | gram | 1 | 300 |  | n/a |  |  |  | Hypersensitivity
636 | OutPt InfusionCenter | immune glob (IVIG) |  | 10 | gram | 100 | mL | 0.1 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 08:00 |  | gram | 1 | 300 |  | n/a |  |  |  | Hypersensitivity
637 | OutPt InfusionCenter | immune glob (IVIG) |  | 10 | gram | 200 | mL | 0.05 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 08:00 |  | gram | 1 | 300 |  | n/a |  |  |  | Hypersensitivity
638 | OutPt InfusionCenter | immune glob (IVIG) |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 02:00 | 25:00 |  | gram | 1 | 300 |  | gram/mL | 0.01 | 0.05 | 0.1 | Hypersensitivity
639 | OutPt InfusionCenter | immune glob (IVIG) |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 02:00 | 36:00 |  | gram/kg | 1 | 2.4 |  | gram/mL | 0.01 | 0.05 | 0.1 | Hypersensitivity
640 | OutPt InfusionCenter | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg | 50 | 2000 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
641 | OutPt InfusionCenter | inFLIXimab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 25:00 |  | mg/kg | 2 | 10 |  | mg/mL | 0.1 | 0.4 | 4 | Hypersensitivity
642 | OutPt InfusionCenter | levoFLOXacin |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
643 | OutPt InfusionCenter | levoFLOXacin |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 | 01:00 | mg | 250 | 750 |  | n/a
644 | OutPt InfusionCenter | levoFLOXacin |  | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 01:15 | 02:00 | 01:30 | mg | 250 | 750 |  | n/a
645 | OutPt InfusionCenter | mag sulfate |  | 1 | gram | 50 | mL | 0.02 | gram/mL | Pump | Primary and Secondary |  | 00:45 | 02:30 | 01:00 | gram | 0.9 | 1.1 |  | n/a
646 | OutPt InfusionCenter | mag sulfate |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:45 | 08:00 |  | gram | 0.9 | 1.1 |  | No
647 | OutPt InfusionCenter | methylPREDNISolone |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 125 | 2000 |  | n/a
648 | OutPt InfusionCenter | methylPREDNISolone |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:45 | 02:00 | 01:00 | mg | 125 | 2000 |  | n/a
649 | OutPt InfusionCenter | methylPREDNISolone |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 |  | mg | 125 | 2000 |  | mg/mL | 0.5 | 1 | 15
650 | OutPt InfusionCenter | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
651 | OutPt InfusionCenter | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 235 | 485 |  | n/a |  |  |  | In-line filter
652 | OutPt InfusionCenter | nivolumab |  | 240 | mg | 100 | mL | 2.4 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
653 | OutPt InfusionCenter | nivolumab |  | 480 | mg | 100 | mL | 4.8 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | n/a |  |  |  | In-line filter
654 | OutPt InfusionCenter | nivolumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 1 | 4 |  | mg/mL |  | 1 | 10 | In-line filter
655 | OutPt InfusionCenter | ocrelizumab |  | 300 | mg | 250 | mL | 1.2 | mg/mL | Pump | Primary and Secondary | 02:20 | 02:25 | 24:00 |  | mg | 290 | 610 | 611 | n/a |  |  |  | 0.22 Filter Needed
656 | OutPt InfusionCenter | ocrelizumab |  | 600 | mg | 500 | mL | 1.2 | mg/mL | Pump | Primary and Secondary | 01:50 | 01:55 | 24:00 |  | mg | 290 | 610 | 611 | n/a |  |  |  | 0.22 Filter Needed
657 | OutPt InfusionCenter | ondansetron |  | 8 | mg | 54 | mL | 0.1481 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 4 | 16 |  | n/a
658 | OutPt InfusionCenter | ondansetron |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 01:00 |  | mg | 4 | 16 |  | mg/mL | 0.02 | 0.04 | 0.64
659 | OutPt InfusionCenter | promethazine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 |  | mg | 6.25 | 25 |  | mg/mL | 1 | 1.25 | 25
660 | OutPt InfusionCenter | raNITIdine |  | 50 | mg | 102 | mL | 0.4902 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:20 | mg | 25 | 50 |  | n/a
661 | OutPt InfusionCenter | rifAMPin |  | 300 | mg | 250 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
662 | OutPt InfusionCenter | rifAMPin |  | 600 | mg | 500 | mL | 1.2 | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | n/a
663 | OutPt InfusionCenter | rifAMPin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 03:00 | 01:00 | mg | 100 | 600 |  | mg/mL | 0.1 | 0.5 | 6
664 | OutPt InfusionCenter | riTUXimab | Based on BSA | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg/m2 | 250 | 500 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
665 | OutPt InfusionCenter | riTUXimab | Rheumatoid Arthritis | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 01:30 | 06:00 |  | mg | 90 | 1200 |  | mg/mL |  | 0.9 | 4.1 | Hypersensitivity
666 | OutPt InfusionCenter | teprotumumab | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 | 00:55 | 08:00 |  | mg/kg | 7 | 21 | 22 | mg/mL | 2.9 | 3 | 18.1
667 | OutPt InfusionCenter | teprotumumab | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 | 00:55 | 08:00 |  | gram/kg | 0.007 | 0.021 | 0.022 | gram/mL | 0.002 | 0.003 | 0.019
668 | OutPt InfusionCenter | tocilizumab |  | 800 | mg | 100 | mL | 8 | mg/mL | Pump | Primary Only | 00:54 | 00:55 | 04:00 | 01:00 | mg | 200 |  | 801 | n/a
669 | OutPt InfusionCenter | tocilizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 00:54 | 00:55 | 04:00 | 01:00 | mg | 200 |  | 801 | mg/mL | 0.9 | 1 | 8.1
670 | OutPt InfusionCenter | tocilizumab |  | 800 | mg | 100 | mL | 8 | mg/mL | Pump | Primary Only | 00:54 | 00:55 | 04:00 | 01:00 | mg/kg | 3.9 | 12.1 | 12.2 | n/a
671 | OutPt InfusionCenter | tocilizumab |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary Only | 00:54 | 00:55 | 04:00 | 01:00 | mg/kg | 3.9 | 12.1 | 12.2 | mg/mL | 0.9 | 1 | 8.1
672 | OutPt InfusionCenter | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 06:00 | 00:30 | mg | 500 | 2000 |  | n/a
673 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 06:00 | 01:00 | mg | 500 | 2000 |  | n/a
674 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 06:00 | 01:00 | mg | 500 | 2000 |  | n/a
675 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 | 01:00 | mg | 500 | 2000 |  | n/a
676 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 | 01:00 | mg | 500 | 2000 |  | n/a
677 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 06:00 | 01:30 | mg | 500 | 2000 |  | n/a
678 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 06:00 | 01:30 | mg | 500 | 2000 |  | n/a
679 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 06:00 | 01:30 | mg | 500 | 2000 |  | n/a
680 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
681 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
682 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
683 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg | 500 | 2000 |  | n/a
684 | OutPt InfusionCenter | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 |  | mg | 500 | 2000 |  | mg/mL | 1 | 3 | 5
685 | OutPt InfusionCenter | vancomycin | .MG Dosing | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 |  | 06:00 | 00:30 | mg/kg | 5 |  | 30 | n/a
686 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 150 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 06:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
687 | OutPt InfusionCenter | vancomycin | .MG Dosing | 750 | mg | 250 | mL | 3 | mg/mL | Pump | Primary and Secondary | 00:40 |  | 06:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
688 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 200 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
689 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1000 | mg | 250 | mL | 4 | mg/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 | 01:00 | mg/kg | 5 |  | 30 | n/a
690 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1250 | mg | 250 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:10 |  | 06:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
691 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 300 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 06:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
692 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1500 | mg | 500 | mL | 3 | mg/mL | Pump | Primary and Secondary | 01:20 |  | 06:00 | 01:30 | mg/kg | 5 |  | 30 | n/a
693 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 350 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
694 | OutPt InfusionCenter | vancomycin | .MG Dosing | 1750 | mg | 500 | mL | 3.5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
695 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 400 | mL | 5 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
696 | OutPt InfusionCenter | vancomycin | .MG Dosing | 2000 | mg | 500 | mL | 4 | mg/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | mg/kg | 5 |  | 30 | n/a
697 | OutPt InfusionCenter | vancomycin | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 |  | mg/kg | 5 |  | 30 | mg/mL | 1 | 3 | 5
698 | OutPt InfusionCenter | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 06:00 | 00:30 | gram | 0.5 | 2 |  | n/a
699 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 06:00 | 01:00 | gram | 0.5 | 2 |  | n/a
700 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 06:00 | 01:00 | gram | 0.5 | 2 |  | n/a
701 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 | 01:00 | gram | 0.5 | 2 |  | n/a
702 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 | 01:00 | gram | 0.5 | 2 |  | n/a
703 | OutPt InfusionCenter | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 06:00 | 01:30 | gram | 0.5 | 2 |  | n/a
704 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 06:00 | 01:30 | gram | 0.5 | 2 |  | n/a
705 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 06:00 | 01:30 | gram | 0.5 | 2 |  | n/a
706 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
707 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
708 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
709 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram | 0.5 | 2 |  | n/a
710 | OutPt InfusionCenter | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 |  | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.003 | 0.005
711 | OutPt InfusionCenter | vancomycin | GM dosing | 0.5 | gram | 100 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:25 |  | 06:00 | 00:30 | gram/kg | 0.005 |  | 0.03 | n/a
712 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 150 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 06:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
713 | OutPt InfusionCenter | vancomycin | GM dosing | 0.75 | gram | 250 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 00:40 |  | 06:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
714 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 200 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
715 | OutPt InfusionCenter | vancomycin | GM dosing | 1 | gram | 250 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 | 01:00 | gram/kg | 0.005 |  | 0.03 | n/a
716 | OutPt InfusionCenter | vancomycin | GM dosing | 1.25 | gram | 250 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:10 |  | 06:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
717 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 300 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 06:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
718 | OutPt InfusionCenter | vancomycin | GM dosing | 1.5 | gram | 500 | mL | 0.003 | gram/mL | Pump | Primary and Secondary | 01:20 |  | 06:00 | 01:30 | gram/kg | 0.005 |  | 0.03 | n/a
719 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 350 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
720 | OutPt InfusionCenter | vancomycin | GM dosing | 1.75 | gram | 500 | mL | 0.0035 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
721 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 400 | mL | 0.005 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
722 | OutPt InfusionCenter | vancomycin | GM dosing | 2 | gram | 500 | mL | 0.004 | gram/mL | Pump | Primary and Secondary | 01:45 |  | 06:00 | 02:00 | gram/kg | 0.005 |  | 0.03 | n/a
723 | OutPt InfusionCenter | vancomycin | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:50 |  | 06:00 |  | gram/kg | 0.005 |  | 0.03 | gram/mL | 0.001 | 0.003 | 0.005
724 | OutPt InfusionCenter | zoledronic acid |  | 4 | mg | 100 | mL | 0.04 | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 02:00 | 00:30 | mg | 2 | 5 |  | n/a
725 | OutPt InfusionCenter | zoledronic acid |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:25 | 02:00 | 00:30 | mg | 2 | 5 |  | mg/mL | 0.01 | 0.03 | 0.05
726 | Pediatrics | acetaminophen |  | 1000 | mg | 100 | mL | 10 | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg | 999 | 1001 |  | n/a
727 | Pediatrics | acetaminophen |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 00:30 | 00:15 | mg/kg | 7 |  | 16 | mg/mL | 9 |  | 11
728 | Pediatrics | acetylcysteine | 1st Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg/kg | 142.5 |  | 157.5 | mg/mL |  | 1 | 75
729 | Pediatrics | acetylcysteine | 1st Dose >100kg | 15000 | mg | 275 | mL | 54.5455 | mg/mL | Pump | Primary and Secondary |  | 00:30 | 02:00 | 01:00 | mg | 14999 |  | 15001 | n/a
730 | Pediatrics | acetylcysteine | 2nd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg/kg | 47.5 |  | 52.5 | mg/mL |  | 1 | 75
731 | Pediatrics | acetylcysteine | 2nd Dose >100kg | 5000 | mg | 525 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 03:35 | 05:00 | 04:00 | mg | 4999 |  | 5001 | n/a
732 | Pediatrics | acetylcysteine | 3rd Dose <=100kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg/kg | 95 |  | 105 | mg/mL |  | 1 | 75
733 | Pediatrics | acetylcysteine | 3rd Dose >100kg | 10000 | mg | 1050 | mL | 9.5238 | mg/mL | Pump | Primary and Secondary |  | 14:20 | 18:00 | 16:00 | mg | 9999 |  | 10001 | n/a
734 | Pediatrics | acyclovir |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 4 | 25 |  | mg/mL | 0.1 | 1 | 7
735 | Pediatrics | albumin 25% |  | 12.5 | gram | 50 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 00:40 | 01:30 | 01:00 | gram | 12 | 26 |  | n/a
736 | Pediatrics | albumin 25% |  | 25 | gram | 100 | mL | 0.25 | gram/mL | Pump | Primary and Secondary |  | 01:30 | 03:00 | 02:00 | gram | 12 | 26 |  | n/a
737 | Pediatrics | ampicillin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:30 |  | mg/kg | 4 | 100 |  | mg/mL | 1 | 5 | 35
738 | Pediatrics | ampicillin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:30 | 00:30 | mg | 500 | 3000 |  | mg/mL | 1 | 5 | 35
739 | Pediatrics | ampicillin/sulbact | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 |  | mg/kg | 4 | 80 |  | mg/mL | 5 | 15 | 35
740 | Pediatrics | ampicillin/sulbact | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 |  | mg | 1000 | 3000 |  | mg/mL | 5 | 15 | 35
741 | Pediatrics | antivenin crotalidae |  | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 01:00 | 11:30 | 10:00 | unit | 1 | 6 |  | unit/mL | 0.004 |  | 0.024 | Antivenin
742 | Pediatrics | azithromycin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 3 | 14 |  | mg/mL | 0.1 | 1 | 2.2
743 | Pediatrics | azithromycin | pt weight >= 40 kg | 250 | mg | 250 | mL | 1 | mg/mL | Pump | Primary and Secondary |  | 00:40 | 02:00 | 01:00 | mg | 100 | 600 |  | n/a
744 | Pediatrics | azithromycin | pt weight >= 40 kg | 500 | mg | 250 | mL | 2 | mg/mL | Pump | Primary and Secondary |  | 00:40 | 02:00 | 01:00 | mg | 100 | 600 |  | n/a
745 | Pediatrics | aztreonam | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:15 | 02:00 |  | mg/kg | 25.5 | 50 |  | mg/mL | 4 | 5 | 40
746 | Pediatrics | aztreonam | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:15 | 02:00 |  | mg | 500 | 2000 |  | mg/mL | 4 | 5 | 40
747 | Pediatrics | caffeine citrate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 06:00 |  | mg | 1 | 1200 |  | mg/mL |  | 0.01 | 10
748 | Pediatrics | caffeine citrate |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:10 | 06:00 |  | mg/kg | 1 | 22 |  | mg/mL |  | 0.01 | 10
749 | Pediatrics | caffeine/sod benzoa |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 04:00 | 01:00 | mg | 50 | 501 | 1001 | mg/mL |  | 0.24 | 6 | 0.22 Filter Needed
750 | Pediatrics | caffeine/sod benzoa |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 | 00:20 | 04:00 | 01:00 | mg/kg | 1 | 22 |  | mg/mL |  | 0.24 | 6 | 0.22 Filter Needed
751 | Pediatrics | ceFAZolin | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg/kg | 9 | 55 |  | mg/mL | 1 | 5 | 40
752 | Pediatrics | ceFAZolin | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
753 | Pediatrics | ceFAZolin | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram/kg | 0.009 | 0.055 |  | gram/mL | 0.001 | 0.005 | 0.04
754 | Pediatrics | ceFAZolin | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:00 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
755 | Pediatrics | cefepime | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 15 | 55 |  | mg/mL | 1 | 5 | 40
756 | Pediatrics | cefepime | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
757 | Pediatrics | cefepime | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | gram/kg | 0.015 | 0.055 |  | gram/mL | 0.001 | 0.005 | 0.04
758 | Pediatrics | cefepime | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
759 | Pediatrics | cefotaxime | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 22 | 80 |  | mg/mL | 5 | 10 | 55
760 | Pediatrics | cefotaxime | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 55
761 | Pediatrics | cefotaxime | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram/kg | 0.022 | 0.08 |  | gram/mL | 0.005 | 0.01 | 0.055
762 | Pediatrics | cefotaxime | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.055
763 | Pediatrics | cefOXitin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg/kg | 20 | 55 |  | mg/mL | 5 | 10 | 40
764 | Pediatrics | cefOXitin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:15 |  | 01:00 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 40
765 | Pediatrics | cefTAZidime | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | mg/kg | 25 | 67 | 135 | mg/mL | 1 | 5 | 50
766 | Pediatrics | cefTAZidime | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | mg | 250 | 2000 |  | mg/mL | 1 | 5 | 50
767 | Pediatrics | cefTAZidime | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | gram/kg | 0.025 | 0.067 | 0.135 | gram/mL | 0.001 | 0.005 | 0.05
768 | Pediatrics | cefTAZidime | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 01:30 | 00:30 | gram | 0.25 | 2 |  | gram/mL | 0.001 | 0.005 | 0.05
769 | Pediatrics | cefTRIAXone | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg/kg | 21 | 105 |  | mg/mL | 5 | 10 | 50
770 | Pediatrics | cefTRIAXone | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg | 1000 | 2000 |  | mg/mL | 5 | 10 | 50
771 | Pediatrics | cefTRIAXone | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | gram/kg | 0.021 | 0.105 |  | gram/mL | 0.005 | 0.01 | 0.05
772 | Pediatrics | cefTRIAXone | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | gram | 1 | 2 |  | gram/mL | 0.005 | 0.01 | 0.05
773 | Pediatrics | cefuroxime | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg/kg | 18 | 55 |  | mg/mL | 5 | 7.5 | 30
774 | Pediatrics | cefuroxime | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 | 00:30 | mg | 750 | 1500 |  | mg/mL | 5 | 7.5 | 30
775 | Pediatrics | chlorothiazide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg | 250 | 1000 |  | mg/mL | 0.5 | 1 | 30
776 | Pediatrics | chlorothiazide |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 01:30 |  | mg/kg | 0.4 | 11 |  | mg/mL | 0.5 | 1 | 30
777 | Pediatrics | clindamycin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg/kg | 4 | 23 |  | mg/mL | 1 | 2 | 20
778 | Pediatrics | clindamycin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | mg | 300 | 2400 |  | mg/mL | 1 | 2 | 20
779 | Pediatrics | erythromycin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg/kg | 2.5 | 15 |  | mg/mL | 1 | 2 | 5.5
780 | Pediatrics | erythromycin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 01:00 | mg | 100 | 1000 |  | mg/mL | 1 | 2 | 5.5
781 | Pediatrics | fluconazole | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 02:00 | 01:00 | mg/kg | 3 | 12 | 14 | mg/mL | 1.5 |  | 2.5
782 | Pediatrics | fluconazole | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg | 100 | 800 |  | mg/mL | 1.5 |  | 2.5
783 | Pediatrics | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:08 | 01:00 | 00:30 | mg | 100 | 1500 |  | mg/mL | 0.9 | 1 | 30
784 | Pediatrics | fosphenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:08 | 01:00 | 00:30 | mg/kg | 2 |  | 29 | mg/mL | 0.9 | 1 | 30
785 | Pediatrics | gentamicin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:15 | 00:30 | mg/kg | 1 | 10 |  | mg/mL | 0.1 | 0.5 | 11
786 | Pediatrics | levETIRAcetam | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:08 | 00:10 | 01:00 |  | mg/kg | 4.3 | 57 | 115 | mg/mL | 0.5 | 1 | 30
787 | Pediatrics | levETIRAcetam | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:08 | 00:10 | 01:00 |  | mg | 200 | 3000 | 5000 | mg/mL | 0.5 | 1 | 30
788 | Pediatrics | mannitol 20% |  | 100 | gram | 500 | mL | 0.2 | gram/mL | Pump | Primary and Secondary |  | 00:31 | 02:00 |  | gram | 1 | 101 |  | n/a
789 | Pediatrics | mannitol 20% |  | 100 | gram | 500 | mL | 0.2 | gram/mL | Pump | Primary and Secondary | 00:25 | 00:31 | 02:00 |  | gram/kg | 0.2 | 2 |  | n/a |  |  |  | In-line filter
790 | Pediatrics | mannitol 20% |  | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:10 | 02:00 |  | gram/kg | 0.2 | 2 |  | gram/mL |  | 0.1 | 0.5 | In-line filter
791 | Pediatrics | meropenem | .mg, pt weight <40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | mg/kg | 10 | 40 |  | mg/mL | 1 | 5 | 40
792 | Pediatrics | meropenem | .mg, pt weight>=40kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
793 | Pediatrics | meropenem | gram, pt wt < 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | gram/kg | 0.01 | 0.04 |  | gram/mL | 0.001 | 0.005 | 0.04
794 | Pediatrics | meropenem | gram, pt wt >= 40kg | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary | 00:05 | 00:25 | 02:00 | 00:30 | gram | 0.5 | 2 |  | gram/mL | 0.001 | 0.005 | 0.04
795 | Pediatrics | metroNIDAZOLE |  | 250 | mg | 50 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 03:00 | 01:00 | mg | 250 | 500 |  | n/a
796 | Pediatrics | metroNIDAZOLE |  | 500 | mg | 100 | mL | 5 | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 03:00 | 01:00 | mg | 250 | 500 |  | n/a
797 | Pediatrics | metroNIDAZOLE |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:25 | 00:30 | 03:00 | 01:00 | mg/kg | 7 | 16 |  | mg/mL | 4 |  | 6
798 | Pediatrics | micafungin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 0.85 | 7 |  | mg/mL | 0.4 |  | 4
799 | Pediatrics | micafungin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg | 40 | 150 |  | mg/mL | 0.4 |  | 4
800 | Pediatrics | nafcillin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg/kg | 10.6 | 57.5 |  | mg/mL | 1 | 5 | 40
801 | Pediatrics | nafcillin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:25 | 01:00 | 00:30 | mg | 500 | 2000 |  | mg/mL | 1 | 5 | 40
802 | Pediatrics | oxacillin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:08 | 01:30 |  | mg/kg | 5 | 50 |  | mg/mL | 5 | 10 | 40
803 | Pediatrics | oxacillin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:08 | 01:30 |  | mg | 500 | 2000 |  | mg/mL | 5 | 10 | 40
804 | Pediatrics | penG x Million Units |  | - - | unit | - - | mL |  | unit/mL | Pump | Primary and Secondary |  | 00:20 | 02:00 | 00:30 | unit | 1 | 6 |  | unit/mL | 0.01 | 0.02 | 0.12 | 1unit=1million
805 | Pediatrics | pentamidine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:45 |  | 04:30 | 02:00 | mg/kg | 2.9 |  | 4.1 | mg/mL |  | 0.6 | 30
806 | Pediatrics | PHENobarbital |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:05 | 00:08 | 02:00 |  | mg/kg | 0.5 | 7.5 |  | mg/mL | 0.1 | 0.5 | 135 | High Risk
807 | Pediatrics | phenytoin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:10 | 00:15 | 02:00 |  | mg/kg | 1 |  | 29 | mg/mL | 1 | 4 | 55 | In-line filter
808 | Pediatrics | piperacillin/tazo | .MG Dosing | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | mg/kg | 20 | 140 |  | mg/mL | 0.5 | 10 | 100
809 | Pediatrics | piperacillin/tazo | GM dosing | 2.25 | gram | 50 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
810 | Pediatrics | piperacillin/tazo | GM dosing | 2.25 | gram | 100 | mL | 0.0225 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
811 | Pediatrics | piperacillin/tazo | GM dosing | 3.375 | gram | 50 | mL | 0.0675 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
812 | Pediatrics | piperacillin/tazo | GM dosing | 3.375 | gram | 100 | mL | 0.0338 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
813 | Pediatrics | piperacillin/tazo | GM dosing | 4.5 | gram | 50 | mL | 0.09 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
814 | Pediatrics | piperacillin/tazo | GM dosing | 4.5 | gram | 100 | mL | 0.045 | gram/mL | Pump | Primary and Secondary |  | 00:20 | 05:00 |  | gram | 1 | 4.8 |  | n/a
815 | Pediatrics | piperacillin/tazo | GM dosing | - - | gram | - - | mL |  | gram/mL | Pump | Primary and Secondary |  | 00:25 | 04:30 |  | gram/kg | 0.025 | 0.14 |  | gram/mL | 0.005 | 0.01 | 0.1
816 | Pediatrics | potassium chloride |  | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:45 | 00:50 | 06:00 |  | mEq | 1 |  | 41 | mEq/mL | 0.04 | 0.05 | 0.2
817 | Pediatrics | potassium chloride |  | - - | mEq | - - | mL |  | mEq/mL | Pump | Primary and Secondary | 00:45 | 00:50 | 06:00 |  | mEq/kg | 0.2 | 1.5 | 2.5 | mEq/mL | 0.04 | 0.05 | 0.2 | High Risk
818 | Pediatrics | potassium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 01:30 | 03:00 | 12:30 |  | mmol | 5 | 20 | 40 | mmol/mL | 0.009 | 0.01 | 0.3
819 | Pediatrics | raNITIdine | pt weight < 50 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:10 | 01:00 | 00:30 | mg/kg | 0.4 | 2.5 |  | mg/mL | 0.1 | 0.5 | 3
820 | Pediatrics | raNITIdine | pt weight >= 50 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:03 | 00:10 | 01:00 |  | mg | 25 | 50 | 150 | mg/mL | 0.1 | 0.5 | 3
821 | Pediatrics | rifAMPin | pt weight < 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 03:00 |  | mg/kg | 4.3 | 23 |  | mg/mL | 0.5 | 1 | 7
822 | Pediatrics | rifAMPin | pt weight >= 40 kg | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 03:00 |  | mg | 300 | 600 |  | mg/mL | 0.5 | 1 | 7
823 | Pediatrics | SMX/TMP |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:30 | mg/kg | 1.7 | 8 | 10 | mg/mL | 0.01 | 0.04 | 1 | 48 Hour Tubing
824 | Pediatrics | sodium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 00:45 | 01:00 | 10:30 |  | mmol | 1 |  | 31 | mmol/mL | 0.009 | 0.01 | 0.3 | Calcium Precipitate
825 | Pediatrics | sodium phosphate |  | - - | mmol | - - | mL |  | mmol/mL | Pump | Primary and Secondary | 00:45 | 01:00 | 10:30 |  | mmol/kg | 0.04 | 0.4 | 0.6 | mmol/mL | 0.009 | 0.01 | 0.3 | Calcium Precipitate
826 | Pediatrics | tobramycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:20 | 00:25 | 01:15 |  | mg/kg | 0.8 | 7 | 12 | mg/mL | 0.01 | 0.1 | 11
827 | Pediatrics | valproate sodium |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:00 |  | mg/kg | 5 | 60 |  | mg/mL | 1 |  | 25
828 | Pediatrics | vancomycin |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:30 | 00:50 | 04:30 |  | mg/kg | 5 | 30 |  | mg/mL | 0.5 | 1 | 5
829 | Pediatrics | voriconazole |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary |  | 00:50 | 02:10 |  | mg/kg | 2.6 | 9 |  | mg/mL | 0.1 | 0.4 | 5.5
830 | Pediatrics | zidovudine |  | - - | mg | - - | mL |  | mg/mL | Pump | Primary and Secondary | 00:40 | 00:50 | 03:00 | 01:00 | mg/kg | 0.8 | 9 |  | mg/mL | 0.5 | 1 | 4
Note on Y25: Existing IV lines must be flushed with D5w prior to infusion. If unable, must start a new line. ***Infusion requires the use of an in-line filter.
Note on Y47: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y48: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y49: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y50: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y51: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y52: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y71: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y73: Use a 0.22 micron in-line filter when administering this product.
Note on Y92: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y93: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y94: Do not infuse any other drugs through this line.
Note on Y95: Do not infuse any other drugs through this line.
Note on Y96: Do not infuse any other drugs through this line.
Note on Y199: Advise patient to lie down during and 30 minutes after infusion to prevent hypotension.
Note on Y223: Protect this drug from light.
Note on Y224: Protect this drug from light.
Note on Y241: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y242: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y243: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y244: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y253: Infusion requires the use of an in-line filter.
Note on Y267: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y268: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y269: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y270: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y271: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y272: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y273: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y274: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y275: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y276: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y277: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y278: Infusion requires the use of an in-line filter.
Note on Y279: Infusion requires the use of an in-line filter.
Note on Y280: INR 2 - 4 ** Max dose of Kcentra should not exceed 2500 units **
Note on Y281: INR 4 - 6 ** Max dose of Kcentra should not exceed 3500 units **
Note on Y282: INR > 6 ** Max dose of Kcentra should not exceed 5000 units **
Note on Y294: Infusion requires the use of an in-line filter.
Note on Y295: Infusion requires the use of an in-line filter.
Note on Y296: Infusion requires the use of an in-line filter.
Note on Y337: Do not infuse any other drugs through this line.
Note on Y339: 1 unit = 1 million units
Note on Y340: 1 unit = 1 million units
Note on Y341: 1 unit = 1 million units
Note on Y342: 1 unit = 1 million units
Note on Y343: 1 unit = 1 million units
Note on Y344: 1 unit = 1 million units
Note on Y350: Infusion requires the use of an in-line filter.
Note on Y385: Premedicate before each infusion.
Note on Y386: Change tubing every 48 hours.
Note on Y388: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y389: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y467: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y468: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y469: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y470: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y471: This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
Note on Y476: Adminster over 60 minutes to decrease risk of bilirubin encephalopathy.
Note on Y482: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y483: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y484: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y485: This medication must be infused using DEHP FREE tubing.
Note on Y486: This medication must be infused using DEHP FREE tubing.
Note on Y497: MAXIMUM dose of brentuximab is 180 mg.
Note on Y498: Use a 0.22 micron in-line filter when administering this product.
Note on Y502: Infusion requires the use of an in-line filter.
Note on Y503: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y504: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y509: Use a 0.22 micron in-line filter when administering this product.
Note on Y510: Use a 0.22 micron in-line filter when administering this product.
Note on Y511: Use a 0.22 micron in-line filter when administering this product.
Note on Y513: This medication must be infused using DEHP FREE tubing.
Note on Y515: Do NOT use an in-line filter for administration.
Note on Y516: Do NOT administer through lines containing dextrose.
Note on Y517: This medication must be infused using Non-PVC tubing.
Note on Y521: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y525: Do NOT use an in-line filter for administration.
Note on Y534: Infusion requires the use of an in-line filter.
Note on Y535: Infusion requires the use of an in-line filter.
Note on Y536: Infusion requires the use of an in-line filter.
Note on Y537: Infusion requires the use of an in-line filter.
Note on Y538: Infusion requires the use of an in-line filter.
Note on Y546: Requires Low protein binding 0.2 micron filter and DEHP free bag
Note on Y547: Do NOT use an in-line filter for administration.
Note on Y548: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y549: Infusion requires the use of an in-line filter.
Note on Y550: Infusion requires the use of an in-line filter.
Note on Y554: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y556: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y557: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y558: Flush line before and after administration. Can be given on same line as sodium chloride 0.9%.
Note on Y562: Drug is a vesicant. Check blood return.
Note on Y563: Drug is a vesicant. Check blood return.
Note on Y565: Use a 0.22 micron in-line filter when administering this product.
Note on Y567: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y568: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y569: This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
Note on Y572: Existing IV lines must be flushed with D5w prior to infusion. If unable, must start a new line. ***Infusion requires the use of an in-line filter.
Note on Y573: Use a low-protein binding in-line or add-on filter between 0.2 and 5 microns.
Note on Y617: Protect this drug from light.
Note on Y618: Protect this drug from light.
Note on Y635: This medication must be infused using a 0.2- micron, low protein binding filter.
Note on Y636: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y637: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y638: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y639: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y640: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y641: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y642: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y643: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y652: Infusion requires the use of an in-line filter.
Note on Y653: Infusion requires the use of an in-line filter.
Note on Y654: Infusion requires the use of an in-line filter.
Note on Y655: Infusion requires the use of an in-line filter.
Note on Y656: Infusion requires the use of an in-line filter.
Note on Y657: Use a 0.22 micron in-line filter when administering this product.
Note on Y658: Use a 0.22 micron in-line filter when administering this product.
Note on Y666: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y667: Risk of hypersensitivity. Check vital signs frequently during infusion.
Note on Y743: 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
Note on Y751: Use a 0.22 micron in-line filter when administering this product.
Note on Y752: Use a 0.22 micron in-line filter when administering this product.
Note on Y791: Infusion requires the use of an in-line filter.
Note on Y792: Infusion requires the use of an in-line filter.
Note on Y806: 1 unit = 1 million units
Note on Y808: This is a HIGH RISK medication. 2nd RN to verify prior to initiating therapy.
Note on Y809: Infusion requires the use of an in-line filter.
Note on Y819: This is a HIGH RISK medication. 2nd RN to verify prior to initiating therapy.
Note on Y825: Change tubing every 48 hours.
Note on Y826: Do not infuse Calcium and Phosphate though the same line.
Note on Y827: Do not infuse Calcium and Phosphate though the same line.

## Fluids
Fluids
Line Item | Profile | Fluid Name | Therapy Name | Module Availability | Supports Secondary (Y/N) | Rate Units | Rate Limits (mL/h) Soft Min | Rate Limits (mL/h) Soft Max | Rate Limits (mL/h) Hard Max | Clinical Advisory
1 | Adult | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 999
2 | Adult | .IVF + KCL | 10 mEq/L | Pump | Y | mL/h | 1 | 999
3 | Adult | .IVF + KCL | 15 mEq/L | Pump | Y | mL/h | 1 | 666 | 700
4 | Adult | .IVF + KCL | 20 mEq/L | Pump | Y | mL/h | 1 | 500 | 550
5 | Adult | .IVF + KCL | 30 mEq/L | Pump | Y | mL/h | 1 | 333 | 366
6 | Adult | .IVF + KCL | 40 mEq/L | Pump | Y | mL/h | 1 | 250 | 275
7 | Adult | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 1 | 999
8 | Adult | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 500 | 550
9 | Adult | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 1 | 250 | 275
10 | Adult | .Maint IVF+Additives |  | Pump | Y | mL/h | 1 |  | 999
11 | Adult | 3% Sodium Chloride |  | Pump | Y | mL/h | 1 | 50 | 100
12 | Adult | Bicarb Drip |  | Pump | Y | mL/h | 1 |  | 999
13 | Adult | Blood (whole) |  | Pump | N | mL/h | 1 | 150 |  | Dedicated IV Line
14 | Adult | Dextrose 10% |  | Pump | Y | mL/h | 1 | 150
15 | Adult | Fresh Frozen Plasma |  | Pump | N | mL/h | 1 | 150 |  | Dedicated IV Line
16 | Adult | Insulin Maternity |  | Pump | Y | mL/h | 70 | 160 | 200
17 | Adult | MVI Only |  | Pump | Y | mL/h | 1 | 999
18 | Adult | MVI+Folate+Thiamine |  | Pump | Y | mL/h | 1 | 999
19 | Adult | Parenteral Nutrition |  | Pump | N | mL/h | 35 | 95 |  | Midline Warning
20 | Adult | PRBCs |  | Pump | N | mL/h | 1 | 150 |  | Dedicated IV Line
21 | Nursery | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 15 | 25
22 | Nursery | Dextrose 10% |  | Pump | Y | mL/h | 1 | 15 | 25
23 | ONC Infusion Center | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 999
24 | ONC Infusion Center | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 1 | 999
25 | ONC Infusion Center | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 667
26 | ONC Infusion Center | .IVF + KCL 30 mEq/L |  | Pump | Y | mL/h | 1 | 334 | 667
27 | ONC Infusion Center | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 1 | 250 | 500
28 | ONC Infusion Center | .Maint IVF+Additives |  | Pump | Y | mL/h | 1 | 250
29 | ONC Infusion Center | Dextrose 10% |  | Pump | Y | mL/h | 15 | 200
30 | ONC Infusion Center | Rituximab (Rituxan) |  | Pump | Y | mL/h | 0.1 | 250
31 | OutPt InfusionCenter | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 999
32 | OutPt InfusionCenter | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 1 | 999
33 | OutPt InfusionCenter | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 667
34 | OutPt InfusionCenter | .IVF + KCL 30 mEq/L |  | Pump | Y | mL/h | 1 | 334 | 667
35 | OutPt InfusionCenter | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 1 | 250 | 500
36 | OutPt InfusionCenter | .Maint IVF+Additives |  | Pump | Y | mL/h | 1 | 250
37 | OutPt InfusionCenter | Dextrose 10% |  | Pump | Y | mL/h | 15 | 200
38 | OutPt InfusionCenter | PRBCs |  | Pump and Syringe | N | mL/h | 1 | 150 |  | Dedicated IV Line
39 | OutPt InfusionCenter | Rituximab (Rituxan) |  | Pump | Y | mL/h | 0.1 | 250
40 | Pediatrics | ..IVF Maintenance |  | Pump | Y | mL/h | 1 | 300
41 | Pediatrics | .IVF + KCL | 10 mEq/L | Pump | Y | mL/h | 1 | 999
42 | Pediatrics | .IVF + KCL | 15 mEq/L | Pump | Y | mL/h | 1 | 666 | 700
43 | Pediatrics | .IVF + KCL | 20 mEq/L | Pump | Y | mL/h | 1 | 500 | 550
44 | Pediatrics | .IVF + KCL | 30 mEq/L | Pump | Y | mL/h | 1 | 333 | 366
45 | Pediatrics | .IVF + KCL | 40 mEq/L | Pump | Y | mL/h | 1 | 250 | 275
46 | Pediatrics | .IVF + KCL 10 mEq/L |  | Pump | Y | mL/h | 0.1 | 200 | 999
47 | Pediatrics | .IVF + KCL 20 mEq/L |  | Pump | Y | mL/h | 1 | 200 | 500
48 | Pediatrics | .IVF + KCL 40 mEq/L |  | Pump | Y | mL/h | 0.1 | 200 | 275
49 | Pediatrics | Cryoprecipitate |  | Pump | N | mL/h | 0.1 | 300
50 | Pediatrics | Dextrose 10% |  | Pump | Y | mL/h | 1 | 50 | 150
51 | Pediatrics | Fresh Frozen Plasma |  | Pump | N | mL/h | 0.1 | 150 |  | Dedicated IV Line
52 | Pediatrics | Parenteral Nutrition |  | Pump | N | mL/h | 35 | 95 |  | Midline Warning
53 | Pediatrics | PRBCs |  | Pump | N | mL/h | 0.1 | 150 |  | Dedicated IV Line
Note on K15: Do not infuse any other drugs through this line.
Note on K17: Do not infuse any other drugs through this line.
Note on K21: Do NOT infuse through a midline catheter.
Note on K22: Do not infuse any other drugs through this line.
Note on K40: Do not infuse any other drugs through this line.
Note on K53: Do not infuse any other drugs through this line.
Note on K54: Do NOT infuse through a midline catheter.
Note on K55: Do not infuse any other drugs through this line.

## PCA Drugs
PCA Drugs
Line Item | Profile | Drug Name | Therapy Name | Drug Amount | Drug Amount Units | Diluent Volume | Diluent Volume Units | Final Concentration | Final Concentration Units | Max Accumulated Lockout | Dosing Units | Max Accumulated Dose Range Soft Min | Max Accumulated Dose Range Soft Max | Max Accumulated Dose Range Hard Max | Max Accumulated Dose Range Initial Value | PCA Dose Limits Soft Min | PCA Dose Limits Soft Max | PCA Dose Limits Hard Max | PCA Dose Limits Initial Value | Lock Out Interval (minutes) Hard Min | Lock Out Interval (minutes) Soft Min | Lock Out Interval (minutes) Soft Max | Lock Out Interval (minutes) Initial Value | Continuous Dose Limits Soft Min | Continuous Dose Limits Soft Max | Continuous Dose Limits Hard Max | Continuous Dose Limits Initial Value | Bolus Dose Limits Soft Min | Bolus Dose Limits Soft Max | Bolus Dose Limits Hard Max | Bolus Dose included in Accum Max Dose (Y/N) | Loading Dose Limits Soft Min | Loading Dose Limits Soft Max | Loading Dose Limits Hard Max | Conc. Units | Conc. Limits Hard Min | Conc. Limits Soft Min | Conc. Limits Soft Max | Supports PCA Pause Protocol (Y/N) | Clinical Advisory
1 | Adult | No drugs exist in this section
2 | Nursery | No drugs exist in this section
3 | ONC Infusion Center | No drugs exist in this section
4 | OutPt InfusionCenter | No drugs exist in this section
5 | Pediatrics | No drugs exist in this section

## PC Unit Config
| PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit | PC Unit
Line Item | Profile | Alarm Audio Profile | Minimum Audio Volume | Default Audio Volume | Anesthesia Mode | Battery Meter Display | Color Display | Key Click Audio | Limit Checking | Max Patient BSA (m2) | Max Patient Weight (kg) | Patient ID Entry | Preventive Maintenance Reminder | Tamper Resistant | Authorized User Mode | No Guardrails Basic Infusion Clinical Advisory | Clinical Advisory Message
1 | Adult | Profile 1 | 3 | 4 | Enabled | Enabled | Enabled | Enabled | Always | 3 | 500 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
2 | Nursery | Profile 1 | 2 | 3 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 7 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
3 | ONC Infusion Center | Profile 1 | 3 | 4 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 500 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
4 | OutPt InfusionCenter | Profile 1 | 3 | 4 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 500 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection
5 | Pediatrics | Profile 1 | 3 | 4 | Disabled | Enabled | Enabled | Enabled | Always | 3 | 250 | Disabled | Enabled | Enabled | Disabled | Enabled | This infusion selection DOES NOT have Guardrails Protection

## Pump Config
| Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module | Pump Module
Line Item | Profile | Accumulated Air-In-Line | Air In Line (microliter) | Auto-Restart Attempts | KVO Rate Adjust (mL/h) | Max Occlusion Pressure (mmHg) | Default Pressure (mmHg) | Maximum Rate (mL/h) | Maximum VTBI (mL) | Pressure Mode Lock Status | Pressure Mode Selection | Secondary Infusion | SEC-PRI Audio Alert
1 | Adult | Enabled | 250 | 9 | 20 | 525 | 525 | 999 | 3000 | Unlocked | Pump | Enabled | Enabled
2 | Nursery | Enabled | 50 | 1 | 2 | 525 | 525 | 25 | 1200 | Unlocked | Pump | Enabled | Enabled
3 | ONC Infusion Center | Enabled | 250 | 9 | 20 | 525 | 525 | 999 | 3000 | Unlocked | Pump | Enabled | Enabled
4 | OutPt InfusionCenter | Enabled | 250 | 9 | 20 | 525 | 525 | 999 | 3000 | Unlocked | Pump | Enabled | Enabled
5 | Pediatrics | Enabled | 75 | 9 | 5 | 525 | 525 | 999 | 2200 | Unlocked | Pump | Enabled | Enabled

## Syringe Config
| Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module | Syringe Module
Line Item | Profile | All Mode | Auto Occlusion Pressure | Back-Off After Occlusion | Fast-Start | KVO Option | KVO Rate Adjust (mL/h) | KVO Volume Adjust (%) | Near End of Infusion (Continuous, Fluids, Basic) | Alert Time (Continuous, Fluids, Basic) (minutes) | Near End of Infusion (Intermittent) | Alert Time (Intermittent) (minutes) | Occlusion Pressure Limit with Disc (mmHg) | Occlusion Pressure Limit No Disc | Priming
1 | Adult | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled
2 | Nursery | Enabled | Enabled | Enabled | Enabled | Enabled | 2 | 5 | Enabled | 2 | Enabled | 2 | 1000 | High | Enabled
3 | ONC Infusion Center | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled
4 | OutPt InfusionCenter | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled
5 | Pediatrics | Disabled | Enabled | Enabled | Enabled | Disabled | n/a | n/a | Disabled | n/a | Disabled | n/a | 1000 | High | Disabled

## Pump & Syringe Shared Config
| Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe) | Shared Infusion Settings (Pump & Syringe)
Line Item | Profile | Delay Start Options | Delay Start Callback | Drug Calculation | Drug Calculation Bolus Dose | Pressure Dynamic Display | Volume/Duration
1 | Adult | Disabled | n/a | Enabled | Enabled | Enabled | Enabled
2 | Nursery | Disabled | n/a | Enabled | Enabled | Enabled | Enabled
3 | ONC Infusion Center | Disabled | n/a | Disabled | n/a | Enabled | Enabled
4 | OutPt InfusionCenter | Disabled | n/a | Disabled | n/a | Enabled | Enabled
5 | Pediatrics | Disabled | n/a | Enabled | Enabled | Enabled | Enabled

## PCA Config
| PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module | PCA Module
Line Item | Profile | Bolus Delivery Rate (mL/h) | Dose Request Cord (Profile) | Max Dose Limit | Time Window (h) | Module Location Enforcement | Near End of Infusion | Near End of Infusion Alert Time (%) | Pressure Limit | Priming | Security Access Level | Security Code
1 | Adult | 150 | 2 | Enabled | 1 | Enabled | Enabled | 5 | High | Enabled | 1 | n/a
2 | Nursery | 150 | 2 | Enabled | 4 | Enabled | Disabled | n/a | High | Enabled | 1 | n/a
3 | ONC Infusion Center | 150 | 2 | Enabled | 4 | Enabled | Disabled | n/a | High | Enabled | 1 | n/a
4 | OutPt InfusionCenter | 150 | 2 | Enabled | 4 | Enabled | Disabled | n/a | High | Enabled | 1 | n/a
5 | Pediatrics | 150 | 2 | Enabled | 1 | Enabled | Enabled | 5 | High | Enabled | 1 | n/a

## PCA Pause Protocol
| PCA Pause Protocol | PCA Pause Protocol | PCA Pause Protocol | PCA Pause Protocol | PCA Pause Protocol
Line Item | Profile | PCA Pause Protocol Enabled | Available For | Monitoring Module Attach Enforcement | PCA Pause Protocol
1 | Adult | Enabled | Etco2 | Disabled | PCA Infusion has paused due to a decline in respiratory status. Check Patient!
2 | Nursery | Disabled | n/a | n/a | n/a
3 | ONC Infusion Center | Disabled | n/a | n/a | n/a
4 | OutPt InfusionCenter | Disabled | n/a | n/a | n/a
5 | Pediatrics | Disabled | n/a | n/a | n/a

## PCA & Syringe Shared Config
| Shared Infusion Settings (PCA & Syringe) | Shared Infusion Settings (PCA & Syringe)
Line Item | Profile | Maximum Rate (mL/h) | NEOI Snooze | NEOI Snooze Time (minutes)
1 | Adult | 999 | Disabled | n/a
2 | Nursery | 25 | Enabled | 5
3 | ONC Infusion Center | 999 | Disabled | n/a
4 | OutPt InfusionCenter | 999 | Disabled | n/a
5 | Pediatrics | 999 | Disabled | n/a

## EtCO2 Config
| EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module | EtCO2 Module
Line Item | Profile | Alarm Limit Type | EtCO2 High (mmHg) | EtCO2 Low (mmHg) | Respiratory Rate High (bpm) | Respiratory Rate Low (bpm) | No Breath Alarm (Seconds) | FiCO2 High (mmHg) | Waveform Time Scale (Seconds) | PCA Pause Resp. Rate Lower Limit (bpm) | PCA Pause Resp. Rate Initial Value (bpm)
1 | Adult | Adult | 60 | 10 | 35 | 7 | 30 | 8 | 10 | 5 | 6
2 | Nursery | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a
3 | ONC Infusion Center | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a
4 | OutPt InfusionCenter | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a
5 | Pediatrics | Adult | 60 | 10 | 35 | 6 | 30 | 8 | 10 | n/a | n/a

## Master Drugs
Master Drugs | Master Drugs | Master Drugs | Master Drugs | Master Drugs
Line Item | Drug Name | Concentration | Used As | Aliases
1 | abatacept | 500 mg / 100 mL (5 mg / mL) | Intermittent | 4144595I
2 | abatacept | 750 mg / 100 mL (7.5 mg / mL) | Intermittent | 4144595I
3 | abatacept | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 4144595I
4 | abciximab | 7.2 mg / 250 mL (0.0288 mg / mL) |  | 2748001I
5 | abciximab | 7.4 mg / 250 mL (0.0296 mg / mL) |  | 2748001I
6 | abciximab | 9 mg / 250 mL (0.036 mg / mL) |  | 2748001I
7 | abciximab | --- mg / --- mL |  | 2748001I
8 | acetaminophen | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2748022I
9 | acetaminophen | --- mg / --- mL | Intermittent | 2748022I
10 | acetaminophen | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2748022I
11 | acetaZOLAMIDE | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2748400I
12 | acetaZOLAMIDE | --- mg / --- mL | Intermittent | 2748400I
13 | acetylcysteine | 5,000 mg / 525 mL (9.5238 mg / mL) | Intermittent | 2748462I
14 | acetylcysteine | 10,000 mg / 1,050 mL (9.5238 mg / mL) | Intermittent | 2748462I
15 | acetylcysteine | 15,000 mg / 275 mL (54.5455 mg / mL) | Intermittent | 2748462I
16 | acetylcysteine | --- mg / --- mL | Intermittent | 2748462I
17 | acyclovir | --- mg / --- mL | Intermittent | 2748487I
18 | adenosine (CATH LAB) | 60 mg / 50 mL (1.2 mg / mL) | Cont/Bolus or PCA | 2748529C
19 | adenosine (CATH LAB) | 90 mg / 90 mL (1 mg / mL) | Cont/Bolus or PCA | 2748529C
20 | ado-trastuzumab | --- mg / --- mL |  | 24501603I
21 | aflibercept | --- mg / --- mL |  | 19949112I
22 | agalsidase beta | --- mg / --- mL |  | 2748541I
23 | albumin 25% | 12.5 gram / 50 mL (0.25 gram / mL) | Intermittent | 2748593I
24 | albumin 25% | 25 gram / 100 mL (0.25 gram / mL) | Intermittent | 2748593I
25 | albumin 25% | --- gram / --- mL | Intermittent | 2748593I
26 | albumin 5% | 12.5 gram / 250 mL (0.05 gram / mL) |  | 2748593
27 | albumin 5% | 25 gram / 500 mL (0.05 gram / mL) |  | 2748593
28 | aldesleukin | --- unit / --- mL |  | 2748678I
29 | alemtuzumab | --- mg / --- mL |  | 2748690I
30 | alfentanil | --- mcg / --- mL |  | 2748710I
31 | allopurinol | --- mg / --- mL |  | 2748758I
32 | alprostadil | 500 mcg / 250 mL (2 mcg / mL) | Cont/Bolus or PCA | 2748821C
33 | alprostadil | --- mcg / --- mL | Cont/Bolus or PCA | 2748821C
34 | alteplase | 100 mg / 100 mL (1 mg / mL) |  | 2748851
35 | alteplase | --- mg / --- mL |  | 2748851
36 | alteplase -stroke | 81 mg / 81 mL (1 mg / mL) | Cont/Bolus or PCA | 2748851C
37 | alteplase -stroke | 90 mg / 90 mL (1 mg / mL) | Cont/Bolus or PCA | 2748851C
38 | alteplase -stroke | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2748851C
39 | alteplase -stroke | --- mg / --- mL | Cont/Bolus or PCA | 2748851C
40 | alteplase-non-stroke | 2 mg / 50 mL (0.04 mg / mL) | Intermittent | 2748851I
41 | alteplase-non-stroke | 2 mg / 100 mL (0.02 mg / mL) | Intermittent | 2748851I
42 | alteplase-non-stroke | 100 mg / 100 mL (1 mg / mL) | Intermittent | 2748851I
43 | alteplase-non-stroke | --- mg / --- mL | Intermittent | 2748851I
44 | amifostine | --- mg / --- mL |  | 2748941I
45 | amikacin | 250 mg / 250 mL (1 mg / mL) | Intermittent | 2748946I
46 | amikacin | --- mg / --- mL | Intermittent | 2748946I
47 | aminocaproic acid | --- mg / --- mL | Intermittent | 2748973I
48 | aminocaproic acid | 5 gram / 250 mL (0.02 gram / mL) | Intermittent | 2748973I
49 | aminocaproic acid | --- gram / --- mL | Intermittent | 2748973I
50 | aminophylline | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2748995C
51 | aminophylline | 500 mg / 500 mL (1 mg / mL) | Cont/Bolus or PCA | 2748995C
52 | aminophylline | 1,000 mg / 500 mL (2 mg / mL) | Cont/Bolus or PCA | 2748995C
53 | aminophylline | --- mg / --- mL | Cont/Bolus or PCA | 2748995C
54 | amiodarone | 150 mg / 100 mL (1.5 mg / mL) |  | 2749009
55 | amiodarone | 450 mg / 250 mL (1.8 mg / mL) |  | 2749009
56 | amiodarone | 900 mg / 500 mL (1.8 mg / mL) |  | 2749009
57 | amiodarone *Loading* | 150 mg / 100 mL (1.5 mg / mL) | Intermittent | 2749009I
58 | amiodarone *Loading* | 150 mg / 103 mL (1.4563 mg / mL) | Intermittent | 2749009I
59 | amiodarone Drip | 360 mg / 200 mL (1.8 mg / mL) | Cont/Bolus or PCA | 2749009C
60 | amiodarone Drip | 450 mg / 250 mL (1.8 mg / mL) | Cont/Bolus or PCA | 2749009C
61 | amiodarone Drip | 900 mg / 500 mL (1.8 mg / mL) | Cont/Bolus or PCA | 2749009C
62 | ammonium chloride | --- mEq / --- mL |  | 2749089I
63 | ampho B (LIPID) | --- mg / --- mL |  | 2749273I
64 | ampho B (liposomal) | --- mg / --- mL | Intermittent | 2749278I
65 | ampho B cholesteryl | --- mg / --- mL |  | 2749266I
66 | amphotericin B | --- mg / --- mL | Intermittent | 2749258I
67 | ampicillin | 250 mg / 50 mL (5 mg / mL) | Intermittent | 2749283I
68 | ampicillin | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2749283I
69 | ampicillin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2749283I
70 | ampicillin | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2749283I
71 | ampicillin | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2749283I
72 | ampicillin | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2749283I
73 | ampicillin | --- mg / --- mL | Intermittent | 2749283I
74 | ampicillin | 0.5 gram / 50 mL (0.01 gram / mL) | Intermittent | 2749283I
75 | ampicillin | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2749283I
76 | ampicillin | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2749283I
77 | ampicillin | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2749283I
78 | ampicillin | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2749283I
79 | ampicillin | --- gram / --- mL | Intermittent | 2749283I
80 | ampicillin/sulbact | 1,500 mg / 50 mL (30 mg / mL) | Intermittent | 2749313I
81 | ampicillin/sulbact | 1,500 mg / 100 mL (15 mg / mL) | Intermittent | 2749313I
82 | ampicillin/sulbact | 3,000 mg / 100 mL (30 mg / mL) | Intermittent | 2749313I
83 | ampicillin/sulbact | --- mg / --- mL | Intermittent | 2749313I
84 | ampicillin/sulbact | 1.5 gram / 50 mL (0.03 gram / mL) | Intermittent | 2749313I
85 | ampicillin/sulbact | 1.5 gram / 100 mL (0.015 gram / mL) | Intermittent | 2749313I
86 | ampicillin/sulbact | 3 gram / 100 mL (0.03 gram / mL) | Intermittent | 2749313I
87 | andulafungin | --- mg / --- mL | Intermittent | 4144679I
88 | antithymocyte EQUIN | --- mg / --- mL |  | 19949425I
89 | antithymocyte RABBI | --- mg / --- mL |  | 19949425
90 | antivenin crotalidae | 4 unit / 250 mL (0.016 unit / mL) | Intermittent | 2749465I
91 | antivenin crotalidae | 6 unit / 250 mL (0.024 unit / mL) | Intermittent | 2749465I
92 | antivenin crotalidae | 10 unit / 250 mL (0.04 unit / mL) | Intermittent | 2749465I
93 | antivenin crotalidae | --- unit / --- mL | Intermittent | 2749465I
94 | aprepitant | 100 mg / 100 mL (1 mg / mL) | Intermittent | 2749846I
95 | aprepitant | 130 mg / 100 mL (1.3 mg / mL) | Intermittent | 2749846I
96 | argatroban | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2749866C
97 | argatroban | 500 mg / 500 mL (1 mg / mL) | Cont/Bolus or PCA | 2749866C
98 | arsenic trioxide | --- mg / --- mL |  | 2749895I
99 | asparaginase | --- unit / --- mL |  | 2750061I
100 | atezolizumab | 800 mg / 250 mL (3.2 mg / mL) | Intermittent | 2554527703I
101 | atezolizumab | 840 mg / 250 mL (3.36 mg / mL) | Intermittent | 2554527703I
102 | atezolizumab | 1,200 mg / 250 mL (4.8 mg / mL) | Intermittent | 2554527703I
103 | atezolizumab | 1,680 mg / 250 mL (6.72 mg / mL) | Intermittent | 2554527703I
104 | atracurium | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2750296C
105 | atracurium | --- mg / --- mL | Cont/Bolus or PCA | 2750296C
106 | azaCITIDine | --- mg / --- mL |  | 3316777I
107 | azaTHIOprine | --- mg / --- mL |  | 2750417I
108 | azithromycin | 250 mg / 250 mL (1 mg / mL) | Intermittent | 2750452I
109 | azithromycin | 500 mg / 250 mL (2 mg / mL) | Intermittent | 2750452I
110 | azithromycin | --- mg / --- mL | Intermittent | 2750452I
111 | aztreonam | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2750475I
112 | aztreonam | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2750475I
113 | aztreonam | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2750475I
114 | aztreonam | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2750475I
115 | aztreonam | --- mg / --- mL | Intermittent | 2750475I
116 | aztreonam | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2750475I
117 | aztreonam | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2750475I
118 | aztreonam | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2750475I
119 | aztreonam | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2750475I
120 | bamlaniv/estesivimab | 700 mg / 110 mL (6.3636 mg / mL) | Intermittent | 1000000200I
121 | bamlaniv/estesivimab | 700 mg / 160 mL (4.375 mg / mL) | Intermittent | 1000000200I
122 | bamlaniv/estesivimab | 700 mg / 210 mL (3.3333 mg / mL) | Intermittent | 1000000200I
123 | bamlaniv/estesivimab | 700 mg / 310 mL (2.2581 mg / mL) | Intermittent | 1000000200I
124 | bamlaniv/estesivimab | --- mg / --- mL | Intermittent | 1000000200I
125 | bamlanivimab | 700 mg / 200 mL (3.5 mg / mL) | Intermittent | 3749717983I
126 | bamlanivimab | 700 mg / 270 mL (2.5926 mg / mL) | Intermittent | 3749717983I
127 | basiliximab | 20 mg / 50 mL (0.4 mg / mL) |  | 2750634I
128 | basiliximab | --- mg / --- mL |  | 2750634I
129 | belinostat | --- mg / --- mL |  | 45868857I
130 | Bendamusti (Bendeka) | --- mg / --- mL | Intermittent | 17213979bend
131 | Bendamusti (Treanda) | --- mg / --- mL | Intermittent | 17213979trea
132 | bendamustine HCl | --- mg / --- mL |  | 17213979I
133 | bevacizumab | --- mg / --- mL | Intermittent | 3316812I
134 | bezlotoxumab | --- mg / --- mL | Intermittent | 2557014467I
135 | bivalirudin | 250 mg / 50 mL (5 mg / mL) | Cont/Bolus or PCA | 2751435C
136 | bivalirudin | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2751435C
137 | bivalirudin | 250 mg / 500 mL (0.5 mg / mL) | Cont/Bolus or PCA | 2751435C
138 | bivalirudin | 500 mg / 100 mL (5 mg / mL) | Cont/Bolus or PCA | 2751435C
139 | bivalirudin | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA | 2751435C
140 | bivalirudin | 1,250 mg / 250 mL (5 mg / mL) | Cont/Bolus or PCA | 2751435C
141 | bleomycin | --- unit / --- mL | Intermittent | 2751447I
142 | bortezomib | --- mg / --- mL |  | 2751476I
143 | brentuximab | --- mg / --- mL | Intermittent | 19950376I
144 | bretylium | 1,000 mg / 250 mL (4 mg / mL) |  | 2751503I
145 | bretylium | 2,000 mg / 250 mL (8 mg / mL) |  | 2751503I
146 | bretylium | 2,000 mg / 500 mL (4 mg / mL) |  | 2751503I
147 | bumetanide | 10 mg / 100 mL (0.1 mg / mL) |  | 2751668I
148 | bumetanide | 12.5 mg / 50 mL (0.25 mg / mL) |  | 2751668I
149 | bumetanide | 12.5 mg / 250 mL (0.05 mg / mL) |  | 2751668I
150 | bumetanide | 25 mg / 100 mL (0.25 mg / mL) |  | 2751668I
151 | bumetanide | --- mg / --- mL |  | 2751668I
152 | bumetanide continuos | --- mg / --- mL | Cont/Bolus or PCA | 2751668C
153 | bupivacaine | --- mg / --- mL |  | 2751679I
154 | busulfan | --- mg / --- mL |  | 2751784I
155 | cabazitaxel | --- mg / --- mL | Intermittent | 19950393I
156 | caffeine citrate | --- mg / --- mL | Intermittent | 2751845I
157 | caffeine/sod benzoa | 500 mg / 1,000 mL (0.5 mg / mL) | Intermittent | 2751881I
158 | caffeine/sod benzoa | --- mg / --- mL | Intermittent | 2751881I
159 | calcium chloride | 1,000 mg / 60 mL (16.6667 mg / mL) | Intermittent | 2752076I
160 | calcium chloride | 1,000 mg / 110 mL (9.0909 mg / mL) | Intermittent | 2752076I
161 | calcium chloride | 2,000 mg / 120 mL (16.6667 mg / mL) | Intermittent | 2752076I
162 | calcium chloride | --- mg / --- mL | Intermittent | 2752076I
163 | calcium chloride | 1 gram / 60 mL (0.0167 gram / mL) | Intermittent | 2752076I
164 | calcium chloride | 1 gram / 110 mL (0.0091 gram / mL) | Intermittent | 2752076I
165 | calcium chloride | 2 gram / 120 mL (0.0167 gram / mL) | Intermittent | 2752076I
166 | calcium chloride | --- gram / --- mL | Intermittent | 2752076I
167 | calcium chloride | --- mEq / --- mL | Intermittent | 2752076I
168 | calcium gluconate | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752096I
169 | calcium gluconate | 1,000 mg / 60 mL (16.6667 mg / mL) | Intermittent | 2752096I
170 | calcium gluconate | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752096I
171 | calcium gluconate | 1,000 mg / 110 mL (9.0909 mg / mL) | Intermittent | 2752096I
172 | calcium gluconate | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752096I
173 | calcium gluconate | 2,000 mg / 70 mL (28.5714 mg / mL) | Intermittent | 2752096I
174 | calcium gluconate | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752096I
175 | calcium gluconate | 2,000 mg / 120 mL (16.6667 mg / mL) | Intermittent | 2752096I
176 | calcium gluconate | --- mg / --- mL | Intermittent | 2752096I
177 | calcium gluconate | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752096I
178 | calcium gluconate | 1 gram / 60 mL (0.0167 gram / mL) | Intermittent | 2752096I
179 | calcium gluconate | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752096I
180 | calcium gluconate | 1 gram / 110 mL (0.0091 gram / mL) | Intermittent | 2752096I
181 | calcium gluconate | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752096I
182 | calcium gluconate | 2 gram / 70 mL (0.0286 gram / mL) | Intermittent | 2752096I
183 | calcium gluconate | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752096I
184 | calcium gluconate | 2 gram / 120 mL (0.0167 gram / mL) | Intermittent | 2752096I
185 | calcium gluconate | --- gram / --- mL | Intermittent | 2752096I
186 | calcium gluconate | --- mEq / --- mL | Intermittent | 2752096I
187 | cangrelor | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 273246945C
188 | carbenicillin | --- mg / --- mL |  | 2752348I
189 | CARBOplatin | --- mg / --- mL | Intermittent | 2752528I
190 | carfilzomib | --- mg / --- mL | Intermittent | 24495877I
191 | carmustine | --- mg / --- mL |  | 2752551I
192 | casirivimab/imdevima | 1,200 mg / 260 mL (4.6154 mg / mL) | Intermittent | 3958645811I
193 | casirivimab/imdevima | 2,400 mg / 250 mL (9.6 mg / mL) | Intermittent | 3958645811I
194 | casirivimab/imdevima | --- mg / --- mL | Intermittent | 3958645811I
195 | caspofungin | 35 mg / 100 mL (0.35 mg / mL) | Intermittent | 2752634I
196 | caspofungin | 50 mg / 100 mL (0.5 mg / mL) | Intermittent | 2752634I
197 | caspofungin | 50 mg / 250 mL (0.2 mg / mL) | Intermittent | 2752634I
198 | caspofungin | 70 mg / 250 mL (0.28 mg / mL) | Intermittent | 2752634I
199 | caspofungin | --- mg / --- mL | Intermittent | 2752634I
200 | cefamandole | --- mg / --- mL |  | ???????????
201 | ceFAZolin | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2752689I
202 | ceFAZolin | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752689I
203 | ceFAZolin | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752689I
204 | ceFAZolin | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752689I
205 | ceFAZolin | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752689I
206 | ceFAZolin | --- mg / --- mL | Intermittent | 2752689I
207 | ceFAZolin | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752689I
208 | ceFAZolin | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752689I
209 | ceFAZolin | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752689I
210 | ceFAZolin | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752689I
211 | ceFAZolin | --- gram / --- mL | Intermittent | 2752689I
212 | cefepime | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2752723I
213 | cefepime | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2752723I
214 | cefepime | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752723I
215 | cefepime | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752723I
216 | cefepime | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752723I
217 | cefepime | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752723I
218 | cefepime | --- mg / --- mL | Intermittent | 2752723I
219 | cefepime | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752723I
220 | cefepime | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752723I
221 | cefepime | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752723I
222 | cefepime | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752723I
223 | cefepime | --- gram / --- mL | Intermittent | 2752723I
224 | cefiderocol | 750 mg / 100 mL (7.5 mg / mL) | Intermittent | 3602322021I
225 | cefiderocol | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 3602322021I
226 | cefiderocol | 1,500 mg / 100 mL (15 mg / mL) | Intermittent | 3602322021I
227 | cefiderocol | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 3602322021I
228 | cefiderocol | 0.75 gram / 100 mL (0.0075 gram / mL) | Intermittent | 3602322021I
229 | cefiderocol | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 3602322021I
230 | cefiderocol | 1.5 gram / 100 mL (0.015 gram / mL) | Intermittent | 3602322021I
231 | cefiderocol | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 3602322021I
232 | cefoperazone | --- mg / --- mL |  | ????????????
233 | cefotaxime | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752732I
234 | cefotaxime | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752732I
235 | cefotaxime | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752732I
236 | cefotaxime | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752732I
237 | cefotaxime | --- mg / --- mL | Intermittent | 2752732I
238 | cefotaxime | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752732I
239 | cefotaxime | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752732I
240 | cefotaxime | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752732I
241 | cefotaxime | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752732I
242 | cefotaxime | --- gram / --- mL | Intermittent | 2752732I
243 | cefoTEtan | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752747I
244 | cefoTEtan | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752747I
245 | cefoTEtan | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752747I
246 | cefoTEtan | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752747I
247 | cefoTEtan | --- mg / --- mL | Intermittent | 2752747I
248 | cefoTEtan | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752747I
249 | cefoTEtan | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752747I
250 | cefoTEtan | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752747I
251 | cefoTEtan | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752747I
252 | cefoTEtan | --- gram / --- mL | Intermittent | 2752747I
253 | cefOXitin | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752760I
254 | cefOXitin | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752760I
255 | cefOXitin | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752760I
256 | cefOXitin | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752760I
257 | cefOXitin | --- mg / --- mL | Intermittent | 2752760I
258 | cefOXitin | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752760I
259 | cefOXitin | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752760I
260 | cefOXitin | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752760I
261 | cefOXitin | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752760I
262 | cefOXitin | --- gram / --- mL | Intermittent | 2752760I
263 | ceftaroline | 200 mg / 50 mL (4 mg / mL) | Intermittent | 19950738I
264 | ceftaroline | 200 mg / 100 mL (2 mg / mL) | Intermittent | 19950738I
265 | ceftaroline | 200 mg / 250 mL (0.8 mg / mL) | Intermittent | 19950738I
266 | ceftaroline | 300 mg / 50 mL (6 mg / mL) | Intermittent | 19950738I
267 | ceftaroline | 300 mg / 100 mL (3 mg / mL) | Intermittent | 19950738I
268 | ceftaroline | 300 mg / 250 mL (1.2 mg / mL) | Intermittent | 19950738I
269 | ceftaroline | 400 mg / 50 mL (8 mg / mL) | Intermittent | 19950738I
270 | ceftaroline | 400 mg / 100 mL (4 mg / mL) | Intermittent | 19950738I
271 | ceftaroline | 400 mg / 250 mL (1.6 mg / mL) | Intermittent | 19950738I
272 | ceftaroline | 600 mg / 50 mL (12 mg / mL) | Intermittent | 19950738I
273 | ceftaroline | 600 mg / 100 mL (6 mg / mL) | Intermittent | 19950738I
274 | ceftaroline | 600 mg / 250 mL (2.4 mg / mL) | Intermittent | 19950738I
275 | ceftaroline | --- mg / --- mL | Intermittent | 19950738I
276 | cefTAZidime | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2752795I
277 | cefTAZidime | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2752795I
278 | cefTAZidime | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752795I
279 | cefTAZidime | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752795I
280 | cefTAZidime | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752795I
281 | cefTAZidime | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752795I
282 | cefTAZidime | --- mg / --- mL | Intermittent | 2752795I
283 | cefTAZidime | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752795I
284 | cefTAZidime | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752795I
285 | cefTAZidime | 1.5 gram / 50 mL (0.03 gram / mL) | Intermittent | 2752795I
286 | cefTAZidime | 1.5 gram / 100 mL (0.015 gram / mL) | Intermittent | 2752795I
287 | cefTAZidime | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752795I
288 | cefTAZidime | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752795I
289 | cefTAZidime | --- gram / --- mL | Intermittent | 2752795I
290 | ceftazidime/avibacta | --- gram / --- mL | Intermittent | 271758255I
291 | ceftizoxime | 1,000 mg / 50 mL (20 mg / mL) |  | 2752834I
292 | ceftizoxime | 2,000 mg / 50 mL (40 mg / mL) |  | 2752834I
293 | ceftizoxime | --- mg / --- mL |  | 2752834I
294 | ceftizoxime | 1 gram / 50 mL (0.02 gram / mL) |  | 2752834I
295 | ceftizoxime | 2 gram / 50 mL (0.04 gram / mL) |  | 2752834I
296 | ceftizoxime | --- gram / --- mL |  | 2752834I
297 | ceftolozane/tazo | 150 mg / 100 mL (1.5 mg / mL) | Intermittent | 202016343I
298 | ceftolozane/tazo | 375 mg / 100 mL (3.75 mg / mL) | Intermittent | 202016343I
299 | ceftolozane/tazo | 750 mg / 100 mL (7.5 mg / mL) | Intermittent | 202016343I
300 | ceftolozane/tazo | 1,500 mg / 100 mL (15 mg / mL) | Intermittent | 202016343I
301 | ceftolozane/tazo | 3,000 mg / 100 mL (30 mg / mL) | Intermittent | 202016343I
302 | ceftolozane/tazo | --- mg / --- mL | Intermittent | 202016343I
303 | ceftolozane/tazo | 1.5 gram / 100 mL (0.015 gram / mL) | Intermittent | 202016343I
304 | ceftolozane/tazo | 3 gram / 100 mL (0.03 gram / mL) | Intermittent | 202016343I
305 | ceftolozane/tazo | --- gram / --- mL | Intermittent | 202016343I
306 | cefTRIAXone | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2752847I
307 | cefTRIAXone | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2752847I
308 | cefTRIAXone | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2752847I
309 | cefTRIAXone | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2752847I
310 | cefTRIAXone | --- mg / --- mL | Intermittent | 2752847I
311 | cefTRIAXone | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2752847I
312 | cefTRIAXone | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2752847I
313 | cefTRIAXone | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2752847I
314 | cefTRIAXone | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2752847I
315 | cefTRIAXone | --- gram / --- mL | Intermittent | 2752847I
316 | cefuroxime | 750 mg / 50 mL (15 mg / mL) | Intermittent | 2752869I
317 | cefuroxime | 750 mg / 100 mL (7.5 mg / mL) | Intermittent | 2752869I
318 | cefuroxime | 1,500 mg / 50 mL (30 mg / mL) | Intermittent | 2752869I
319 | cefuroxime | 1,500 mg / 100 mL (15 mg / mL) | Intermittent | 2752869I
320 | cefuroxime | --- mg / --- mL | Intermittent | 2752869I
321 | cefuroxime | 0.75 gram / 50 mL (0.015 gram / mL) | Intermittent | 2752869I
322 | cefuroxime | 0.75 gram / 100 mL (0.0075 gram / mL) | Intermittent | 2752869I
323 | cefuroxime | 1.5 gram / 50 mL (0.03 gram / mL) | Intermittent | 2752869I
324 | cefuroxime | 1.5 gram / 100 mL (0.015 gram / mL) | Intermittent | 2752869I
325 | cefuroxime | --- gram / --- mL | Intermittent | 2752869I
326 | cemiplimab | --- mg / --- mL | Intermittent | 2970621255I
327 | cetuximab | --- mg / --- mL | Intermittent | 3316890I
328 | chloramphenicol | --- mg / --- mL | Intermittent | 2753038I
329 | chlorothiazide | --- mg / --- mL | Intermittent | 2753128I
330 | chlorproMAZINE | --- mg / --- mL | Intermittent | 2753638I
331 | cidofovir | --- mg / --- mL |  | 2753804I
332 | cimetidine | 300 mg / 50 mL (6 mg / mL) |  | 2753818I
333 | cimetidine | --- mg / --- mL |  | 2753818I
334 | ciprofloxacin | 200 mg / 100 mL (2 mg / mL) | Intermittent | 2753846I
335 | ciprofloxacin | 400 mg / 200 mL (2 mg / mL) | Intermittent | 2753846I
336 | ciprofloxacin | --- mg / --- mL | Intermittent | 2753846I
337 | cisatracurium | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2753896C
338 | cisatracurium | 100 mg / 250 mL (0.4 mg / mL) | Cont/Bolus or PCA | 2753896C
339 | cisatracurium | 200 mg / 200 mL (1 mg / mL) | Cont/Bolus or PCA | 2753896C
340 | cisatracurium | 200 mg / 250 mL (0.8 mg / mL) | Cont/Bolus or PCA | 2753896C
341 | cisatracurium | --- mg / --- mL | Cont/Bolus or PCA | 2753896C
342 | CISplatin | --- mg / --- mL | Intermittent | 2753903I
343 | cladribine | --- mg / --- mL |  | 2753956I
344 | clindamycin | 300 mg / 50 mL (6 mg / mL) | Intermittent | 2753995I
345 | clindamycin | 300 mg / 100 mL (3 mg / mL) | Intermittent | 2753995I
346 | clindamycin | 600 mg / 50 mL (12 mg / mL) | Intermittent | 2753995I
347 | clindamycin | 600 mg / 100 mL (6 mg / mL) | Intermittent | 2753995I
348 | clindamycin | 900 mg / 50 mL (18 mg / mL) | Intermittent | 2753995I
349 | clindamycin | 900 mg / 100 mL (9 mg / mL) | Intermittent | 2753995I
350 | clindamycin | --- mg / --- mL | Intermittent | 2753995I
351 | clodronate | --- mg / --- mL |  | 4333632I
352 | clofarabine | --- mg / --- mL |  | 3617188I
353 | cloxacillin | --- mg / --- mL |  | 2754214I
354 | CMV immunoglobulin | --- mg / --- mL |  | 2754840I
355 | colistimethate | --- mg / --- mL | Intermittent | 2754536I
356 | conivaptan | 20 mg / 100 mL (0.2 mg / mL) | Cont/Bolus or PCA | 4145032C
357 | conivaptan | 20 mg / 250 mL (0.08 mg / mL) | Cont/Bolus or PCA | 4145032C
358 | conivaptan | --- mg / --- mL | Cont/Bolus or PCA | 4145032C
359 | conivaptan Loading | 20 mg / 100 mL (0.2 mg / mL) | Intermittent | 4145032I
360 | cotrimoxazole | --- mg / --- mL |  | ????????
361 | cyclophosphamide | --- mg / --- mL | Intermittent | 2754749I
362 | cycloSPORINE | --- mg / --- mL |  | 2754780I
363 | cycloSPORINE cont | --- mg / --- mL |  | 2754780C
364 | cytarabine | --- mg / --- mL |  | 2754823I
365 | dacarbazine | --- mg / --- mL | Intermittent | 2754845I
366 | daclizumab | --- mg / --- mL |  | 2754853I
367 | DACTINomycin | --- mcg / --- mL |  | 2754858I
368 | danaparoid | --- unit / --- mL |  | 4333829I
369 | dantrolene | --- mg / --- mL | Intermittent | 2754894I
370 | DAPTOmycin | --- mg / --- mL | Intermittent | 2754915I
371 | daratumumab | --- mg / --- mL | Intermittent | 275110833I
372 | DAUNOrubicin | --- mg / --- mL |  | 2754939I
373 | DAUNOrubicin liposo | --- mg / --- mL |  | 2754945I
374 | decitabine | --- mg / --- mL |  | 4145048I
375 | deferoxamine | --- mg / --- mL | Intermittent | 2754950I
376 | denileukin | --- mcg / --- mL |  | 2754986I
377 | desferoxamine | --- mg / --- mL |  | 2754950
378 | desmopressin | --- mcg / --- mL | Intermittent | 2755032I
379 | dexamethasone | --- mg / --- mL | Intermittent | 2755093I
380 | dexmedetomidine | 200 mcg / 50 mL (4 mcg / mL) | Cont/Bolus or PCA | 2755261C
381 | dexmedetomidine | 200 mcg / 100 mL (2 mcg / mL) | Cont/Bolus or PCA | 2755261C
382 | dexmedetomidine | 400 mcg / 100 mL (4 mcg / mL) | Cont/Bolus or PCA | 2755261C
383 | dexmedetomidine | --- mcg / --- mL | Cont/Bolus or PCA | 2755261C
384 | dexrazoxane | --- mg / --- mL |  | 2755283I
385 | dextran 40 | 50 gram / 500 mL (0.1 gram / mL) |  | 2755295I
386 | dextran 40 | --- gram / --- mL |  | 2755295I
387 | diazePAM | 50 mg / 50 mL (1 mg / mL) |  | 2755701I
388 | diazePAM | --- mg / --- mL |  | 2755701I
389 | digoxin | --- mcg / --- mL |  | 2755862I
390 | digoxin immune fab | --- mg / --- mL | Intermittent | 2755888I
391 | dihydroergotamine | --- mg / --- mL |  | 2755900I
392 | dilTIAZem | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2755935C
393 | dilTIAZem | 125 mg / 125 mL (1 mg / mL) | Cont/Bolus or PCA | 2755935C
394 | dilTIAZem | --- mg / --- mL | Cont/Bolus or PCA | 2755935C
395 | dimenhyDRINATE | --- mg / --- mL |  | 2756003I
396 | diphenhydrAMINE | --- mg / --- mL |  | 2756043I
397 | dipyridamole | --- mg / --- mL |  | 2756254I
398 | DOBUTamine | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2756310C
399 | DOBUTamine | 500 mg / 250 mL (2 mg / mL) | Cont/Bolus or PCA | 2756310C
400 | DOBUTamine | 500 mg / 500 mL (1 mg / mL) | Cont/Bolus or PCA | 2756310C
401 | DOBUTamine | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA | 2756310C
402 | DOBUTamine | 1,000 mg / 500 mL (2 mg / mL) | Cont/Bolus or PCA | 2756310C
403 | DOBUTamine | --- mg / --- mL | Cont/Bolus or PCA | 2756310C,2756310I
404 | Dobutamine Stress | --- mg / --- mL |  | 2756310
405 | DOCEtaxel | --- mg / --- mL | Intermittent | 2756319I
406 | dolasetron | --- mg / --- mL |  | 2756447I
407 | DOPamine | 400 mg / 250 mL (1.6 mg / mL) | Cont/Bolus or PCA | 2756465C
408 | DOPamine | 800 mg / 250 mL (3.2 mg / mL) | Cont/Bolus or PCA | 2756465C
409 | DOPamine | --- mg / --- mL | Cont/Bolus or PCA | 2756465C
410 | doripenem | 250 mg / 50 mL (5 mg / mL) | Intermittent | 4316756I
411 | doripenem | 250 mg / 100 mL (2.5 mg / mL) | Intermittent | 4316756I
412 | doripenem | 500 mg / 100 mL (5 mg / mL) | Intermittent | 4316756I
413 | doxaCURIUM | --- mg / --- mL |  | 2756488I
414 | doxapram | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2756493C
415 | doxapram | --- mg / --- mL | Cont/Bolus or PCA | 2756493C
416 | DOXOrubicin | --- mg / --- mL | Intermittent | 2756540I
417 | DOXOrubicin, liposo | --- mg / --- mL | Intermittent | 2756552I
418 | doxycycline | 100 mg / 100 mL (1 mg / mL) | Intermittent | 2756557I
419 | doxycycline | 200 mg / 250 mL (0.8 mg / mL) | Intermittent | 2756557I
420 | doxycycline | --- mg / --- mL | Intermittent | 2756557I
421 | droperidol | --- mg / --- mL |  | 2756618I
422 | drotrecogin | --- mg / --- mL |  | 2756633I
423 | eculizumab | 300 mg / 60 mL (5 mg / mL) |  | 4245702I
424 | eculizumab | 600 mg / 120 mL (5 mg / mL) |  | 4245702I
425 | eculizumab | 900 mg / 180 mL (5 mg / mL) |  | 4245702I
426 | eculizumab | 1,200 mg / 240 mL (5 mg / mL) |  | 4245702I
427 | eculizumab | --- mg / --- mL |  | 4245702I
428 | EDTA | --- mg / --- mL |  | edta
429 | enalaprilat | --- mcg / --- mL |  | 2757060I
430 | epidural w/ fentanyl | 100 unit / 100 mL (1 unit / mL) | Cont/Bolus or PCA | epiduralfent
431 | epidural w/o fentany | 100 unit / 100 mL (1 unit / mL) | Cont/Bolus or PCA | epidnofent
432 | EPINEPHrine | 4 mg / 250 mL (0.016 mg / mL) | Cont/Bolus or PCA | 2757140C
433 | EPINEPHrine | 5 mg / 250 mL (0.02 mg / mL) | Cont/Bolus or PCA | 2757140C
434 | EPINEPHrine | 8 mg / 250 mL (0.032 mg / mL) | Cont/Bolus or PCA | 2757140C
435 | EPINEPHrine | 10 mg / 500 mL (0.02 mg / mL) | Cont/Bolus or PCA | 2757140C
436 | EPINEPHrine | 16 mg / 250 mL (0.064 mg / mL) | Cont/Bolus or PCA | 2757140C
437 | EPINEPHrine | 30 mg / 250 mL (0.12 mg / mL) | Cont/Bolus or PCA | 2757140C
438 | EPINEPHrine | --- mg / --- mL | Cont/Bolus or PCA | 2757140C
439 | epiRUBicin | --- mg / --- mL |  | 2757216I
440 | epoprostenol | 500 mcg / 100 mL (5 mcg / mL) | Cont/Bolus or PCA | 2757252C
441 | epoprostenol | 1,000 mcg / 100 mL (10 mcg / mL) | Cont/Bolus or PCA | 2757252C
442 | epoprostenol | 3,000 mcg / 100 mL (30 mcg / mL) | Cont/Bolus or PCA | 2757252C
443 | epoprostenol | --- mcg / --- mL | Cont/Bolus or PCA | 2757252C
444 | eptifibatide | 75 mg / 100 mL (0.75 mg / mL) | Cont/Bolus or PCA | 2757273C
445 | eptifibatide | 200 mg / 100 mL (2 mg / mL) | Cont/Bolus or PCA | 2757273C
446 | eriBULin | --- mg / --- mL | Intermittent | 19952644I
447 | ertapenem | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2757309I
448 | ertapenem | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2757309I
449 | ertapenem | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2757309I
450 | ertapenem | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2757309I
451 | ertapenem | 0.5 gram / 50 mL (0.01 gram / mL) | Intermittent | 2757309I
452 | ertapenem | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2757309I
453 | ertapenem | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2757309I
454 | ertapenem | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2757309I
455 | erythromycin | 250 mg / 100 mL (2.5 mg / mL) | Intermittent | 2757314I
456 | erythromycin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2757314I
457 | erythromycin | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 2757314I
458 | erythromycin | --- mg / --- mL | Intermittent | 2757314I
459 | erythromycin | --- gram / --- mL | Intermittent | 2757314I
460 | esmolol | 2,500 mg / 250 mL (10 mg / mL) | Cont/Bolus or PCA | 2757437C
461 | esmolol | 5,000 mg / 250 mL (20 mg / mL) | Cont/Bolus or PCA | 2757437C
462 | esmolol | --- mg / --- mL | Cont/Bolus or PCA | 2757437C
463 | esomeprazole | --- mg / --- mL |  | 2757446C,2757446I
464 | ethanol 10% | --- gram / --- mL |  | 2757676I
465 | ethanol 5% | --- gram / --- mL |  | 2757676
466 | etidronate | --- mg / --- mL |  | 2757876I
467 | etomidate | --- mg / --- mL |  | 2757905I
468 | etoposide | --- mg / --- mL | Intermittent | 2757912I
469 | factor IX | --- unit / --- mL |  | 2757938I
470 | factor VIII | --- unit / --- mL |  | 24606459I
471 | factor VIIIa | --- unit / --- mL |  | 24606459
472 | famotidine | 20 mg / 50 mL (0.4 mg / mL) | Intermittent | 2757966I
473 | famotidine | --- mg / --- mL | Intermittent | 2757966I
474 | fat emulsion 10% | 10 gram / 100 mL (0.1 gram / mL) |  | 2757989
475 | fat emulsion 20% | 20 gram / 100 mL (0.2 gram / mL) | Intermittent | 2757989I
476 | fat emulsion 20% | 25 gram / 125 mL (0.2 gram / mL) | Intermittent | 2757989I
477 | fat emulsion 20% | 50 gram / 250 mL (0.2 gram / mL) | Intermittent | 2757989I
478 | fat emulsion 30% | 30 gram / 100 mL (0.3 gram / mL) |  | 2757989-30%
479 | fenoldopam | 10 mg / 250 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2758046C
480 | fentaNYL | 1,000 mcg / 100 mL (10 mcg / mL) | Cont/Bolus or PCA | 2758060C
481 | fentaNYL | 1,250 mcg / 25 mL (50 mcg / mL) | Cont/Bolus or PCA | 2758060C
482 | fentaNYL | 2,500 mcg / 100 mL (25 mcg / mL) | Cont/Bolus or PCA | 2758060C
483 | fentaNYL | 2,500 mcg / 250 mL (10 mcg / mL) | Cont/Bolus or PCA | 2758060C
484 | fentaNYL | --- mcg / --- mL | Cont/Bolus or PCA | 2758060C
485 | ferric carboxymaltos | 750 mg / 265 mL (2.8302 mg / mL) | Intermittent | 24606832I
486 | ferric carboxymaltos | --- mg / --- mL | Intermittent | 24606832I
487 | ferric gluconate Na | 125 mg / 100 mL (1.25 mg / mL) |  | 2758129I
488 | filgrastim | --- mcg / --- mL |  | 2758220I
489 | floxuridine | --- mg / --- mL |  | 2758251I
490 | fluconazole | 100 mg / 50 mL (2 mg / mL) | Intermittent | 2758256I
491 | fluconazole | 200 mg / 100 mL (2 mg / mL) | Intermittent | 2758256I
492 | fluconazole | 400 mg / 200 mL (2 mg / mL) | Intermittent | 2758256I
493 | fluconazole | 800 mg / 400 mL (2 mg / mL) | Intermittent | 2758256I
494 | fluconazole | --- mg / --- mL | Intermittent | 2758256I
495 | fludarabine | --- mg / --- mL |  | 2758282I
496 | fluorouracil | --- mg / --- mL | Intermittent | 2758517C,2758517I
497 | folic acid | --- mg / --- mL | Intermittent | 2758687I
498 | fomepizole | --- mg / --- mL | Intermittent | 2758716I
499 | fosaprepitant | 150 mg / 150 mL (1 mg / mL) | Intermittent | 17213916I
500 | foscarnet | 100 mg / 50 mL (2 mg / mL) |  | 2758754I
501 | foscarnet | --- mg / --- mL |  | 2758754I
502 | fosphenytoin | --- mg / --- mL | Intermittent | 2758779I
503 | furosemide | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2758798C
504 | furosemide | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2758798C
505 | furosemide | 300 mg / 30 mL (10 mg / mL) | Cont/Bolus or PCA | 2758798C
506 | furosemide | 500 mg / 100 mL (5 mg / mL) | Cont/Bolus or PCA | 2758798C
507 | furosemide | --- mg / --- mL | Cont/Bolus or PCA | 2758798C
508 | ganciclovir | --- mg / --- mL | Intermittent | 2758863I
509 | gancyclovir | --- mg / --- mL |  | 2758863
510 | ganisetron | --- mcg / --- mL |  | 2759194
511 | gatifloxacin | 200 mg / 100 mL (2 mg / mL) |  | 2758883I
512 | gatifloxacin | 400 mg / 200 mL (2 mg / mL) |  | 2758883I
513 | gatifloxacin | --- mg / --- mL |  | 2758883I
514 | gemcitabine | --- mg / --- mL | Intermittent | 2758906I
515 | gemtuzumab | --- mg / --- mL |  | 2758923I
516 | gentamicin | 80 mg / 100 mL (0.8 mg / mL) | Intermittent | 2758928I
517 | gentamicin | 120 mg / 100 mL (1.2 mg / mL) | Intermittent | 2758928I
518 | gentamicin | --- mg / --- mL | Intermittent | 2758928I
519 | glucagon | 10 mg / 50 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2759022C
520 | glucagon | 10 mg / 100 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2759022C
521 | glucagon | --- mg / --- mL | Cont/Bolus or PCA | 2759022C
522 | golimumab | --- mg / --- mL | Intermittent | 19952385I
523 | granisetron | --- mg / --- mL |  | 2759194I
524 | haloperidol | --- mg / --- mL |  | 2759713I
525 | heparin | 100 unit / 1 mL (100 unit / mL) | Cont/Bolus or PCA | 2759747C
526 | heparin | 1,250 unit / 50 mL (25 unit / mL) | Cont/Bolus or PCA | 2759747C
527 | heparin | 2,500 unit / 25 mL (100 unit / mL) | Cont/Bolus or PCA | 2759747C
528 | heparin | 2,500 unit / 50 mL (50 unit / mL) | Cont/Bolus or PCA | 2759747C
529 | heparin | 5,000 unit / 50 mL (100 unit / mL) | Cont/Bolus or PCA | 2759747C
530 | heparin | 5,000 unit / 100 mL (50 unit / mL) | Cont/Bolus or PCA | 2759747C
531 | heparin | 25,000 unit / 250 mL (100 unit / mL) | Cont/Bolus or PCA | 2759747C
532 | heparin | 25,000 unit / 500 mL (50 unit / mL) | Cont/Bolus or PCA | 2759747C
533 | heparin | --- unit / --- mL | Cont/Bolus or PCA | 2759747C
534 | hetastarch 6% | 30 gram / 500 mL (0.06 gram / mL) | Intermittent | 2759815C,2759815I
535 | hydrALAZINE | --- mg / --- mL |  | 2759877I
536 | hydrochloric acid | 0.1 mEq / 1 mL (0.1 mEq / mL) |  | HCL
537 | hydrochloric acid | 0.2 mEq / 1 mL (0.2 mEq / mL) |  | HCL
538 | hydrochloric acid | --- mEq / --- mL |  | HCL
539 | hydrocortisone | --- mg / --- mL | Intermittent | 2760102I
540 | HYDROmorphone | 20 mg / 100 mL (0.2 mg / mL) |  | 2760421C
541 | HYDROmorphone | 100 mg / 100 mL (1 mg / mL) |  | 2760421C
542 | HYDROmorphone | --- mg / --- mL |  | 2760421C
543 | ibuprofen lysine | --- mg / --- mL |  | 2760660I
544 | ibutilide | 0.5 mg / 50 mL (0.01 mg / mL) | Intermittent | 2760741I
545 | ibutilide | 0.5 mg / 100 mL (0.005 mg / mL) | Intermittent | 2760741I
546 | ibutilide | 1 mg / 50 mL (0.02 mg / mL) | Intermittent | 2760741I
547 | ibutilide | 1 mg / 60 mL (0.0167 mg / mL) | Intermittent | 2760741I
548 | ibutilide | 1 mg / 100 mL (0.01 mg / mL) | Intermittent | 2760741I
549 | ibutilide | --- mg / --- mL | Intermittent | 2760741I
550 | IDArubicin | --- mg / --- mL |  | 2760752I
551 | Ifex / Mesna | --- mg / --- mL |  | 2760768I
552 | Ifex / Mesna | --- gram / --- mL |  | 2760768I
553 | ifosfamide | --- mg / --- mL | Intermittent | 2760761I
554 | ifosfamide | --- gram / --- mL | Intermittent | 2760761I
555 | iloprost | --- mcg / --- mL |  | 3617215I
556 | imiglucerase | --- unit / --- mL |  | 2760783I
557 | imipenem/cilastatin | 250 mg / 100 mL (2.5 mg / mL) | Intermittent | 2760792I
558 | imipenem/cilastatin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2760792I
559 | imipenem/cilastatin | 750 mg / 250 mL (3 mg / mL) | Intermittent | 2760792I
560 | imipenem/cilastatin | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 2760792I
561 | imipenem/cilastatin | --- mg / --- mL | Intermittent | 2760792I
562 | imipenem/cilastatin | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2760792I
563 | imipenem/cilastatin | 1 gram / 250 mL (0.004 gram / mL) | Intermittent | 2760792I
564 | imipenem/cilastatin | --- gram / --- mL | Intermittent | 2760792I
565 | immune glob (IVIG) | 5 gram / 50 mL (0.1 gram / mL) | Intermittent | 2760832I
566 | immune glob (IVIG) | 5 gram / 100 mL (0.05 gram / mL) | Intermittent | 2760832I
567 | immune glob (IVIG) | 10 gram / 100 mL (0.1 gram / mL) | Intermittent | 2760832I
568 | immune glob (IVIG) | 10 gram / 200 mL (0.05 gram / mL) | Intermittent | 2760832I
569 | immune glob (IVIG) | 20 gram / 200 mL (0.1 gram / mL) | Intermittent | 2760832I
570 | immune glob (IVIG) | --- gram / --- mL | Intermittent | 2760832I
571 | immune globulin Rho | --- gram / --- mL |  | 2768135I
572 | inamrinone | 100 mg / 100 mL (1 mg / mL) |  | 2760881I
573 | inamrinone | 300 mg / 60 mL (5 mg / mL) |  | 2760881I
574 | indomethacin | --- mg / --- mL |  | 2760908I
575 | inFLIXimab | --- mg / --- mL | Intermittent | 2760925I
576 | insulin | 10 unit / 1,000 mL (0.01 unit / mL) | Cont/Bolus or PCA | 2761029C
577 | insulin | 100 unit / 100 mL (1 unit / mL) | Cont/Bolus or PCA | 2761029C
578 | insulin | 250 unit / 250 mL (1 unit / mL) | Cont/Bolus or PCA | 2761029C
579 | insulin | --- unit / --- mL | Cont/Bolus or PCA | 2761029C
580 | interferon alpha-2b | --- unit / --- mL |  | 2761065I
581 | ipilimumab | --- mg / --- mL | Intermittent | 19952974I
582 | irinotecan | --- mg / --- mL | Intermittent | 2761340I
583 | irinotecan liposomal | --- mg / --- mL | Intermittent | 275110257I
584 | iron dextran | 25 mg / 50 mL (0.5 mg / mL) |  | 2761345I
585 | iron dextran | --- mg / --- mL |  | 2761345I
586 | iron sucrose | 200 mg / 110 mL (1.8182 mg / mL) | Intermittent | 2761376I
587 | iron sucrose | 500 mg / 275 mL (1.8182 mg / mL) | Intermittent | 2761376I
588 | iron sucrose | --- mg / --- mL | Intermittent | 2761376I
589 | isoproterenol | 1 mg / 250 mL (0.004 mg / mL) | Cont/Bolus or PCA | 2761419C
590 | isoproterenol | 2 mg / 250 mL (0.008 mg / mL) | Cont/Bolus or PCA | 2761419C
591 | isoproterenol | 2 mg / 500 mL (0.004 mg / mL) | Cont/Bolus or PCA | 2761419C
592 | isoproterenol | --- mg / --- mL | Cont/Bolus or PCA | 2761419C
593 | itraconazole | 200 mg / 60 mL (3.3333 mg / mL) | Intermittent | 2761517I
594 | itraconazole | --- mg / --- mL | Intermittent | 2761517I
595 | ixabepilone | --- mg / --- mL |  | 4316936I
596 | Kcentra (Factor IX) | --- unit / --- mL | Intermittent | 24608305I
597 | ketamine | --- mg / --- mL |  | 2761553I
598 | ketamine ANES | 50 mg / 50 mL (1 mg / mL) | Cont/Bolus or PCA | 2761553C
599 | ketamine ANES | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2761553C
600 | ketamine ANES | --- mg / --- mL | Cont/Bolus or PCA | 2761553C
601 | Kinevac (sincalide) | --- mcg / --- mL | Intermittent | 2768854I
602 | Kinevac (sincalide) | --- mg / --- mL | Intermittent | 2768854I
603 | labetalol | 200 mg / 200 mL (1 mg / mL) | Cont/Bolus or PCA | 2761620C
604 | labetalol | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA | 2761620C
605 | labetalol | 500 mg / 250 mL (2 mg / mL) | Cont/Bolus or PCA | 2761620C
606 | labetalol | --- mg / --- mL | Cont/Bolus or PCA | 2761620C
607 | lacosamide | --- mg / --- mL | Intermittent | 19952919I
608 | lepirudin | 100 mg / 250 mL (0.4 mg / mL) |  | 2761819I
609 | lepirudin | --- mg / --- mL |  | 2761819I
610 | leucovorin | --- mg / --- mL | Intermittent | 2761829I
611 | levETIRAcetam | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2761878I
612 | levETIRAcetam | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2761878I
613 | levETIRAcetam | 1,500 mg / 100 mL (15 mg / mL) | Intermittent | 2761878I
614 | levETIRAcetam | --- mg / --- mL | Intermittent | 2761878I
615 | levETIRAcetam | --- gram / --- mL | Intermittent | 2761878I
616 | levobupivicaine | --- mg / --- mL |  | 2761905I
617 | levOCARNitine | --- mg / --- mL |  | 2761919I
618 | levoFLOXacin | 250 mg / 50 mL (5 mg / mL) | Intermittent | 2761938I
619 | levoFLOXacin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2761938I
620 | levoFLOXacin | 750 mg / 150 mL (5 mg / mL) | Intermittent | 2761938I
621 | levoFLOXacin | --- mg / --- mL | Intermittent | 2761938I
622 | LEVOleucovorin | --- mg / --- mL |  | 17213770I
623 | levothyroxine | --- mcg / --- mL |  | 2761978I
624 | lidocaine | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA | 2762076C
625 | lidocaine | 2,000 mg / 500 mL (4 mg / mL) | Cont/Bolus or PCA | 2762076C
626 | lidocaine | --- mg / --- mL | Cont/Bolus or PCA | 2762076C
627 | linezolid | 200 mg / 100 mL (2 mg / mL) | Intermittent | 2762196I
628 | linezolid | 600 mg / 300 mL (2 mg / mL) | Intermittent | 2762196I
629 | linezolid | --- mg / --- mL | Intermittent | 2762196I
630 | liothyronine | --- mcg / --- mL |  | 2762207I
631 | LORazepam | 20 mg / 100 mL (0.2 mg / mL) |  | 2762369C
632 | LORazepam | 40 mg / 100 mL (0.4 mg / mL) |  | 2762369C
633 | LORazepam | 120 mg / 60 mL (2 mg / mL) |  | 2762369C
634 | LORazepam | --- mg / --- mL |  | 2762369C
635 | mag OB maintenance | 20 gram / 500 mL (0.04 gram / mL) | Cont/Bolus or PCA | 2762580C
636 | mag sulfate | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2762580I
637 | mag sulfate | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2762580I
638 | mag sulfate | 4 gram / 100 mL (0.04 gram / mL) | Intermittent | 2762580I
639 | mag sulfate | --- gram / --- mL | Intermittent | 2762580I
640 | magnesium sulfate | 1 gram / 50 mL (0.02 gram / mL) |  | 2762580
641 | magnesium sulfate | 4 gram / 100 mL (0.04 gram / mL) |  | 2762580
642 | magnesium sulfate | 40 gram / 500 mL (0.08 gram / mL) |  | 2762580
643 | magnesium sulfate | --- gram / --- mL |  | 2762580
644 | mannitol 20% | 50 gram / 250 mL (0.2 gram / mL) | Intermittent | 2762620I
645 | mannitol 20% | 100 gram / 500 mL (0.2 gram / mL) | Intermittent | 2762620I
646 | mannitol 20% | --- gram / --- mL | Intermittent | 2762620I
647 | mannitol 10% | 100 gram / 1,000 mL (0.1 gram / mL) |  | 2762620
648 | mannitol 10% | --- gram / --- mL |  | 2762620
649 | mannitol 15% | 75 gram / 500 mL (0.15 gram / mL) |  | 2762620-15
650 | mannitol 15% | --- gram / --- mL |  | 2762620-15
651 | mannitol 20% | 50 gram / 250 mL (0.2 gram / mL) |  | 2762620-20
652 | mannitol 20% | --- gram / --- mL |  | 2762620-20
653 | mannitol 25% | 12.5 gram / 50 mL (0.25 gram / mL) |  | 2762620-12.5
654 | mannitol 25% | --- gram / --- mL |  | 2762620-12.5
655 | mannitol 5% | 50 gram / 1,000 mL (0.05 gram / mL) |  | 2762620-5
656 | mannitol 5% | --- gram / --- mL |  | 2762620-5
657 | mechlorethamine | --- mg / --- mL |  | 2762671I
658 | melphalan | --- mg / --- mL |  | 2762749I
659 | meperidine | 10 mg / 1 mL (10 mg / mL) |  | 2762823I
660 | meperidine | 500 mg / 50 mL (10 mg / mL) |  | 2762823C
661 | meperidine | --- mg / --- mL |  | 2762823C
662 | meropenem | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2762894I
663 | meropenem | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2762894I
664 | meropenem | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2762894I
665 | meropenem | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2762894I
666 | meropenem | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2762894I
667 | meropenem | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2762894I
668 | meropenem | --- mg / --- mL | Intermittent | 2762894I
669 | meropenem | 0.5 gram / 50 mL (0.01 gram / mL) | Intermittent | 2762894I
670 | meropenem | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2762894I
671 | meropenem | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2762894I
672 | meropenem | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2762894I
673 | meropenem | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2762894I
674 | meropenem | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2762894I
675 | meropenem | --- gram / --- mL | Intermittent | 2762894I
676 | meropenem/vaborbacta | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2670044615I
677 | meropenem/vaborbacta | 1 gram / 250 mL (0.004 gram / mL) | Intermittent | 2670044615I
678 | meropenem/vaborbacta | 2 gram / 250 mL (0.008 gram / mL) | Intermittent | 2670044615I
679 | meropenem/vaborbacta | 2 gram / 500 mL (0.004 gram / mL) | Intermittent | 2670044615I
680 | meropenem/vaborbacta | 4 gram / 250 mL (0.016 gram / mL) | Intermittent | 2670044615I
681 | meropenem/vaborbacta | 4 gram / 500 mL (0.008 gram / mL) | Intermittent | 2670044615I
682 | meropenem/vaborbacta | 4 gram / 1,000 mL (0.004 gram / mL) | Intermittent | 2670044615I
683 | meropenem/vaborbacta | --- gram / --- mL | Intermittent | 2670044615I
684 | mesna | --- mg / --- mL | Intermittent | 2762918I
685 | methadone | --- mg / --- mL |  | 2763004C
686 | methocarbamol | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 2763080I
687 | methocarbamol | --- mg / --- mL | Intermittent | 2763080I
688 | methohexitol | --- mg / --- mL |  | 2763092I
689 | methotrexate | --- mg / --- mL |  | 2763101I
690 | methotrexate | --- gram / --- mL |  | 2763101I
691 | methyldopa | --- mg / --- mL |  | 2763256I
692 | methylPREDNISolone | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2763323I
693 | methylPREDNISolone | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2763323I
694 | methylPREDNISolone | --- mg / --- mL | Intermittent | 2763323I
695 | metoclopramide | --- mg / --- mL |  | 2763400I
696 | metoprolol | --- mg / --- mL |  | 2763424C,2763424I
697 | metroNIDAZOLE | 250 mg / 50 mL (5 mg / mL) | Intermittent | 2763443I
698 | metroNIDAZOLE | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2763443I
699 | metroNIDAZOLE | --- mg / --- mL | Intermittent | 2763443I
700 | mezlocillin | --- mg / --- mL |  | unknown
701 | micafungin | 50 mg / 100 mL (0.5 mg / mL) | Intermittent | 3830619I
702 | micafungin | 100 mg / 100 mL (1 mg / mL) | Intermittent | 3830619I
703 | micafungin | 150 mg / 100 mL (1.5 mg / mL) | Intermittent | 3830619I
704 | micafungin | --- mg / --- mL | Intermittent | 3830619I
705 | miconazole | --- mg / --- mL |  | 19953615I
706 | midazolam | 50 mg / 50 mL (1 mg / mL) | Cont/Bolus or PCA | 2763568C
707 | midazolam | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2763568C
708 | midazolam | 125 mg / 125 mL (1 mg / mL) | Cont/Bolus or PCA | 2763568C
709 | midazolam | --- mg / --- mL | Cont/Bolus or PCA | 2763568C
710 | milrinone | 20 mg / 100 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2763605C
711 | milrinone | 40 mg / 200 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2763605C
712 | milrinone | --- mg / --- mL | Cont/Bolus or PCA | 2763605C
713 | minocycline | --- mg / --- mL |  | 2763628I
714 | mitoMYcin | --- mg / --- mL | Intermittent | 2763694I
715 | mitoXANTRONE | --- mg / --- mL |  | 2763708I
716 | morphine | 1 mg / 1 mL (1 mg / mL) |  | 2763791I
717 | morphine | 30 mg / 30 mL (1 mg / mL) |  | 2763791C,2763791I
718 | morphine | 50 mg / 50 mL (1 mg / mL) |  | 2763791C
719 | morphine | 100 mg / 100 mL (1 mg / mL) |  | 2763791C
720 | morphine | --- mg / --- mL |  | 2763791C
721 | moxifloxacin | 400 mg / 250 mL (1.6 mg / mL) |  | 2763901I
722 | moxifloxacin | --- mg / --- mL |  | 2763901I
723 | muromonab | --- mg / --- mL |  | 2764672I
724 | mycophenolate | --- mg / --- mL |  | 3320344I
725 | nafcillin | 500 mg / 50 mL (10 mg / mL) | Intermittent | 2764713I
726 | nafcillin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2764713I
727 | nafcillin | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2764713I
728 | nafcillin | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2764713I
729 | nafcillin | 2,000 mg / 50 mL (40 mg / mL) | Intermittent | 2764713I
730 | nafcillin | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2764713I
731 | nafcillin | --- mg / --- mL | Intermittent | 2764713I
732 | nafcillin | 0.5 gram / 50 mL (0.01 gram / mL) | Intermittent | 2764713I
733 | nafcillin | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2764713I
734 | nafcillin | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2764713I
735 | nafcillin | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2764713I
736 | nafcillin | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2764713I
737 | nafcillin | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2764713I
738 | nafcillin | --- gram / --- mL | Intermittent | 2764713I
739 | nalbuphine | --- mg / --- mL |  | 2764738I
740 | naloxone | 2 mg / 500 mL (0.004 mg / mL) | Cont/Bolus or PCA | 2764760C
741 | naloxone | --- mg / --- mL | Cont/Bolus or PCA | 2764760C
742 | nelarabine | --- mg / --- mL |  | 3976209I
743 | nesiritide | 1.5 mg / 250 mL (0.006 mg / mL) | Cont/Bolus or PCA | 2764982C
744 | niCARdipine | 20 mg / 200 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2765030C
745 | niCARdipine | 25 mg / 250 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2765030C
746 | niCARdipine | 40 mg / 200 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2765030C
747 | niCARdipine | 50 mg / 500 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2765030C
748 | niCARdipine | --- mg / --- mL | Cont/Bolus or PCA | 2765030C
749 | nitroglycerin | 25 mg / 250 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2765191C
750 | nitroglycerin | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2765191C
751 | nitroglycerin | --- mg / --- mL | Cont/Bolus or PCA | 2765191C
752 | nitroprusside | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2765267C
753 | nitroprusside | 50 mg / 500 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2765267C
754 | nitroprusside | 100 mg / 250 mL (0.4 mg / mL) | Cont/Bolus or PCA | 2765267C
755 | nitroprusside | --- mg / --- mL | Cont/Bolus or PCA | 2765267C
756 | nivolumab | 240 mg / 100 mL (2.4 mg / mL) | Intermittent | 202016817I
757 | nivolumab | 480 mg / 100 mL (4.8 mg / mL) | Intermittent | 202016817I
758 | nivolumab | --- mg / --- mL | Intermittent | 202016817I
759 | norepinephrine | 4 mg / 250 mL (0.016 mg / mL) | Cont/Bolus or PCA | 2765336C
760 | norepinephrine | 8 mg / 250 mL (0.032 mg / mL) | Cont/Bolus or PCA | 2765336C
761 | norepinephrine | 8 mg / 500 mL (0.016 mg / mL) | Cont/Bolus or PCA | 2765336C
762 | norepinephrine | 16 mg / 250 mL (0.064 mg / mL) | Cont/Bolus or PCA | 2765336C
763 | norepinephrine | 16 mg / 500 mL (0.032 mg / mL) | Cont/Bolus or PCA | 2765336C
764 | norepinephrine | --- mg / --- mL | Cont/Bolus or PCA | 2765336C
765 | obinutuzumab | 100 mg / 100 mL (1 mg / mL) | Intermittent | 25920903I
766 | obinutuzumab | 100 mg / 104 mL (0.9615 mg / mL) | Intermittent | 25920903I
767 | obinutuzumab | 900 mg / 250 mL (3.6 mg / mL) | Intermittent | 25920903I
768 | obinutuzumab | 900 mg / 286 mL (3.1469 mg / mL) | Intermittent | 25920903I
769 | obinutuzumab | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 25920903I
770 | obinutuzumab | 1,000 mg / 290 mL (3.4483 mg / mL) | Intermittent | 25920903I
771 | obinutuzumab | --- mg / --- mL | Intermittent | 25920903I
772 | ocrelizumab | 300 mg / 250 mL (1.2 mg / mL) | Intermittent | 2581289359I
773 | ocrelizumab | 600 mg / 500 mL (1.2 mg / mL) | Intermittent | 2581289359I
774 | octreotide | 500 mcg / 100 mL (5 mcg / mL) | Cont/Bolus or PCA | 2765439C
775 | octreotide | 600 mcg / 240 mL (2.5 mcg / mL) | Cont/Bolus or PCA | 2765439C
776 | octreotide | 1,000 mcg / 250 mL (4 mcg / mL) | Cont/Bolus or PCA | 2765439C
777 | octreotide | 1,250 mcg / 250 mL (5 mcg / mL) | Cont/Bolus or PCA | 2765439C
778 | octreotide | --- mcg / --- mL | Cont/Bolus or PCA | 2765439C
779 | ofatumumab | --- mg / --- mL |  | 19954113I
780 | ondansetron | 4 mg / 50 mL (0.08 mg / mL) | Intermittent | 2765642I
781 | ondansetron | 8 mg / 50 mL (0.16 mg / mL) | Intermittent | 2765642I
782 | ondansetron | 8 mg / 54 mL (0.1481 mg / mL) | Intermittent | 2765642I
783 | ondansetron | 32 mg / 50 mL (0.64 mg / mL) | Intermittent | 2765642I
784 | ondansetron | --- mg / --- mL | Intermittent | 2765642I
785 | oritavacin | 1,200 mg / 1,000 mL (1.2 mg / mL) |  | 155136453I
786 | oxacillin | 1,000 mg / 50 mL (20 mg / mL) | Intermittent | 2765734I
787 | oxacillin | 2,000 mg / 100 mL (20 mg / mL) | Intermittent | 2765734I
788 | oxacillin | --- mg / --- mL | Intermittent | 2765734I
789 | oxacillin | 1 gram / 50 mL (0.02 gram / mL) | Intermittent | 2765734I
790 | oxacillin | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2765734I
791 | oxaliplatin | --- mg / --- mL | Intermittent | 2765754I
792 | oxytocin | 10 unit / 1,000 mL (0.01 unit / mL) | Cont/Bolus or PCA | 2765930C
793 | oxytocin | 20 unit / 1,000 mL (0.02 unit / mL) | Cont/Bolus or PCA | 2765930C
794 | oxytocin | 30 unit / 500 mL (0.06 unit / mL) | Cont/Bolus or PCA | 2765930C
795 | oxytocin | 30 unit / 1,000 mL (0.03 unit / mL) | Cont/Bolus or PCA | 2765930C
796 | oxytocin | 40 unit / 1,000 mL (0.04 unit / mL) | Cont/Bolus or PCA | 2765930C
797 | oxytocin | --- unit / --- mL | Cont/Bolus or PCA | 2765930C
798 | PACLitaxel | --- mg / --- mL | Intermittent | 2765935I
799 | PACLitaxel protein-b | --- mg / --- mL | Intermittent | 3830361I
800 | pamidronate | 90 mg / 250 mL (0.36 mg / mL) | Intermittent | 2765957I
801 | pamidronate | --- mg / --- mL | Intermittent | 2765957I
802 | pancuronium | --- mg / --- mL | Cont/Bolus or PCA | 2766061C
803 | panitumumab | --- mg / --- mL | Intermittent | 4246362I
804 | pantoprazole | 40 mg / 100 mL (0.4 mg / mL) |  | 2766068
805 | pantoprazole | 80 mg / 100 mL (0.8 mg / mL) |  | 2766068
806 | pantoprazole | 96 mg / 240 mL (0.4 mg / mL) |  | 2766068
807 | pantoprazole | --- mg / --- mL |  | 2766068
808 | pantoprazole gtt | 40 mg / 100 mL (0.4 mg / mL) | Cont/Bolus or PCA | 2766068C
809 | pantoprazole gtt | 96 mg / 240 mL (0.4 mg / mL) | Cont/Bolus or PCA | 2766068C
810 | pantoprazole gtt | --- mg / --- mL | Cont/Bolus or PCA | 2766068C
811 | pantoprazole ivpb | 40 mg / 100 mL (0.4 mg / mL) | Intermittent | 2766068I
812 | pantoprazole ivpb | 80 mg / 100 mL (0.8 mg / mL) | Intermittent | 2766068I
813 | papaverine | 300 mg / 300 mL (1 mg / mL) |  | 2766098I
814 | papaverine | --- mg / --- mL |  | 2766098I
815 | pegaspargase | --- unit / --- mL |  | 2766160I
816 | pembrolizumab | 200 mg / 100 mL (2 mg / mL) | Intermittent | 114589861I
817 | pembrolizumab | --- mg / --- mL | Intermittent | 114589861I
818 | PEMEtrexed | --- mg / --- mL | Intermittent | 3320387I
819 | penG 2.5 Mill Units | 2.5 unit / 50 mL (0.05 unit / mL) | Intermittent | 19955177I2.5
820 | penG 2.5 Mill Units | 2.5 unit / 100 mL (0.025 unit / mL) | Intermittent | 19955177I2.5
821 | penG 3 Million Units | 3 unit / 50 mL (0.06 unit / mL) | Intermittent | 19955177I
822 | penG 5 Million Units | 5 unit / 50 mL (0.1 unit / mL) | Intermittent | 19955177
823 | penG 5 Million Units | 5 unit / 100 mL (0.05 unit / mL) | Intermittent | 19955177
824 | penG x Million Units | --- unit / --- mL | Intermittent | 19955177-x
825 | pentamidine | --- mg / --- mL | Intermittent | 2766284I
826 | PENTobarbital | 2,500 mg / 500 mL (5 mg / mL) |  | 2766302I
827 | PENTobarbital | --- mg / --- mL |  | 2766302I
828 | pentostatin | --- mg / --- mL |  | 2766314I
829 | pertuzumab | 240 mg / 250 mL (0.96 mg / mL) | Intermittent | 24607638I
830 | pertuzumab | 420 mg / 250 mL (1.68 mg / mL) | Intermittent | 24607638I
831 | pertuzumab | 420 mg / 264 mL (1.5909 mg / mL) | Intermittent | 24607638I
832 | pertuzumab | 840 mg / 250 mL (3.36 mg / mL) | Intermittent | 24607638I
833 | pertuzumab | 840 mg / 278 mL (3.0216 mg / mL) | Intermittent | 24607638I
834 | PHENobarbital | 130 mg / 52 mL (2.5 mg / mL) | Intermittent | 2766460I
835 | PHENobarbital | 130 mg / 102 mL (1.2745 mg / mL) | Intermittent | 2766460I
836 | PHENobarbital | 260 mg / 54 mL (4.8148 mg / mL) | Intermittent | 2766460I
837 | PHENobarbital | 260 mg / 104 mL (2.5 mg / mL) | Intermittent | 2766460I
838 | PHENobarbital | --- mg / --- mL | Intermittent | 2766460I
839 | phentolamine | --- mg / --- mL |  | 2766536I
840 | phenylephrine | 10 mg / 250 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2766541C
841 | phenylephrine | 10 mg / 251 mL (0.0398 mg / mL) | Cont/Bolus or PCA | 2766541C
842 | phenylephrine | 20 mg / 250 mL (0.08 mg / mL) | Cont/Bolus or PCA | 2766541C
843 | phenylephrine | 20 mg / 500 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2766541C
844 | phenylephrine | 20 mg / 502 mL (0.0398 mg / mL) | Cont/Bolus or PCA | 2766541C
845 | phenylephrine | 40 mg / 250 mL (0.16 mg / mL) | Cont/Bolus or PCA | 2766541C
846 | phenylephrine | 40 mg / 1,000 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2766541C
847 | phenylephrine | --- mg / --- mL | Cont/Bolus or PCA | 2766541C
848 | phenylephrine ANES | 10 mg / 100 mL (0.1 mg / mL) | Cont/Bolus or PCA | 2766541
849 | phenylephrine ANES | 10 mg / 101 mL (0.099 mg / mL) | Cont/Bolus or PCA | 2766541
850 | phenylephrine ANES | 10 mg / 250 mL (0.04 mg / mL) | Cont/Bolus or PCA | 2766541
851 | phenylephrine ANES | 10 mg / 251 mL (0.0398 mg / mL) | Cont/Bolus or PCA | 2766541
852 | phenylephrine ANES | --- mg / --- mL | Cont/Bolus or PCA | 2766541
853 | phenytoin | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 2766638I
854 | phenytoin | --- mg / --- mL | Intermittent | 2766638I
855 | phytonadione | --- mg / --- mL | Intermittent | 2766694I
856 | piperacillin | 2.25 gram / 50 mL (0.045 gram / mL) |  | 2766817I
857 | piperacillin | 3.37 gram / 50 mL (0.0674 gram / mL) |  | 2766817I
858 | piperacillin | 4.5 gram / 50 mL (0.09 gram / mL) |  | 2766817I
859 | piperacillin/tazo | --- mg / --- mL | Intermittent | 2766819I
860 | piperacillin/tazo | 2.25 gram / 50 mL (0.045 gram / mL) | Intermittent | 2766819I
861 | piperacillin/tazo | 2.25 gram / 100 mL (0.0225 gram / mL) | Intermittent | 2766819I
862 | piperacillin/tazo | 3.375 gram / 50 mL (0.0675 gram / mL) | Intermittent | 2766819I
863 | piperacillin/tazo | 3.375 gram / 100 mL (0.0338 gram / mL) | Intermittent | 2766819I
864 | piperacillin/tazo | 4.5 gram / 50 mL (0.09 gram / mL) | Intermittent | 2766819I
865 | piperacillin/tazo | 4.5 gram / 100 mL (0.045 gram / mL) | Intermittent | 2766819I
866 | piperacillin/tazo | --- gram / --- mL | Intermittent | 2766819I
867 | plicamycin | --- mcg / --- mL |  | 2766901I
868 | porfirmer | --- mg / --- mL |  | 2767007I
869 | potassium acetate | 10 mEq / 100 mL (0.1 mEq / mL) | Intermittent | 2767012I
870 | potassium acetate | 20 mEq / 100 mL (0.2 mEq / mL) | Intermittent | 2767012I
871 | potassium acetate | 30 mEq / 100 mL (0.3 mEq / mL) | Intermittent | 2767012I
872 | potassium acetate | 40 mEq / 100 mL (0.4 mEq / mL) | Intermittent | 2767012I
873 | potassium acetate | 40 mEq / 520 mL (0.0769 mEq / mL) | Intermittent | 2767012I
874 | potassium acetate | --- mEq / --- mL | Intermittent | 2767012I
875 | potassium chloride | 10 mEq / 50 mL (0.2 mEq / mL) | Intermittent | 1000000100I
876 | potassium chloride | 10 mEq / 100 mL (0.1 mEq / mL) | Intermittent | 1000000100I
877 | potassium chloride | 20 mEq / 100 mL (0.2 mEq / mL) | Intermittent | 1000000100I
878 | potassium chloride | 40 mEq / 100 mL (0.4 mEq / mL) | Intermittent | 1000000100I
879 | potassium chloride | 40 mEq / 520 mL (0.0769 mEq / mL) | Intermittent | 1000000100I
880 | potassium chloride | --- mEq / --- mL | Intermittent | 1000000100I
881 | potassium phosphate | --- mEq / --- mL | Intermittent | 2767185I
882 | potassium phosphate | --- mmol / --- mL | Intermittent | 2767185I
883 | pralidoxime | --- mg / --- mL | Intermittent | 2767273I
884 | pralidoxime | --- gram / --- mL | Intermittent | 2767273I
885 | procainamide | 1,000 mg / 50 mL (20 mg / mL) | Cont/Bolus or PCA | 2767500C
886 | procainamide | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA | 2767500C
887 | procainamide | 2,000 mg / 250 mL (8 mg / mL) | Cont/Bolus or PCA | 2767500C
888 | procainamide | 2,000 mg / 500 mL (4 mg / mL) | Cont/Bolus or PCA | 2767500C
889 | prochlorperazine | --- mg / --- mL |  | 2767546I
890 | promethazine | --- mg / --- mL | Intermittent | 2767590I
891 | propofol | 500 mg / 50 mL (10 mg / mL) | Cont/Bolus or PCA | 2767660C
892 | propofol | 1,000 mg / 100 mL (10 mg / mL) | Cont/Bolus or PCA | 2767660C
893 | propofol | --- mg / --- mL | Cont/Bolus or PCA | 2767660C
894 | propofol ANES | 10 mg / 1 mL (10 mg / mL) | Cont/Bolus or PCA | 2767660I
895 | propofol ANES | 500 mg / 50 mL (10 mg / mL) | Cont/Bolus or PCA | 2767660I
896 | propofol ANES | 1,000 mg / 100 mL (10 mg / mL) | Cont/Bolus or PCA | 2767660I
897 | propofol ANES | --- mg / --- mL | Cont/Bolus or PCA | 2767660I
898 | protamine sulfate | --- mg / --- mL |  | 2767711I
899 | quiNIDine | --- mg / --- mL |  | 2767995I
900 | ramucirumab | --- mg / --- mL | Intermittent | 38821169I
901 | raNITIdine | 50 mg / 50 mL (1 mg / mL) | Intermittent | 2768055I
902 | raNITIdine | 50 mg / 100 mL (0.5 mg / mL) | Intermittent | 2768055I
903 | raNITIdine | 50 mg / 102 mL (0.4902 mg / mL) | Intermittent | 2768055I
904 | raNITIdine | 150 mg / 250 mL (0.6 mg / mL) | Intermittent | 2768055I
905 | raNITIdine | 250 mg / 250 mL (1 mg / mL) | Intermittent | 2768055I
906 | raNITIdine | --- mg / --- mL | Intermittent | 2768055I
907 | rasburicase | --- mg / --- mL |  | 2768080I
908 | remdesivir | 100 mg / 230 mL (0.4348 mg / mL) | Intermittent | 3520886709I
909 | remdesivir | 100 mg / 250 mL (0.4 mg / mL) | Intermittent | 3520886709I
910 | remdesivir | 200 mg / 210 mL (0.9524 mg / mL) | Intermittent | 3520886709I
911 | remdesivir | 200 mg / 250 mL (0.8 mg / mL) | Intermittent | 3520886709I
912 | remifentanil | 50 mcg / 1 mL (50 mcg / mL) |  | 2768094
913 | remifentanil | --- mcg / --- mL |  | 2768094
914 | remifentanil ANES | 50 mcg / 1 mL (50 mcg / mL) | Cont/Bolus or PCA | 2768094C
915 | remifentanil ANES | 1,000 mcg / 40 mL (25 mcg / mL) | Cont/Bolus or PCA | 2768094C
916 | remifentanil ANES | 2,000 mcg / 80 mL (25 mcg / mL) | Cont/Bolus or PCA | 2768094C
917 | remifentanil ANES | --- mcg / --- mL | Cont/Bolus or PCA | 2768094C
918 | reteplase | --- unit / --- mL |  | 2768131I
919 | rifAMPin | 300 mg / 250 mL (1.2 mg / mL) | Intermittent | 2768179I
920 | rifAMPin | 600 mg / 500 mL (1.2 mg / mL) | Intermittent | 2768179I
921 | rifAMPin | --- mg / --- mL | Intermittent | 2768179I
922 | ritodrine | --- mg / --- mL |  | 2768244I
923 | riTUXimab | 500 mg / 250 mL (2 mg / mL) | Intermittent | 2768259I
924 | riTUXimab | 1,000 mg / 500 mL (2 mg / mL) | Intermittent | 2768259I
925 | riTUXimab | --- mg / --- mL | Intermittent | 2768259I
926 | rocuronium | 500 mg / 500 mL (1 mg / mL) | Cont/Bolus or PCA | 2768289C
927 | rocuronium | --- mg / --- mL | Cont/Bolus or PCA | 2768289C
928 | rocuronium ANES | 500 mg / 500 mL (1 mg / mL) | Cont/Bolus or PCA | 2768289I
929 | rocuronium ANES | --- mg / --- mL | Cont/Bolus or PCA | 2768289I
930 | romiDEPsin | --- mg / --- mL |  | 19955002I
931 | ropivacaine | 200 mg / 100 mL (2 mg / mL) |  | 2768324C
932 | ropivacaine | 400 mg / 200 mL (2 mg / mL) |  | 2768324C
933 | ropivacaine | --- mg / --- mL |  | 2768324C
934 | salbutamol | --- mg / --- mL |  | salbutamol
935 | sargramostim | --- mcg / --- mL |  | 2768593I
936 | siltuximab | --- mg / --- mL |  | 45868529I
937 | SMX/TMP | --- mg / --- mL | Intermittent | 2769391I
938 | sodium bicarbonate | --- mEq / --- mL |  | 2768888
939 | sodium chloride 3% | 256.6 mEq / 500 mL (0.5132 mEq / mL) |  | 2778727
940 | sodium chloride 5% | 427.5 mEq / 500 mL (0.855 mEq / mL) |  | 2778729
941 | sodium phosphate | --- mmol / --- mL | Intermittent | 2769071I
942 | sotrovimab | 500 mg / 50 mL (10 mg / mL) | Intermittent | 4112134745I
943 | sotrovimab | 500 mg / 100 mL (5 mg / mL) | Intermittent | 4112134745I
944 | streptozocin | --- mg / --- mL |  | 2769343I
945 | succinylcholine | --- mg / --- mL |  | 2769352I
946 | SUFentanil | --- mcg / --- mL |  | 2769370I
947 | SUFentanil ANES | 100 mcg / 100 mL (1 mcg / mL) | Cont/Bolus or PCA | 2769370C
948 | SUFentanil ANES | --- mcg / --- mL | Cont/Bolus or PCA | 2769370C
949 | tacrolimus | --- mg / --- mL |  | 2769525I
950 | tedizolid phosphate | 200 mg / 250 mL (0.8 mg / mL) | Intermittent | 45868623I
951 | temozolomide | --- mg / --- mL | Intermittent | 2769599I
952 | temsirolimus | --- mg / --- mL |  | 4246596I
953 | teniposide | --- mg / --- mL |  | 2769615I
954 | teprotumumab | --- mg / --- mL | Intermittent | 3602321791I
955 | teprotumumab | --- gram / --- mL | Intermittent | 3602321791I
956 | terbutaline | --- mg / --- mL |  | 2769660I
957 | theophylline | 400 mg / 500 mL (0.8 mg / mL) |  | 2769843I
958 | theophylline | 800 mg / 500 mL (1.6 mg / mL) |  | 2769843I
959 | theophylline | --- mg / --- mL |  | 2769843I
960 | thiamine | 100 mg / 50 mL (2 mg / mL) | Intermittent | 2769929I
961 | thiamine | 100 mg / 100 mL (1 mg / mL) | Intermittent | 2769929I
962 | thiamine | --- mg / --- mL | Intermittent | 2769929I
963 | thiopental | --- mg / --- mL |  | 2769954I
964 | thiotepa | --- mg / --- mL |  | 2769992I
965 | ticarcillin | 3.1 gram / 100 mL (0.031 gram / mL) |  | 2770096I
966 | ticarcillin/clav | --- mg / --- mL | Intermittent | 2770103I
967 | ticarcillin/clav | 2 gram / 50 mL (0.04 gram / mL) | Intermittent | 2770103I
968 | ticarcillin/clav | 2 gram / 100 mL (0.02 gram / mL) | Intermittent | 2770103I
969 | ticarcillin/clav | 3.1 gram / 50 mL (0.062 gram / mL) | Intermittent | 2770103I
970 | ticarcillin/clav | 3.1 gram / 100 mL (0.031 gram / mL) | Intermittent | 2770103I
971 | tigecycline | 50 mg / 50 mL (1 mg / mL) | Intermittent | 3831121I
972 | tigecycline | 50 mg / 100 mL (0.5 mg / mL) | Intermittent | 3831121I
973 | tigecycline | 100 mg / 100 mL (1 mg / mL) | Intermittent | 3831121I
974 | tirofiban | 12.5 mg / 250 mL (0.05 mg / mL) | Cont/Bolus or PCA | 2770172C
975 | tobramycin | 80 mg / 50 mL (1.6 mg / mL) | Intermittent | 2770186I
976 | tobramycin | 80 mg / 100 mL (0.8 mg / mL) | Intermittent | 2770186I
977 | tobramycin | 120 mg / 100 mL (1.2 mg / mL) | Intermittent | 2770186I
978 | tobramycin | --- mg / --- mL | Intermittent | 2770186I
979 | tocilizumab | 800 mg / 100 mL (8 mg / mL) | Intermittent | 19955605I
980 | tocilizumab | --- mg / --- mL | Intermittent | 19955605I
981 | tolazoline | --- mg / --- mL |  | tolazoline
982 | topotecan | --- mg / --- mL | Intermittent | 2770314I
983 | torsemide | --- mg / --- mL |  | 2770324I
984 | tranexamic BOLUS | 1,000 mg / 100 mL (10 mg / mL) | Intermittent | 2770412I
985 | tranexamic BOLUS | --- mg / --- mL | Intermittent | 2770412I
986 | tranexamic BOLUS | 1 gram / 100 mL (0.01 gram / mL) | Intermittent | 2770412I
987 | tranexamic BOLUS | --- gram / --- mL | Intermittent | 2770412I
988 | trastuzumab | --- mg / --- mL | Intermittent | 2770424I
989 | trimetrexate | --- mg / --- mL |  | 2770672I
990 | tromethamine | --- mEq / --- mL |  | 2770725I
991 | urokinase | --- unit / --- mL |  | 2770879I
992 | valproate sodium | --- mg / --- mL | Intermittent | 2770918I
993 | valproic acid | --- mg / --- mL |  | 2770918
994 | vancomycin | 500 mg / 100 mL (5 mg / mL) | Intermittent | 2770944I
995 | vancomycin | 750 mg / 150 mL (5 mg / mL) | Intermittent | 2770944I
996 | vancomycin | 750 mg / 250 mL (3 mg / mL) | Intermittent | 2770944I
997 | vancomycin | 1,000 mg / 200 mL (5 mg / mL) | Intermittent | 2770944I
998 | vancomycin | 1,000 mg / 250 mL (4 mg / mL) | Intermittent | 2770944I
999 | vancomycin | 1,250 mg / 250 mL (5 mg / mL) | Intermittent | 2770944I
1000 | vancomycin | 1,500 mg / 300 mL (5 mg / mL) | Intermittent | 2770944I
1001 | vancomycin | 1,500 mg / 500 mL (3 mg / mL) | Intermittent | 2770944I
1002 | vancomycin | 1,750 mg / 350 mL (5 mg / mL) | Intermittent | 2770944I
1003 | vancomycin | 1,750 mg / 500 mL (3.5 mg / mL) | Intermittent | 2770944I
1004 | vancomycin | 2,000 mg / 400 mL (5 mg / mL) | Intermittent | 2770944I
1005 | vancomycin | 2,000 mg / 500 mL (4 mg / mL) | Intermittent | 2770944I
1006 | vancomycin | --- mg / --- mL | Intermittent | 2770944I
1007 | vancomycin | 0.5 gram / 100 mL (0.005 gram / mL) | Intermittent | 2770944I
1008 | vancomycin | 0.75 gram / 150 mL (0.005 gram / mL) | Intermittent | 2770944I
1009 | vancomycin | 0.75 gram / 250 mL (0.003 gram / mL) | Intermittent | 2770944I
1010 | vancomycin | 1 gram / 200 mL (0.005 gram / mL) | Intermittent | 2770944I
1011 | vancomycin | 1 gram / 250 mL (0.004 gram / mL) | Intermittent | 2770944I
1012 | vancomycin | 1.25 gram / 250 mL (0.005 gram / mL) | Intermittent | 2770944I
1013 | vancomycin | 1.5 gram / 300 mL (0.005 gram / mL) | Intermittent | 2770944I
1014 | vancomycin | 1.5 gram / 500 mL (0.003 gram / mL) | Intermittent | 2770944I
1015 | vancomycin | 1.75 gram / 350 mL (0.005 gram / mL) | Intermittent | 2770944I
1016 | vancomycin | 1.75 gram / 500 mL (0.0035 gram / mL) | Intermittent | 2770944I
1017 | vancomycin | 2 gram / 400 mL (0.005 gram / mL) | Intermittent | 2770944I
1018 | vancomycin | 2 gram / 500 mL (0.004 gram / mL) | Intermittent | 2770944I
1019 | vancomycin | --- gram / --- mL | Intermittent | 2770944I
1020 | vasopressin | 0.02 unit / 1 mL (0.02 unit / mL) | Cont/Bolus or PCA | 2770987C
1021 | vasopressin | 0.04 unit / 1 mL (0.04 unit / mL) | Cont/Bolus or PCA | 2770987C
1022 | vasopressin | 1 unit / 1 mL (1 unit / mL) | Cont/Bolus or PCA | 2770987C
1023 | vasopressin | 4 unit / 100 mL (0.04 unit / mL) | Cont/Bolus or PCA | 2770987C
1024 | vasopressin | 20 unit / 100 mL (0.2 unit / mL) | Cont/Bolus or PCA | 2770987C
1025 | vasopressin | 40 unit / 100 mL (0.4 unit / mL) | Cont/Bolus or PCA | 2770987C
1026 | vasopressin | 80 unit / 100 mL (0.8 unit / mL) | Cont/Bolus or PCA | 2770987C
1027 | vasopressin | 100 unit / 100 mL (1 unit / mL) | Cont/Bolus or PCA | 2770987C
1028 | vasopressin | --- unit / --- mL | Cont/Bolus or PCA | 2770987C
1029 | vecuronium | 1 mg / 1 mL (1 mg / mL) | Cont/Bolus or PCA | 2770992C
1030 | vecuronium | 2 mg / 1 mL (2 mg / mL) | Cont/Bolus or PCA | 2770992C
1031 | vecuronium | 20 mg / 50 mL (0.4 mg / mL) | Cont/Bolus or PCA | 2770992C
1032 | vecuronium | 25 mg / 50 mL (0.5 mg / mL) | Cont/Bolus or PCA | 2770992C
1033 | vecuronium | 30 mg / 50 mL (0.6 mg / mL) | Cont/Bolus or PCA | 2770992C
1034 | vecuronium | 50 mg / 50 mL (1 mg / mL) | Cont/Bolus or PCA | 2770992C
1035 | vecuronium | 50 mg / 100 mL (0.5 mg / mL) | Cont/Bolus or PCA | 2770992C
1036 | vecuronium | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA | 2770992C
1037 | vecuronium | 100 mg / 100 mL (1 mg / mL) | Cont/Bolus or PCA | 2770992C
1038 | vecuronium | --- mg / --- mL | Cont/Bolus or PCA | 2770992C
1039 | verapamil | --- mg / --- mL |  | 2771023I
1040 | vidarabine | --- mg / --- mL |  | vidarabine
1041 | vinBLAStine | --- mg / --- mL | Intermittent | 2771078I
1042 | vinCRIStine | --- mg / --- mL | Intermittent | 2771083I
1043 | vinCRIStine Liposom | --- mg / --- mL |  | 24608273I
1044 | vinorelbine | --- mg / --- mL | Intermittent | 2771090I
1045 | voriconazole | 200 mg / 50 mL (4 mg / mL) | Intermittent | 2771180I
1046 | voriconazole | 400 mg / 100 mL (4 mg / mL) | Intermittent | 2771180I
1047 | voriconazole | --- mg / --- mL | Intermittent | 2771180I
1048 | zidovudine | --- mg / --- mL | Intermittent | 2771278I
1049 | ziv-aflibercept | --- mg / --- mL | Intermittent | 24608076I
1050 | zoledronic acid | 4 mg / 100 mL (0.04 mg / mL) | Intermittent | 2771383I
1051 | zoledronic acid | 5 mg / 100 mL (0.05 mg / mL) | Intermittent | 2771383I
1052 | zoledronic acid | --- mg / --- mL | Intermittent | 2771383I
1053 | zz amiodarone | 150 mg / 100 mL (1.5 mg / mL) | Cont/Bolus or PCA
1054 | zz amiodarone | 900 mg / 500 mL (1.8 mg / mL) | Cont/Bolus or PCA
1055 | zz DOBUTamine | 2,000 mcg / 1 mL (2,000 mcg / mL) | Cont/Bolus or PCA
1056 | zz DOBUTamine | 250 mg / 250 mL (1 mg / mL) | Cont/Bolus or PCA
1057 | zz DOBUTamine | 500 mg / 250 mL (2 mg / mL) | Cont/Bolus or PCA
1058 | zz DOBUTamine | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA
1059 | zz DOBUTamine | --- mg / --- mL | Cont/Bolus or PCA
1060 | zz DOPamine | 1,600 mcg / 1 mL (1,600 mcg / mL) | Cont/Bolus or PCA
1061 | zz DOPamine | 400 mg / 250 mL (1.6 mg / mL) | Cont/Bolus or PCA
1062 | zz DOPamine | 800 mg / 250 mL (3.2 mg / mL) | Cont/Bolus or PCA
1063 | zz DOPamine | --- mg / --- mL | Cont/Bolus or PCA
1064 | zz EPINEPHrine | 10 mcg / 1 mL (10 mcg / mL) | Cont/Bolus or PCA
1065 | zz EPINEPHrine | 4 mg / 250 mL (0.016 mg / mL) | Cont/Bolus or PCA
1066 | zz EPINEPHrine | 5 mg / 250 mL (0.02 mg / mL) | Cont/Bolus or PCA
1067 | zz EPINEPHrine | 8 mg / 250 mL (0.032 mg / mL) | Cont/Bolus or PCA
1068 | zz EPINEPHrine | 10 mg / 100 mL (0.1 mg / mL) | Cont/Bolus or PCA
1069 | zz EPINEPHrine | 10 mg / 500 mL (0.02 mg / mL) | Cont/Bolus or PCA
1070 | zz EPINEPHrine | 16 mg / 250 mL (0.064 mg / mL) | Cont/Bolus or PCA
1071 | zz lidocaine | 8 mg / 1 mL (8 mg / mL) | Cont/Bolus or PCA
1072 | zz lidocaine | 1,000 mg / 250 mL (4 mg / mL) | Cont/Bolus or PCA
1073 | zz lidocaine | 2,000 mg / 500 mL (4 mg / mL) | Cont/Bolus or PCA
1074 | zz lidocaine | --- mg / --- mL | Cont/Bolus or PCA
1075 | zz lidocaine | 2 gram / 500 mL (0.004 gram / mL) | Cont/Bolus or PCA
1076 | zz mag sulfate | --- mg / --- mL | Intermittent
1077 | zz mag sulfate | 1 gram / 50 mL (0.02 gram / mL) | Intermittent
1078 | zz mag sulfate | 1 gram / 100 mL (0.01 gram / mL) | Intermittent
1079 | zz mag sulfate | 2 gram / 50 mL (0.04 gram / mL) | Intermittent
1080 | zz mag sulfate | 4 gram / 100 mL (0.04 gram / mL) | Intermittent
1081 | zz mag sulfate | 40 gram / 500 mL (0.08 gram / mL) | Intermittent
1082 | zz mag sulfate | --- gram / --- mL | Intermittent
1083 | zz nitroglycerin | 25 mg / 250 mL (0.1 mg / mL) | Cont/Bolus or PCA
1084 | zz nitroglycerin | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA
1085 | zz nitroglycerin | --- mg / --- mL | Cont/Bolus or PCA
1086 | zz nitroprusside | 50 mg / 250 mL (0.2 mg / mL) | Cont/Bolus or PCA
1087 | zz nitroprusside | 100 mg / 250 mL (0.4 mg / mL) | Cont/Bolus or PCA
1088 | zz nitroprusside | --- mg / --- mL | Cont/Bolus or PCA
1089 | zz norepinephrine | 16 mcg / 1 mL (16 mcg / mL) | Cont/Bolus or PCA
1090 | zz norepinephrine | 4 mg / 250 mL (0.016 mg / mL) | Cont/Bolus or PCA
1091 | zz norepinephrine | 8 mg / 250 mL (0.032 mg / mL) | Cont/Bolus or PCA
1092 | zz norepinephrine | 8 mg / 500 mL (0.016 mg / mL) | Cont/Bolus or PCA
1093 | zz norepinephrine | 16 mg / 250 mL (0.064 mg / mL) | Cont/Bolus or PCA
1094 | zz norepinephrine | 16 mg / 500 mL (0.032 mg / mL) | Cont/Bolus or PCA
1095 | zz norepinephrine | --- mg / --- mL | Cont/Bolus or PCA
1096 | zz phenylephrine | 10 mg / 250 mL (0.04 mg / mL) | Cont/Bolus or PCA
1097 | zz phenylephrine | 20 mg / 250 mL (0.08 mg / mL) | Cont/Bolus or PCA
1098 | zz phenylephrine | 40 mg / 250 mL (0.16 mg / mL) | Cont/Bolus or PCA
1099 | zz phenylephrine | 50 mg / 100 mL (0.5 mg / mL) | Cont/Bolus or PCA
1100 | zz phenylephrine | --- mg / --- mL | Cont/Bolus or PCA

## Master Fluids
Master Fluids | Master Fluids | Master Fluids | Master Fluids
Line Item | Fluid Name | Supports Secondary | Aliases
1 | ..IVF Maintenance | Yes | 1000000000F,1000000000I
2 | .Bolus IVF | Yes
3 | .IVF + KCL | Yes | 1000000100F
4 | .IVF + KCL 10 mEq/L | Yes | 1000000010F
5 | .IVF + KCL 20 mEq/L | Yes | 1000000020F
6 | .IVF + KCL 30 mEq/L | Yes
7 | .IVF + KCL 40 mEq/L | Yes | 1000000040F
8 | .IVF + Oxytocin | Yes
9 | .IVF + Vitamins | Yes
10 | .Maint IVF + KCL | Yes
11 | .Maint IVF+Additives | Yes
12 | 0.45%NS+KCL10 mEq/L | Yes
13 | 0.45%NS+KCL20 mEq/L | Yes
14 | 0.45%NS+KCL30 mEq/L | Yes
15 | 0.45%NS+KCL40 mEq/L | Yes
16 | 0.9% Normal Saline | Yes
17 | 0.9%NS+KCL 10mEq/L | Yes
18 | 0.9%NS+KCL 20mEq/L | Yes
19 | 0.9%NS+KCL 30mEq/L | Yes
20 | 0.9%NS+KCL 40mEq/L | Yes
21 | 1/2NS Hep 0.5unit/mL | Yes
22 | 1/4NS Hep 0.5unit/mL | Yes
23 | 3% Sodium Chloride | Yes | 2778727F,2778727I
24 | Bicarb Drip | Yes | 2768888F,2768888I
25 | Blood (RBCs) | No
26 | Blood (whole) | No
27 | Blood Reinfusion | No
28 | Cryoprecipitate | No
29 | D10W 1/4NS | Yes
30 | D10W 1/4NS Hep 0.5un | Yes
31 | D10W Hep 0.5units/mL | Yes
32 | D5 + Normal Saline | Yes
33 | D5 1/2Normal Saline | Yes
34 | D5 1/2NS+KCL10mEq/L | Yes
35 | D5 1/2NS+KCL20mEq/L | Yes
36 | D5 1/2NS+KCL30mEq/L | Yes
37 | D5 1/2NS+KCL40mEq/L | Yes
38 | D5 1/3Normal Saline | Yes
39 | D5 1/3NS+KCL10mEq/L | Yes
40 | D5 1/3NS+KCL20mEq/L | Yes
41 | D5 1/3NS+KCL30mEq/L | Yes
42 | D5 1/3NS+KCL40mEq/L | Yes
43 | D5 1/4Normal Saline | Yes
44 | D5 1/4NS Hep 0.5unit | Yes
45 | D5 1/4NS+KCL10mEq/L | Yes
46 | D5 1/4NS+KCL20mEq/L | Yes
47 | D5 1/4NS+KCL30mEq/L | Yes
48 | D5 1/4NS+KCL40mEq/L | Yes
49 | D5 Lactated Ringers | Yes
50 | D5NS+KCL 10 mEq/L | Yes
51 | D5NS+KCL 20 mEq/L | Yes
52 | D5NS+KCL 30 mEq/L | Yes
53 | D5NS+KCL 40 mEq/L | Yes
54 | D5W | Yes
55 | D5W Hep 0.5units/mL | Yes
56 | D5W+KCL 10 mEq/L | Yes
57 | D5W+KCL 20 mEq/L | Yes
58 | D5W+KCL 30 mEq/L | Yes
59 | D5W+KCL 40 mEq/L | Yes
60 | Dextrose 10% | Yes | 2778755F,2778755I
61 | Dextrose in Ringers | Yes
62 | FLUSH | Yes
63 | Fresh Frozen Plasma | No
64 | Hetastarch 6% NS | Yes
65 | Infliximab(Remicade) | Yes
66 | Insulin Maternity | Yes | 2761029F
67 | IVF +Hep 0.5units/mL | Yes
68 | Lactated Ringers | Yes
69 | Lipids 20% | No
70 | Mannitol | No
71 | MVI Only | Yes | 1000000003F,1000000003I
72 | MVI+Folate+Thiamine | Yes | 1000000002F,1000000002I
73 | NS w Hep 0.5units/mL | Yes
74 | Parenteral Nutrition | No | 1000000001F,1000000001I
75 | Platelets | No
76 | PPN | No
77 | PRBCs | No
78 | Replacement Fluid | Yes
79 | Rituximab (Rituxan) | Yes
80 | TPN | No
81 | UAC | No
82 | UVC | No

## Master Clinical Advisories
Master Clinical Advisories | Master Clinical Advisories | Master Clinical Advisories
Line Item | Advisory Name | Advisory Text
1 | 0.2 to 5 micron filt | Use a low-protein binding in-line or add-on filter between 0.2 and 5 microns.
2 | 0.22 Filter Needed | Use a 0.22 micron in-line filter when administering this product.
3 | 1 unit = 1 mL | 1 unit = 1 mL
4 | 12 Hour Tubing | Tubing must be changed every 12 hours
5 | 1unit=1million | 1 unit = 1 million units
6 | 2nd RN to verify | 2 RNs to verify and document concentration and programming prior to infusion.
7 | 48 Hour Tubing | Change tubing every 48 hours.
8 | Alaris Unit | 1 Unit = 1 million ex. 0.6units = 600,000 units
9 | Allergic Reaction | Infuse slowly for the first 10 minutes (25-50 mL/hr) to watch for allergic reaction, if none, rate may be increased to 250 mL/hr.
10 | Alteplase wt based | Dose and volume to be infused need to be set according to weight based protocol to avoid potential over infusion of entire contents of bag.
11 | AmBisome | Existing IV lines must be flushed with D5w prior to infusion. If unable, must start a new line. ***Infusion requires the use of an in-line filter.
12 | Anaphylaxis Warning | Observe patient for 1 hour after starting infusion for signs and symptoms of anaphylaxis.
13 | Antivenin | 1 unit = 1 vial Infuse over 25-50 mL/hr for the first 10 minutes, if no allergic reaction, may increase rate.
14 | Bleeding Precautions | Monitor for signs of bleeding.
15 | Calcium Precipitate | Do not infuse Calcium and Phosphate though the same line.
16 | CeftriSyringe | Adminster over 60 minutes to decrease risk of bilirubin encephalopathy.
17 | Central Line Only | Run through a CENTRAL line only.
18 | Cerebyx | Not to be infused faster than 150 mgPE / min, due to risk of hypotension.
19 | Check Concentration | *CHECK CONCENTRATION* 2nd RN to verify and document concentration and programming prior to starting infusion.
20 | Code only drugs | This selection is reserved for use during Code Blue, ACLS events. DO NOT USE for routine infusions!
21 | Dedicated IV Line | Do not infuse any other drugs through this line.
22 | DEHP | This medication must be infused using DEHP FREE tubing.
23 | Double Concentrated | ****DOUBLE CONCENTRATED DRIP**** Double-check bag concentration and pump settings.
24 | EKG Required | EKG Montoring required for infusions greater than 10mEq per hour.
25 | Filter and DEHP Free | Requires Low protein binding 0.2 micron filter and DEHP free bag
26 | Fluid Bolus | Administer a fluid bolus prior to starting the cidofovir infusion.
27 | Flush Before/After | Flush line before and after administration. Can be given on same line as sodium chloride 0.9%.
28 | Heparin | Monitor PTT. Do not administer other drugs through same IV line.
29 | High Risk | This is a HIGH RISK medication. 2nd RN to verify prior to initiating therapy.
30 | High Risk Central | This is a HIGH RISK medication that should be infused via a CENTRAL LINE. 2nd RN to verify prior to initiating therapy.
31 | Hypersensitivity | Risk of hypersensitivity. Check vital signs frequently during infusion.
32 | IFEX | Dose is based on Ifosfamide (IFEX) content
33 | In-line filter | Infusion requires the use of an in-line filter.
34 | INR > 6 | INR > 6 ** Max dose of Kcentra should not exceed 5000 units **
35 | INR 2-4 | INR 2 - 4 ** Max dose of Kcentra should not exceed 2500 units **
36 | INR 4-6 | INR 4 - 6 ** Max dose of Kcentra should not exceed 3500 units **
37 | Insulin Maternity | 75 ml/hr = 0.75 u/hr 100 ml/hr = 1 u/hr 125 ml/hr = 1.25 u/hr 150 ml/hr = 1.5 u/hr
38 | Levophed | *****Must have physician order for rates greater than 60 mcg/min*****
39 | Low Binding 0.2or1.2 | This medication must be infused using a 0.2- to 1.2-micron, low protein brinding filter.
40 | Low Binding Filter | This medication must be infused using a 0.2- micron, low protein binding filter.
41 | Low Binding Tubing | * Use Low Binding Tubing *
42 | MAX DOSE 180 mg | MAXIMUM dose of brentuximab is 180 mg.
43 | Midline Warning | Do NOT infuse through a midline catheter.
44 | Monitoring required | Monitoring module required.
45 | MVI | 1 unit = 1 dose
46 | Neuromuscular BA | *Train of Four Required* Mechanical Ventilation and concurrent Sedation is mandatory.
47 | No Dextrose | Do NOT administer through lines containing dextrose.
48 | NO FILTER | Do NOT use an in-line filter for administration.
49 | Non-PVC | This medication must be infused using Non-PVC tubing.
50 | NS only | Do not dilute with D5W Dilute with NS 0.9% only
51 | oritavacin | Contraindicated to run ****Heparin**** within 48 hours of Oritavacin administration. May increase risk of bleeding with oral anticoagulant therapy.
52 | Patient Lie Down | Advise patient to lie down during and 30 minutes after infusion to prevent hypotension.
53 | Phenylephrine | *****Must have physician order for rates greater than 300 mcg/min*****
54 | Premedicate | Premedicate before each infusion.
55 | Protect From Light | Protect this drug from light.
56 | SCr 2.0-4.0 | SCr 2.0 - 4.0
57 | SCr<2.0 | SCr < 2.0
58 | Siltuximab | Use PVC with DEHP or PU containing administration set containing a 0.2 micron filter.
59 | Stress Test | For STRESS TEST only.
60 | Vesicant | Drug is a vesicant. Check blood return.
61 | Vesicant No Filter | Drug is a vesicant. Check blood return. Do not use an in-line filter.
62 | Vincristine Lipo | Use dedicated line. Do not filter.

## Master Therapies
Master Therapies | Master Therapies
Line Item | Therapy
1 | ..Less than 3 kg
2 | ..Standard
3 | ..Std. Cardiac/PVD
4 | .Double Concentrated
5 | .Induction
6 | .INR 2 to < 4
7 | .INR 4 - 6
8 | .MG Dosing
9 | .mg, pt weight <40kg
10 | .mg, pt weight>=40kg
11 | .Standard & wt>121kg
12 | .Standard Dose
13 | .Standard NON STEMI
14 | .STANDARD STEMI PCI
15 | .Standard wt >=50 kg
16 | .Std Cardiac or PVD
17 | .Test Dose
18 | .VTE (DVT, PE)
19 | .VTE/DVT/PE
20 | .wt less than 100kg
21 | 0.2 mg / 1 mL
22 | 1 mg/1mL
23 | 1 mg/mL *HIGH CONC*
24 | 10 mEq/L
25 | 15 mEq/L
26 | 1st Dose <=100kg
27 | 1st Dose >100kg
28 | 20 mEq/L
29 | 2nd Dose <=100kg
30 | 2nd Dose >100kg
31 | 3 - 9.9 kg
32 | 30 mEq/L
33 | 3rd Dose <=100kg
34 | 3rd Dose >100kg
35 | 40 mEq/L
36 | 5 mg/1mL *HIGH CONC*
37 | Anavip
38 | Anesthesia
39 | Anticonvulsant Thera
40 | AUC 2-3 q week
41 | AUC 3-7.5 q3weeks
42 | Based on AUC
43 | Based on BSA
44 | Bleeding
45 | Bleeding/Hemophilia
46 | BSA Based Dose
47 | Cardiac Arrest
48 | CENTRAL < 20kg
49 | CENTRAL >=20kg
50 | Central IV Occlusion
51 | Central Line
52 | Cerebrovascular
53 | CrCl <= 60
54 | CrCl > 60
55 | Crofab
56 | Cycle 1
57 | Cycle 2
58 | Cyclic
59 | Diabetes Insipidus
60 | DVT / Ischemic Limb
61 | Echo Stress Test
62 | ESRD
63 | Every 2 weeks
64 | Every 3 weeks
65 | First Dose
66 | GI Bleed
67 | GM dosing
68 | gram, pt wt < 40kg
69 | gram, pt wt >= 40kg
70 | Hepatic Dose
71 | High Dose
72 | High Dose - GRAMS
73 | Induction
74 | INR > 6
75 | Insulin Maternity
76 | Load <67kg WtBasedGM
77 | Load <67kg WtBasedMG
78 | Load >67kg WtBasedGM
79 | Load >67kg WtBasedMG
80 | Load NonWtBased MG
81 | Load NonWtBasedGM
82 | Loading Dose
83 | LoDose & >= 153 kg
84 | LoDose & CrCl < 30
85 | LoDose & CrCl >= 30
86 | LoDose,Renal&>=153kg
87 | Maintenance Dose
88 | Maintenance Dose GM
89 | Maintenance Dose MG
90 | Non-Anesthesia
91 | Non-Based Based Dose
92 | OB/L&D
93 | PERIPH < 20kg
94 | PERIPH >=20kg
95 | Peripheral Line
96 | Postpartum
97 | pt age < = 12 years
98 | pt age > = 12 years
99 | pt BSA < 1 m2
100 | pt BSA >=1
101 | pt weight < 20 kg
102 | pt weight < 30 kg
103 | pt weight < 40 kg
104 | pt weight < 50 kg
105 | pt weight < 60 kg
106 | pt weight >= 20 kg
107 | pt weight >= 30 kg
108 | pt weight >= 40 kg
109 | pt weight >= 50 kg
110 | pt weight >= 60 kg
111 | Pulmonary Embolism
112 | Renal Dose
113 | Renal Dose CrCl < 30
114 | Renal Dose wt >121kg
115 | Renal NON STEMI
116 | Renal STEMI PCI
117 | Rheumatoid Arthritis
118 | Second Dose
119 | Shock
120 | Spinal bolus
121 | Stress Test
122 | Stroke < 101 kg
123 | Stroke 101kg or more
124 | Subsequent Doses
125 | Third Dose
126 | Undergoing PCI
127 | Weekly
128 | Weight < 50 kg
129 | Weight above 110kg
130 | Weight Based
131 | Weight over 80kg
132 | wt < or = 1.5 kg
133 | WT = 10 - 19.9 kg
134 | WT = 20 - 39.9 kg
135 | WT = 40 kg and over
136 | wt 1.5 - 2.5 kg
137 | wt above 2.5 kg
138 | wt more than 100kg

## Channel Labels
Channel Labels | Channel Labels | Channel Labels | Channel Labels | Channel Labels
Line Item | Channel Labels are not used | Channel Labels are not used | Channel Labels are not used | Channel Labels are not used

## Master Syringe List
Master Syringe List | Master Syringe List | Master Syringe List
Line Item | Syringe Type | Included?
1 | BD Plastipak 1 mL | False
2 | BD Plastipak 3 mL | True
3 | BD Plastipak 5 mL | True
4 | BD Plastipak 10 mL | True
5 | BD Plastipak 20 mL | True
6 | BD Plastipak 30 mL | True
7 | BD Plastipak 50 and 60 mL | True
8 | Monoject 3 mL | True
9 | Monoject 6 mL | False
10 | Monoject 12 mL | False
11 | Monoject 20 mL | True
12 | Monoject 35 mL | False
13 | Monoject 60 mL | False
14 | Terumo 3 mL | True
15 | Terumo 5 mL | True
16 | Terumo 10 mL | True
17 | Terumo 20 mL | True
18 | Terumo 30 mL | True
19 | Terumo 60 mL | True

## Master PCA Syringe List
Master PCA Syringe List | Master PCA Syringe List | Master PCA Syringe List
Line Item | PCA Syringe Type | Included?
1 | BD Plastipak 20 mL | True
2 | BD Plastipak 30 mL | True
3 | BD Plastipak 50 and 60 mL | True
4 | Monoject 20 mL | True
5 | Monoject 35 mL | True
6 | Monoject 60 mL | True
7 | Terumo 20 mL | True
8 | Terumo 30 mL | True
9 | Terumo 60 mL | True
10 | IMS 30 mL Prefill | True